# BD MAX™ MRSA Assay Procedure

PURPOSE

- This procedure provides instructions for preparing samples and testing on the BD MAX™ instrument

#### POLICY STATEMENT

- MRSA PCR testing is performed 1 – 2 times daily, 0700 –1530

- Alert Value: PMRSA (first time positive patients) - Report results by telephone to the patient’s caregiver; document

ABBREVIATIONS

| - MRSP: methicillin-resistant S. aureus PCR - IC: internal control - INC: incomplete run - IND: indeterminate - LIS: laboratory information system - MM: master mix - MREJ: SCC mec right extremity junction - NA: nucleic acid - NEGC: negative control | - PCR: polymerase chain reaction - POSC: positive control - pp: primer – probe - RT – Room temperature - UNR: unresolved - Area/Room 1: Clean room - Area/Room 2: Processing room - Area/Room 3: Amplification room |
| --- | --- |

## DOCUMENTATION/RECORDS

- BD MAX run-specific Results Report

- LIS Incomplete and Completed worksheets

- Daily Maintenance Log

## SAFETY CONSIDERATIONS

- Standard precautions. Refer to MB 2.02 Biohazard Containment

- Use of engineering controls: Refer to MB 3.01 Engineering Controls to Prevent Nucleic Acid Contamination

| Equipment | Reagents | Supplies |
| --- | --- | --- |
| Room 2 - Refrigerator 2 – 8° C - -70⁰ C freezer - VWR Multi-tube vortex - BSC BSL-2 Room 3 - BD MAX instrument | BD MAX MRSA Assay kit: Product No. 443461 | Orange barrier wipes |
|  | MRSA positive control | Nitrile gloves (powder-free) |
|  | MSSA negative control | BD MAX PCR cartridges |
|  | Sani-Cloth Bleach wipes | Nalgene cryogenic vial holder |
|  | 70% alcohol | BD MAX Sample rack |
|  | water | Septum Caps |

#### MATERIALS REQUIRED

## QUALITY CONTROL

- Assay Controls

  - A POSC and NEGC must be included in each assay run.

  - An IC is incorporated into each extraction tube

- QC Monitors:

| Control | Control Monitor |
| --- | --- |
| MRSA Positive Control (POSC) | Cell lysis, reagent failure and primer-probe integrity |
| MSSA Negative Control (NEGC) | Reagent and/or environmental contamination, cumulative effect |
| Internal Control (IC) | PCR inhibition in specimen, reagent failure or process error |

- Before reporting patient results, all controls must yield valid results. Refer to MB 10.05, Procedures F and G, Evaluating and Interpreting Results.

PROCEDURE A: Follow the steps in the table below to organize and label samples, prepare specimen for testing

Specimen Preparation

| Activity |  | ###### Step |  | Action | Related Doc |
| --- | --- | --- | --- | --- | --- |
| Note: Create worklist and load strips in racks prior to vortexing step #15. Refer to MRSA 005 Procedures B, C |  |  |  |  |  |
|  | 1 |  | Call worksheet MRSP; use this worksheet for sample identification throughout testing. |  | MB 1.01 Specimen Management |
| Sample Organization Room 2 | 2 |  | Process up to 22 patient samples plus one POSC and NEGC per run. Position samples and controls as follows: Sample; Position. Patient samples; 1 – nn. POSC ( + ); 2nd to last position. NEGC ( - ); Last tube. | MB 3.01 Engineering Controls |
|  | 3 |  | Using the MRSP worksheet as a layout, organize patient samples and labels - Color code worksheets and labels per run - Number patients on worksheet in consecutive order - Number patient labels including 1 aliquot label corresponding to worksheet |  |  |
| Numbering BufferTubes and swabs | 4 |  | Number and label one buffer tube per patient sample, POSC and NEGC - Place Sample buffer tubes in rack - Number each cap and tube consecutively - Place aliquot label on corresponding tube, matching number to the cap - Do not cover barcode |  |  |
|  | 5 |  | Number each patient swab according to MRSP worksheet |  |  |
|  | 6 |  | Place swabs in the rack behind the corresponding tube |  |  |
| Processing Room 2 | 7 |  | Loosen caps on each sample buffer tube, allowing the caps to sit lightly on tube - Only one tube can be open at a time |  |  |
|  | 8 |  | Remove test swab from the sample transport tube |  |  |
|  | 9 |  | Lift cap from corresponding tube and discard |  |  |
|  | 10 |  | Place swab into the sample buffer tube |  |  |
|  | 11 |  | Break swab off in sample buffer as follows: - Using an orange barrier protector, hold the swab near the rim of the tube - Lift the swab up 1 cm from the bottom of the tube - Bend the swab against the edge of the tube to break - Return swab shaft to original transport tube - Discard barrier protector |  |  |
|  | 12 |  | Place septum cap on tube, touching only the edges to secure |  |  |
|  | 13 |  | Repeat steps 8 – 12 for all samples to be tested |  |  |
|  | 14 |  | Change gloves |  |  |
| Vortex Room 2 | 15 |  | Vortex tubes in rack for 1 minute (Speed 10, Timer 10) - Place protective barrier on top of septum caps prior to vortexing - Perform assay immediately after vortexing step, Procedure D |  | MB 10.05 Procedure D: Running MRSA Assay |
| Room 3 | 16 |  | Remove lab coat; move to room 3 with test samples |  |  |

PROCEDURE B: Follow the steps in the table below for logging samples into the system

Log in Sample Tubes

| Activity | ###### Step | Action |  | Related Doc |
| --- | --- | --- | --- | --- |
| Log In | 1 | Log In to the system |  | BD MAX User Manual |
|  | 2 | Enter username and password **Case sensitive |  |  |
| Work List | 3 | Click Run button |  |  |
|  | 4 | Click the Work List tab (see Figure 1) |  |  |
| Method of entry | 5 | Click on position A1; an arrow will appear to the right of the position column (see Fig. 1) To set up Work list; Then. Row by row; Click arrow to point. Column by column; Click arrow to point. |  |
| Select Assay | 6 | Click the Assay field in row A1 |  |  |
|  | 7 | Select BD MAX MRSA XT from drop down box; all Rack A fields will populate |  |  |
| ###### Figure 1: Work List Screen |  |  |  |  |
| 14 5 |  |  |  |  |
|  | 8 | Sample Tube Barcode field: Scan Sample buffer tube barcodes just prior to starting run, Procedure D |  |  |
| Assigning test position | 9 | Click in the Patient ID field and type in the assigned position number 1 – nn corresponding to MRSP worksheet/label/tube |  |  |
|  | 10 | Click in the Accession field and scan the corresponding barcode label - Controls: type in MRSA/MSSA followed by date as a unique identifier, i.e., MRSA 09.13.16 |  |  |
|  | 11 | Select positive or negative from the drop down box in the External Control field to designate type of control |  |  |
| Lot Number | 12 | In the Lot Number field, select the lot number in use from the drop down box to populate all rows in Rack A Note: If second lot is necessary, note location and scan into the work list |  | MB 10.10 Creating Lot Numbers in BD MAX |
|  | 13 | To continue in Rack B, click on position B1 for the remaining samples - Refer to steps 7 – 12 |  |  |
| Clearing information for (and below) a selected line | 14 | To clear information, click the desired row once; an “ X” appears right of the Position field |  |  |
|  | 15 | Click the “X” button twice to clear information in that row and the rows below |  |  |
|  | 16 | If additional samples are added, re-select the assay and lot numbers |  |  |

PROCEDURE C: Follow the steps in the table below for preparation of reagents and racks

BD MAX™ prepare racks

| Activity | ###### Step | Action |  | ##### Related Doc |
| --- | --- | --- | --- | --- |
|  |  | Note: For optimal performance, run samples in multiples of 4 when possible |  |  |
| Gloves | 1 | Clean gloves must be used for handling all reagents |  |  |
|  | 2 | Remove the required number of reagent strips from the MRSA kit. |  |  |
| Reagent strips Room 3 | 3 | Position strips in the sample rack according to work list - Place the pipette tips toward the front - Seat the back of the strip under the back rail of the sample rack and front edge under the front lip - Pull strip back slightly to make sure it is seated properly |  |  |
|  | 4 | Tap rack on a hard surface to ensure all liquids are at the bottom of the tubes |  |  |
| Extraction and MM tubes | 5 | Snap the MRSA extraction tube into position 1 (white) of each reagent strip | Figure 2. Snap in reagents 6 - green 5 - white |  |
|  | 6 | Snap the MRSA MM tube into position 2 (green) of each reagent strip |  |  |
|  | 7 | Reseal securely each reagent pouch containing desiccant and remaining tubes |  |  |
|  | 8 | Use MM and extraction reagents within 3 h after removal from pouch |  |  |

PROCEDURE D: Follow the steps in the table below for running the BD MAX MRSA assay

Running the BD MAX MRSA assay

| Activity | ###### Step | Action |  |  |  | Related Doc |
| --- | --- | --- | --- | --- | --- | --- |
| Room 3 | 1 | Place Sample Buffer Tubes into the Sample Rack in the assigned positions Note: Face buffer tube barcodes outward for scanning |  |  |  |  |
|  | 2 | Click in the Sample Tube Barcode field on worklist |  |  |  |  |
|  | 3 | Place tubes into Racks A and B following the same sample order as the work list |  |  |  |  |
|  | 4 | Click on Unlock Door button, push down on lid gently and lift to open - The door is unlocked for 15 seconds; after 15 seconds the door relocks |  |  |  |  |
| Sample Racks | 5 | Place sample racks into system - Rack A on left, Rack B on right - The racks have pegs to guide and secure rack in place - Lower handle to the back of the rack once in place |  |  |  |  |
| Insert Cartridges | 6 | Place PCR cartridges into the cartridge drawer - Position the notched edge into the upper left corner - Make sure the cartridge lies snuggly on the reader drawer - Cartridges can be used 2X |  | Figure 3. Inserting cartridge |  |  |
| Begin Run | 7 | Close the door |  |  |  |  |
|  | 8 | Click the Start Run button |  |  |  |  |
| Initialize | 9 | Instrument will perform a catalog check, reading barcodes on reagent strips and tubes |  |  |  |  |
| Catalog check errors | 10 | The Status screen automatically appears as the run begins. If there are any errors in the catalog check, the alert icon flashes red and a red alert icon appears for that consumable - Run will stop - Click on flashing icon to display message and to acknowledge the alert Alerts button information - Blue: no active alerts - Flashing red: active unacknowledged alerts - Red: active acknowledged alert | Figure 4. Catalog errors |  |  |  |
| Run Status | 11 | Alerts Display - Expand alert field to get text describing the alert, click thebutton - Persistent errors must be corrected - Non-persistent errors can be cleared from the display |  |  | Figure 5. Alerts Display |  |
|  | 12 | Catalog errors must be corrected before the run can continue |  |  |  |  |
|  | 13 | Click the unlock door to correct error |  |  |  |  |
| Correct Catalog Errors | 14 | Correct source of error (examples) - Replace cartridge - Missing reagent tube - Swapped extraction and PCR reagents - Sample buffer tubes in wrong location |  |  |  |  |
|  | 15 | Once corrected, close door and click Start Run to resume testing - Warning: never try to open door when the “Unlock Door” is grayed out |  |  |  |  |
|  | 16 | Monitor Status display for successful catalog check - All reagents and strips are green |  |  |  |  |
| Run icons - Rack - Robotic pipettor - PCR | 17 | Figure 6: Monitoring Run Status while in progress Reader A ¦ Robot ¦ Reader B Rack PCR/Reader - Rack icon: rack icon may be level, tilted to the left or right, or grayed out if rack not present. If tilted, an alert icon will appear over the rack - Robotic pipettor & PCR icons: when active, moving circles appear; alerts display above the icon |  |  |  |  |
| Run Completed | 18 | Run time approx. 2 h |  |  |  | MB 10.11 Interleaved Runs |
|  | 19 | Log off when testing is completed |  |  |  |  |

PROCEDURE E: Follow the steps in the table below for rack removal and sample storage

Rack removal and sample storage

| Activity | Step | Action | Related doc |
| --- | --- | --- | --- |
| Status | 1 | Racks can be remove when: - The run is completed or - The robot status is idle and the PCR/reader status displays a moving circle Figure 7. Status Messages Status messages in light blue bars |  |
| Completed |  |  |  |
|  | 2 | Click on Unlock Door button, open door |  |
|  | 3 | Remove racks containing reagent strips from the run |  |
| Room 3 | 4 | Remove reagent strips from racks and place in a Ziploc bag; dispose bag in red biohazard container in room 3 |  |
|  | 5 | Positive patient samples: Replace septum cap with screw cap; hold in designated rack in refrigerator for 1 month in room 2 |  |
|  | 6 | Hold negative sample buffer tubes in room 3 organized by Run until end of day; if no problems identified, discard tubes in room 3 |  |
|  | 7 | Remove PCR cartridges if used 2X and place in Ziploc bag for disposal |  |
|  | 8 | If another run can be performed on the cartridge, load the new sample racks on the instrument deck |  |
|  | 9 | If there are no more runs, store cartridges in instrument until next run (good for 7 days) |  |
| Cleaning | 10 | Clean room and equipment; refer to Decontamination procedure for schedule | MB 10.08 Equipment and Room Decontam |

PROCEDURE F: Follow the activities below for PCR Analysis and Printing Results

PCR Analysis and Printing Results

| Activity | Step | Action |  | Related Doc |
| --- | --- | --- | --- | --- |
|  | 1 | ### Click the Results button ; the run list appears, Fig. 8 |  |  |
|  | 2 | ### Any runs in progress will display the run icon |  |  |
|  | 3 | ### Click the desired run from the list |  |  |
|  | 4 | ### Click the Print button; the print preview window appears, Fig. 9 |  |  |
|  | 5 | ### Wait for the Print Preview to fully populate |  |  |
|  | 6 | ### Click Print for report and graphs Analyte; Instrument Channel. MREJ (Staphylococcus aureus specific); 475/520. Methicillin resistance gene (mec A or mec C); 585/630. Internal control; 680/715. |  |
|  | 7 | ### Change gloves and lab coat before leaving room 3 |  |  |
| Figure 8: Results List |  |  | Figure 9: Print Preview |  |

PROCEDURE G: Follow the activities below for evaluating the acceptability of patient results

Evaluating and Interpreting QC and Patient Results

| Activity | Step |  | Action |  | Related doc |
| --- | --- | --- | --- | --- | --- |
|  | 1 |  | ### Check QC for acceptability before reporting patient results |  |  |
| QC and Patient acceptability | 2 |  | If | Then | BD MAX User Manual 7.2 System Problems 7.4 Error list MB 10.06 Trouble-shooting Guide Refer to Procedure H: Repeat Testing |
|  |  |  | Valid assay: Controls as expected | - Report patient results |  |
|  |  |  | Invalid assay: POSC/ NEGC fail | - Do not report patient results - Review the sample handling/ preparation technique - Repeat testing using failed control and a new control |  |
|  |  |  | NEGC positive | - Do not report patient results - Review the sample handling/ preparation technique - Review specimen mix-up or contamination - Repeat testing, refer to Procedure H - Tech spec/director review required – If positive patient in failed run and not a previous positive, recollect patient sample |  |
|  |  |  | External control: Unresolved, Indeterminate or Incomplete result | - Failure caused by reagent or system failure - Check for error messages - Repeat testing - Problem repeats, refer to trouble-shooting guide |  |
|  |  |  | Problem unresolved | - Call BD technical service, 1-800-638-8663, option #2 - Notify Technical Specialist |  |
| Problem |  | 3 | ### Record problem/action in the QC failure log. |  |  |
| Interpretation |  | 4 | ### The system software will automatically interpret the specimen results; Refer to Table 1 |  | Refer to Table 1 and MB 10.07 For Reporting |
| Results Interpretation and entry |  | 5 | ### Enter results in Sunquest Microbiology Result Entry using Culture Mode |  |  |
|  |  | 6 | ### Alert Value: PMRSA (first time positive patients) - Report results by telephone to the patient’s caregiver; document |  |  |
| Review |  | 7 | ### Call completed MRSP worksheet to review reporting. |  |  |
| Archive |  | 8 | ### Initial LIS result worksheet and SC run specific report. |  |  |
|  |  | 9 | ### Place reports in MRSA result log book. |  |  |

###### Table 1: Interpretation of results – For additional information, refer to MB 10.07 Reporting and Archiving Results and

MB 10.05, Procedure H: Repeat Testing

| Sample Type | Test Result | Sunquest Code | Repeat testing | Interpretation of result |
| --- | --- | --- | --- | --- |
| Patient Samples |  | NMRSA |  | No MRSA DNA detected, MRSA nasal colonization unlikely Target; MREJ; mec A/ mecC; IC. NMRSA; -; -; +. NMRSA (mec A / C drop out); +; -; +. NMRSA (S. epi); -; +; +. |
|  |  | PMRSA |  | MRSA DNA detected, MRSA nasal colonization Target; MREJ; mec A/ mecC; IC. NMRSA; +; +; NA. |
|  | UNR | UNR | √ | Unresolved – inhibitory specimen or reagent failure - No target or IC amplification - If sample remains UNR, call caregiver for new collection |
|  | IND | ICLA or EQ | √ | Inconclusive results, repeat testing [free text: to be performed on date] - Sample was indeterminate or had a LLS error; an instrument error occurred for this sample - Refer to BD MAX User Manual for interpretation of error codes - Repeat testing - If sample remains IND, report EQ (equivocal result) |
|  | INC | ICLA or EQ | √ | Inconclusive results, repeat testing [free text: to be performed on date] - Sample prep or PCR did not meet expected time limits - Refer to BD MAX User Manual for interpretation of error codes - Repeat testing - If sample remains INC, report EQ (equivocal result) |
| POSC |  | ------- |  | Valid POSC; valid run when NEGC is also valid. |
|  |  | ------- | √ | Invalid POSC; invalid run. Patient results cannot be reported. |
| NEGC |  | ------- |  | Valid NEGC; valid run when POSC is also valid |
|  |  | ------ | √ | Invalid NEGC; invalid run. Patient results cannot be reported. |

INC – Incomplete run; IND – Indeterminate; UNR – unresolved

PROCEDURE H: Follow the activities below for repeat testing

Repeat Testing

| Activity | Step | Action | Related doc |
| --- | --- | --- | --- |
| Type of Failure | 1 | ### For repeat testing, refer to following table for appropriate sample. Repeat testing can be performed one time from Sample buffer tube. Failure; Sample Buffer Tube; Recollect New sample; New Controls. Negative POSC; √ and; √. Positive NEGC; √ and; √. UNR; √; √ if 2X. INC; √; √ if 2X. IND; √; √ if 2X. POSC is negative: repeat all patient samples; √. NEGC is positive: patient is positive; Previous positive √; Newly identified positive √. NEGC is positive: patient is negative; √. | MB 10.06 Troubleshooting BD MAX User Manual 7.2 System Problems 7.4 Error list |
| Timeframe | 2 | Retesting from RT sample buffer tube must be performed within 36 hours after the end of the run or within 5 days if stored at 2 – 8⁰ C |  |
|  | 3 | Repeat samples may be retested in the same run as new samples |  |
| Vortex | 4 | Vortex the sample buffer tube 1 min prior to retesting |  |
| Error Codes | 5 | For interpretation of error codes or warnings, refer to Troubleshooting guide |  |

#### METHOD PERFORMANCE

- Clinical Sensitivity/Specificity: 93% / 95.9%

PROFICIENCY TESTING

- WSLH PT – MR: MRSA/VRE screen

#### ALTERNATE METHOD

- MRSA Culture/Surveillance MRSA Culture

- Sunquest Order code: MRSS, SMRSS

- Logistics:

  - Acceptable specimen: Anterior nares

  - Transport:

- BD CultureSwab

- RT

## LIMITATIONS

- A positive test result does not necessarily indicate the presence of viable organisms.

- This assay does not detect the mecA gene directly or the penicillin binding protein (PBP 2a) encoded by this gene. A false positive MRSA result may occur if an “empty cassette” S. aureus variant is present.

- This assay may cross-react with some strains of MSSA when they are present in extremely high concentrations.

- Excess blood in the specimen may inhibit the assay.

- Low levels of target below the LoD of the assay may be detected, but results may not be reproducible.

REFERENCES

- BD MAX™MRSA Assay, Product Circular Reference 443461, P0167(04), 2016-02, GeneOhm Sciences Canada, Inc, 2555Boul. Du Parc Technologique, Quebec (QC), G1P 4S5, Canada

- BD MAX™ System User’s Manual (US IVD Version), 2014-06 Document Number: 8089572(01), Becton Dickinson and Company, 7 Loveton Circle, Sparks, Maryland 21152 USA

| Historical Record |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Version | Written/Revised by: | Effective Date: | Summary of Revisions |
|  | 1 | Pat Ackerman | 6/10/2010 | Initial Version |
|  | 2 | P. Ackerman | 12.12.13 | Updated with BD MAX assay information |
|  | 3 | P. Ackerman | 5.08.14 | Added Procedure H: Repeat testing |
|  | 4 | P. Ackerman | 09.14.2016 | Reformatted for CMS upload; updated logo; revised table 1; updated screen shots; renumbered prev MRSA 005 |
|  | 4 | J. Laramie | 09.14.2016 | Biennial review: 04.05.2018, JL |
|  | 5 | J. Laramie | 09.03.2018 | Updated alert value notes to specify first time patients only |