| **Magnesium on Abbott** |
| --- |
| **Purpose** | This procedure provides instructions for MAGNESIUM ON ABBOTT INSTRUMENTATION. The magnesium method is an *in vitro* diagnostic test for the quantitative measurement of magnesium in human serum and plasma on the Abbott Architect c4000 or Abbott Alinity c automated chemistry analyzers. |
| **Policy Statements** | This procedure applies to all personnel responsible for operating the Abbott Architect c4000 or Abbott Alinity c at Children’s Minnesota Laboratory. |
| **Principle** | Magnesium present in the sample is a cofactor in an enzymatic reaction with isocitrate dehydrogenase. The rate of increase in absorbance at 340 nm, due to the formation of NADPH, is directly proportional to the magnesium concentration. Isocitrate dehydrogenaseD-isocitric acid + NADP → 2-oxoglutarate + CO2 + NADPH Mg2+Methodology: Enzymatic |
| **Clinical Significance** | Magnesium is an essential nutrient which is involved in many biochemical functions. It has a structural role in nucleic acids and ribosomal particles, is required as an activator for many enzymes, and has a role in energy producing oxidative phosphorylation. The normal body contains 21 to 28 g magnesium, more than 50% of which is complexed with calcium and phosphate in bone. Only approximately 1% of the total magnesium is found in the extracellular fluid. It tends to enter and leave cells under the same conditions as potassium. Approximately 35% of plasma magnesium is protein-bound, mainly to albumin, and therefore changes in albumin concentration may affect magnesium. Hypomagnesemia results in impairment of neuromuscular function, carbohydrate intolerance, and cardiac arrhythmias. Hypermagnesemia results in hypotension, bradycardia, and respiratory depression, among other conditions. |
| **Analyzer** | **Minneapolis: Abbott Alinity c (Sunquest method code: MALIC)****St. Paul: Abbott Architect c4000 (Sunquest method code: ARCH4)** |
| **Sunquest Test Codes** | **MG**: Magnesium in serum or plasma |
| **Specimen** | Sample: Plasma (lithium heparin with or without gel) preferred. Sodium heparin plasma or serum (with or without gel) are also acceptable. Refer to specimen collection procedures.**Minimum sample volume:** 200 µL preferred, 150 µL minimum**Stability when separated from cells/gel:** RT / 7 day, 2-8 °C / 7 days, < -20°C / 1 year**Rejection criteria:** Unlabeled tube, sample type other than serum or heparinized plasma.**Preparation:** 1. Whole blood specimens should be centrifuged following complete clot formation, according to Specimen Processing procedures prior to analysis.
2. Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection.
3. Specimens should be free of particulate matter.
4. Transfer serum or plasma to a properly labeled pilot tube or aliquot cup. Architect and Alinity systems utilize a specimen level detect mechanism, so special racks specific to tube-type are not required.
5. Minimum labeling includes sample accession ID, and/ or patient name, medical record number, collection date and time.
 |
| **Reagents** | **Alinity c and Architect c4000:****Reagent Handling** Upon receipt, place reagent cartridges in an upright position for 8 hours before use to allow bubbles that may have formed to dissipate. If a reagent cartridge is dropped, place in an upright position for 24 hours before use to allow bubbles that may have formed to dissipate. Reagents are susceptible to the formation of foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results. Use a pipette to remove all bubbles prior to loading on the Alinity or Architect system.* Do not use reagents beyond the expiration date.
* Do not pool reagents within a kit or between kits.
* Do not use components from one lot with components from another lot.

**Alinity c:**Reagents may be stored on or off the system. If removed from the system, store reagents with new replacement caps in an upright position at 2 – 8°C. For reagents stored off the system, it is recommended that they be stored in their original trays or boxes to ensure they remain upright.**Alinity c:**

|  |  |  |
| --- | --- | --- |
| ***Product Description*** | ***Product Code*** | ***Stability*** |
| Abbott Alinity c Magnesium ReagentCHC# 33276 | 08P19-20 | **Store at:** 2 – 8°C**Unopened:** Manufacturer’s printed expiration date**On-board: 30 days****Opened, off the analyzer (with clean caps):** Manufacturer’s printed expiration date. (Reagents may be stored on or off the system. The system tracks time onboard.) |
| Abbott Alinity c Multiconstituent CalibratorCHC# 32633 | 08P60-01 | **Store at:** 2 – 8°C**Unopened:** Manufacturer’s printed expiration date**On board expiration: 5 days when stored onboard.** The Alinity c tracks time on the system.**Opened expiration: 7 days** when opened and stored off the system. |

 |
|  | **Architect c4000:**

|  |  |  |
| --- | --- | --- |
| ***Product Description*** | ***Product Code*** | ***Stability*** |
| Abbott Architect Magnesium ReagentCHC# 32551 | 07D70-21 | **Store at:** 2 – 8°C**Unopened:** Manufacturer’s printed expiration date.**On-board: 30 days** |
| Abbott Architect Multiconstituent CalibratorCHC# 32557 | 01E65-05 | **Store at:** 2 – 8°C**Unopened:** Manufacturer’s printed expiration date**Opened:** 7 Days |

 |
| **Risk and Safety** |

|  |
| --- |
| **CAUTION:** For in vitro diagnostic use. This product requires the handling of human specimens. It is recommended that all human sourced materials be considered potentially infectious and be handled in accordance with the OSHA Standard on Bloodborne Pathogens. Appropriate Personal Protective Equipment (PPE) must be worn according to Children’s Minnesota Laboratory policies. Current SDSs are kept on the [Children’s StarNet](https://msdsmanagement.msdsonline.com/a07dc954-23d8-42a9-b591-ef5763cdfd33/ebinder/?nas=True) page.**Alinity c** and **Architect c4000:**R1 Reagent: Contains sodium azide and polyethylene glycol octylphenyl ether. Contact with acids liberates very toxic gas. Harmful to aquatic life. Harmful to aquatic life with long lasting effects.R2 Reagents: Contains sodium azide. Contact with acids liberates very toxic gas.Reagents and Multiconstituent Calibrators should be disposed of in Regulated Medical Waste (red trash). |

 |
| **Calibration** | **Alinity c** and **Architect c4000:**

|  |  |
| --- | --- |
| Assay Range: | 0.6 – 9.5 mg/dL |
| Reference Material: | Abbott Alinity Multiconstituent CalibratorAbbott Architect Multiconstituent Calibrator |
| Suggested Calibration Levels: |  See lot-specific assay set point documentation |
| Calibration Scheme: | 2 levels, Linear data reduction method |
| Calibration Frequency: | 30 Days |
| AMR | AMR is verified twice annually using the Maine Standards GC1 Product # 1100ab by running all applicable levels in triplicate. Assay results are submitted to Maine Standards for compilation and comparison to peers. Results are reviewed and approved by the Technical Specialist. Any questionable results are investigated and corrective actions documented.  |

 |
| **Quality Control** | **Alinity c and Architect c4000:** Bio-Rad Liquichek™ Unassayed Chemistry Control (Human) Levels 1 & 2**Frequency:** Two levels each day of use**Stability:** Once thawed, opened, and stored tightly capped at 2 to 8°C, this product will be labeled with an expiration date equal to the shortest stability of the included analytes, which is **6 days.****Preparation**: This product should be treated the same as patient specimens and run in accordance with the instructions accompanying the instrument, kit, or reagent being used. * To thaw the product, allow it to stand at room temperature (18° to 25°C) until completely thawed but no longer than one (1) hour.
* After thawing, the product **MUST** be gently swirled and inverted several times to ensure homogeneity.
* For optimal analyte stability in the thawed state, promptly return to 2 to 8°C storage after each use and minimize the time at room temperature to no more than 20 minutes daily.
* **Before each use**, gently swirl the contents until homogeneous with no visible signs of precipitate.

**Acceptable ranges:** * Non-Bio-Rad controls will utilize manufacturer ranges and 2 SD Westgard rules.
* New lots of Bio-Rad controls should be run for 20 days in parallel with the current lot whenever possible prior to switching to the new lot.
* Refer to the Westgard Rules in Chemistry procedure (insert hyperlink) for current Westgard rules in place for each analyte.
* **Acceptable ranges are current in Unity Real Time only.** Quality Control results must be rejected in Sunquest when the results cross the interface.
* In the event of a QC failure, refer to the [Unity Real Time QC Review, General User](https://starnet.childrenshc.org/References/labsop/chem/quality/ch-2.17-unity-real-time-qc-review-general-user.pdf) and navigate to the QC Troubleshooting section.
* Do not load or release patients until QC is acceptable in Unity Real Time.
 |
| **Interferences** | **Alinity c and Architect c4000:** Interference studies were conducted by Abbott Diagnostics Division based on guidance from NCCLS EP7-P and effects were assessed by Dose Response and Paired Difference methods at the medical decision levels of the analyte. Interference is less than 7.5% at:* Hemoglobin: up to 1200 mg/dL
* Bilirubin: up to 60 mg/dL
* Triglycerides: up to 2000 mg/dL

Interferences from medication or endogenous substances may affect results.For accurate results, serum and plasma specimens should be free of fibrin, red blood cells, and other particulate matter. Serum specimens from patients receiving anticoagulant or thrombolytic therapy may contain fibrin due to incomplete clot formation. See Abbott package inserts for all potentially interfering substances tested. |
|  | **Alinity c and Architect c4000:**  |
| **Reference Intervals** |

|  |  |
| --- | --- |
| Age | Magnesium |
| 0 – 14 days | 1.8 – 4.1 mg/dL |
| 15 - 364 days | 1.8 – 3.2 mg/dL |
| 1 – 18 years | 2.0 – 2.9 mg/dL |
| Adult | 1.8 – 4.5 mg/dL |
|  |

 |
| **Critical Values** | < 1.2 or > 6.6 mg/dL.Critical values must be called and documented according to the critical values policy. |
| **Limitations** | The instrument reporting system contains flags and comments to provide the user with information regarding instrument processing errors, instrument status information and potential errors in magnesium results. Refer to the [Abbott Architect](https://starnet.childrenshc.org/References/labsop/chem/operator/abbott-architect-operations-manual.pdf) or [Alinity](https://starnet.childrenshc.org/References/labsop/chem/operator/alinity-ci-series-operations-manual.pdf) Operator’s Guides for the meaning of report flags and comments, and instructions for addressing them. Do not report results until a report containing flags and/or comments is resolved.For accurate results, serum and plasma specimens should be free of fibrin, red blood cells, and other particulate matter. Serum specimens from patients receiving anticoagulant or thrombolytic therapy may contain fibrin due to incomplete clot formation. |
| **Dilutions** |

|  |
| --- |
| **Alinity c and Architect c4000:**  |
| Auto Dilution: | None |
| Maximum Manual Dilution: | 1:2 |
| Diluent: | Saline |
| Manual Dilution: | Follow Abbott [Architect Operator’s Manual](https://starnet.childrenshc.org/References/labsop/chem/operator/abbott-architect-operations-manual.pdf) or [Alinity Operator’s Manual](https://starnet.childrenshc.org/References/labsop/chem/operator/alinity-ci-series-operations-manual.pdf) instructions for programming manual dilutions. The operator must enter the dilution factor when ordering the manual dilution. The system will use this dilution factor to automatically calculate the concentration of the sample and report the result.If a diluted sample result is less than the lower value of the measuring interval of 0.6 mg/dL, do not report the result. Rerun and/or investigate for other possible causes of error. |

 |
| **Result Reporting** | **Alinity c and Architect c4000:** * Results between 1.2 mg/dL – 6.6 mg/dL without error messages are released
* Results < 1.2 mg/dL and > 6.6 mg/dL must be called and documented according to the critical values policy.
* Results < 0.6 mg/dL are reported as <0.6 mg/dL.
* Results > 9.5 mg/dL should be diluted 1:2 according to the dilution policy above. Results > 19.0 mg/dL after manual dilution are reported as > 19.0 mg/dL
 |
| **Specimen Storage** | Promptly stopper tested specimen and store upright in a specimen rack. Every 8 hours remove specimens to refrigerator/freezer storage. Samples are retained 7 days in specimen storage freezer. |
| **References** | 1. Jacobs & DeMott Laboratory Test Handbook, Lexi-Comp, Inc., Hudson, OH, 5th Edition, 2001
2. Architect Magnesium Package Insert, Abbott Laboratories Diagnostics Division, Abbott Park, IL 60064, November 2018.
3. Alinity Magnesium Package Insert, Abbott Laboratories Diagnostics Division, Abbott Park, IL 60064, November 2018.
4. Alinity c Multiconstituent Calibrator Package Insert, Abbott Laboratories Diagnostics Division, Abbott Park, IL 60064, December 2017.
5. Architect Multiconstituent Calibrator Package Insert, Abbott Laboratories Diagnostics Division, Abbott Park, IL 60064, August 2017.
6. Bio-Rad Liquichek Unassayed Chemistry Control Product Insert, Bio-Rad Laboratories, Irvine, CA 92618
7. CALIPER pediatric reference range database. (2019). Retrieved October 3, 2019, from https://caliper.research.sickkids.ca/#/
 |
| **Historical Record** | **Version** | **Written/Revised by:** | **Effective Date:** | **Summary of Revisions** |
| 1 | Erin Bartos | 10/15/2019 | New Procedure for Abbott analyzers |
|  |  |  |  |