# RVP Control and Reagent Preparation

**PURPOSE**

* This procedure provides instructions for preparation of reagents and procedural controls

#### ABBREVIATIONS

* Ct: crossing threshold
* EXC: extraction control
* FABR: Flu A, B & RSV PCR
* Hyb: hybridization solution
* IC: internal control
* MM: master mix
* NA: Nucleic Acid
* NEGC: negative control
* NFW: nuclease free water
* RT-PCR: reverse transcription polymerase chain reaction
* PCTL: process control
* POSC: positive control
* RT: room temperature
* RVP: Respiratory Viral Panel
* VTM: viral transport media

Area/Room 1: Clean room

Area/Room 2: Processing room

Area/Room 3: Amplification room

#### MATERIALS REQUIRED

| **Equipment** | **Reagents** | **Supplies** |
| --- | --- | --- |
| Room 1 * Adjustable pipettes
* Cold block
* Freezer, -20° C
* Laminar air-flow hood
* Refrigerator 2 – 8° C
* Vortex mixer

Room 2* Adjustable pipettes
* BioHit 8 channel pipette
* Bio-Safety Cabinet (BSC)
* Cold Block
* Freezer, -70° C
* Magnetic rack
* Mini-centrifuge
* NucliSens easyMag
* Refrigerator 2 – 8° C
* Tube racks, 1.5 – 2 ml
* Vortex mixer

Room 3* Adjustable pipettes
* Cold Block
* Freezer, -20° C
* GenMark eSensor XT-8 instrument
* Mini-centrifuge
* PCR thermocycler
* PCR workstation
* Vortex mixer
 | eSensor *RVP* kit: Product No. MT005102 | Sterile filtered 10 μl pipette tips |
| easyMAG Lysis buffer, 2 ml | Sterile filtered 30 μl pipette tips |
| easyMAG Buffer 1 | Sterile filtered 100 μl pipette tips |
| easyMAG Buffer 2 | Sterile filtered 200 μl pipette tips |
| easyMAG Buffer 3 | Sterile filtered 1000 μl pipette tips |
| MagSil | Micro tubes 1.5 ml, RNase/DNase free |
| Molecular grade water, nuclease free | Nitrile gloves (powder-free) |
| Viral transport media (VTM) | PCR 8 tube strips with caps |
| Viral isolates: H1, H3, RSV, Flu B | easyMag disposable vessel strips and tips |
| Patient hMPV sample | BioHit pipette tips |
| Sani-Cloth Bleach Wipes (10%) | BioHazard wipes |
| 70% alcohol | Gripper rack |
| Household bleach | Sharps disposal container |
| MMQCI RVP Control Panel  |  |
| Extran 300 |  |
|  |  |
|  |  |
|  |  |
|  |  |
|  |  |

**PROCEDURE A:** Follow the activities in the table below for preparing RVP assay controls

Preparing RVP Control Panel, Positive/Extraction Controls, Internal Control, and Negative Control

| Control | Step | **Action** |
| --- | --- | --- |
| **MMQCI RVP Control Panel3** | 1 | The RVP control panel consists of 2 vials M244 and M245, single use only |
| Room 2 | 2 | Allow the vials to warm to RT |
| 1X use | 3 | Vortex each vial for 5 s prior to use |
|  | 4 | Spin for 5 s to pull down matrix |
|  | 5 | Extract 200 µl supernatant including 10 µl IC using the EasyMag; final elution 60 µl |
|  | 6 | Vortex the eluate for 5 sec; allow to sit in magnetic rack for 10 min |
| **MMQCI Testing schedule** | 7 | Analyze RVP control panel according to the RVP protocol* Test monthly alternating M244 and M245; record results MB 11.08.F4
 |
|  | 8 | Freeze eluates at -70⁰ C |
| **Internal Control** | 1 | No preparation necessary |
| Room 2 | 2 | Remove Internal Control (IC) from -70⁰ C freezer; warm to RT before use |
|  | 3 | Add 10 µl to each sample in easyMag vessel to be extracted  |
|  | 4 | Mark cap after each use representing one F/T; return to -70⁰ C freezer  |
| F/T 5X | 5 | Freeze/thaw cycles up to 5X |
| **Negative Control** | 1 | Aliquot 300 µl VTM in 1.5 ml micro-centrifuge tubes |
| **(NEGC)** | 2 | Label tubes with NEGC and prep date using preprinted labels |
| Room 1 | 3 | Label box with VTM lot number and expiry date |
|  | 4 | Store at 2 – 8° C in room 1 |
|  | 5 | Store a 1 week working supply in room 2 |
| **RSV, 2009 H1, H3, Flu B, and hMPV****PCTL/EXC**Prepare virusVirology Lab | 1 | Cultivate from stock viral suspensions

|  |  |  |  |
| --- | --- | --- | --- |
| Virus | Cell Line | CPE | Approx CX days |
| 2009 H1N1 | RMK | 3 – 4+ | 3 |
| Seasonal Flu H3 | RMK | 3 – 4+ | 3 |
| Influenza B | RMK | 3 – 4+ | 3 |
| RSV | Hep-2 | 3 – 4+ | 3-4 |
| hMPV | A549 & MDCK | 1 – 4+ | 3-4 |

 |
|  | 2 | Scrape down cell culture tube to make a new stock suspension |
|  | 3 | Serially dilute each stock suspension using NFW to prepare a 10-4 working dilution (total volume approx 25 ml) |
| **Working suspension** | 4 | Add 5 ml of VTM to the suspension; mix well |
| Room 2 | 5 | Extract 200 µl of the working dilution (each control) |
|  | 6 | Perform an RVP test to determine the nA value Target nA values\*:

|  |  |
| --- | --- |
| Virus | nA |
| Flu A / 2009 H1N1 | 80 - 250 |
| Flu A / Seasonal Flu H3 | 80 - 250 |
| Influenza B | 80 - 200 |
| RSV | 80 - 200 |
| hMPV | 80 - 200 |

\*Subject to adjustment – approval by Technical Specialist or Technical Director required |
| **Working suspension cont.** | 7 | If necessary, adjust suspension to obtain projected range with NFW / VTM based on the previous nA value |
|  | 8 | Repeat RVP testing from new suspension. |
| **Aliquot and freeze** | 9 | Label a set of 1.5 mL micro-centrifuge tubes for each extraction control using preprinted labels with prep date |
| Room 2 | 10 | Pipette 1.1 mL of working suspension into tubes |
|  | 11 | Store in –70° C freezer, EXC box |
| **Test aliquots before use** | 12 | Before use:* Thaw one PCTL aliquot
* Test on day 1 and 5 (or 2 if making a hMPV control)
 |
|  | 13 | Document Ct values on RVP PCTL New Reagent Worksheet MB 11.04.F1 |
|  | 15 | Place worksheet and reports in the RVP Binder  |
| **Stability** | 16 | * Flu A 2009 H1N1, Flu A Seasonal Flu H3, Influenza B, RSV
	+ 1 year at ≤70° C
	+ Once thawed, 5 days at 2 – 8° C
* hMPV
	+ 1 year at ≤70° C
	+ Once thawed, 2 days at 2 – 8° C

**NOTE:** Label box with prep and expiration date. |
|  | 17 | Do not refreeze (only 1 F/T cycle) |
| **RVP PCTL/EXC Rotation** | 18 | Rotate EXC as follows:

|  |  |
| --- | --- |
| Order | Viral Extraction Control |
| 1 | 2009 H1N1 |
| 2 | Seasonal Flu H3 |
| 3 | Influenza B |
| 4 | RSV |
| 5 | hMPV |

 |

**PROCEDURE B:** Follow the activity below for preparing master mix (MM)

Preparing RT-PCR Master Mix (MM)

| Activity | Step | **Action** |
| --- | --- | --- |
|  | 1 | ***MM must be used within 30 min of preparation****.* |
| **MM** | 2 | Wear lab coat and gloves dedicated in the Clean room 1. |
|  | 3 | Clean hood and equipment * Extran (5%) followed by 70% alcohol
 |
| Room 1 | 4 | Thaw RVP PCR mix at RT up to 1 h. Place Enzyme mix in a cold block, refrigerated until use. |
|  | 5 | Vortex PCR mix 3 – 5 s, making sure it is completely thawed |
|  | 6 | Centrifuge the enzyme and PCR mix; place both reagents in cold block |
|  | 7 | Prepare MM according to number of reactions needed including POSC and NEGC; Refer to Set-up Table MB 11.04.A1 |
|  | 8 | Vortex MM and centrifuge; place in cold block until use  |
| **Refreeze rgts** | 9 | Refreeze reagents. Place a hatch mark on each cap to represent one F/T cycle (up to 5X) |
| **MM calculations** | 10 | Volume Calculations for MM: N = Total number of reactions in run including POSC and NEGC, MB 11.04.A1

|  |  |  |  |
| --- | --- | --- | --- |
| Component | Volume/reaction | Calculation | Volume (µl) |
| RVP PCR Mix | 28.6 µl | 28.6 \* N \*1.1 = |  |
| RVP Enzyme | 1.4 µl |  1.4 \* N \* 1.1 = |  |
| Total volume | 30 µl |  30 \* N \* 1.1 = |  |

 |

**PROCEDURE C:** Follow the activity below for preparing Hybridization Solution

Preparing Hybridization Solution “Hyb”

| Activity | Step | **Action** |
| --- | --- | --- |
|  | 1 | Clean hood and equipment Sani-Cloth Bleach Wipes (10%) followed by water twice |
| **Hyb solution** | 2 | Thaw Signal buffer, Buffer 1 and Buffer 2 at RT |
| Room 3 | 3 | Vortex and centrifuge or tap lightly  |
|  | 4 | Prepare hybridization buffer according to number of reactions needed; Refer to Hybridization buffer set-up table MB 11.04.A1; stable up to 4 hours at RT |
| **Hyb solution****Cont.** | 5 | Label 2 ml tube “Hyb” (may need to prepare 2 tubes for sufficient volume)

|  |  |
| --- | --- |
| Step | Action |
| a | Add reagents to Hyb tube in order1. Signal buffer
2. Buffer 1
3. Buffer2 (white precipitate will appear after addition)
 |
| b | Vortex at setting 10 for 3 – 5 s to clear precipitate |
| c | Centrifuge 3 – 5 s |
| d | ***Note***: Warm with hands if precipitate does not disappear; vortex |

 |
|  | 6 | Mark the cap of the buffer tubes to represent one F/T cycle |
|  | 7 | Change gloves; return detection reagents to -20⁰ C freezer |
| **Hyb solution calculations** | 8 | Volume Calculations for Hyb solution: N = Total number of reactions in run including POSC and NEGC, MB 11.04.A1

|  |  |  |  |
| --- | --- | --- | --- |
| Component | Volume/reaction | Calculation | Volume (µl) |
| RVP Signal Buffer | 70 µl | 70 \* N \* 1.1 = |  |
| Buffer 1 | 10 µl |  10 \* N \* 1.1 = |  |
| Buffer2  | 20 µl |  20 \* N \* 1.1 = |  |
| Total Hyb solution volume = |  |

 |

**PROCEDURE D:** Follow the activity below for preparing miscellaneous reagents

Preparing miscellaneous reagents

| Reagent | Step | **Action** |
| --- | --- | --- |
| **10% Bleach**  | 1 | Prepare in dish room. Made daily.  |
| Dish room | 2 | Make working solution as follows:

|  |  |  |
| --- | --- | --- |
| Step | Reagent (10% bleach) | Volume |
| 1 | House hold bleach (5 – 6 %) | 500 ml |
| 2 | Water | 2000 ml |
| 3 | Alconox (add for contamination clean-up) | 25 g |

Labeling:Refer to: MB 2.01 Safe Work Practices in Molecular |
| **70% alcohol** | 1 | Prepare from 100% Dehydrant alcohol located in the Flammable cabinet in the Recycling room. |
| Room 3 or Recycling room | 2 | Make working solution as follows:

|  |  |  |
| --- | --- | --- |
| Working Volume | 100% Dehydrant | Water |
| 1000 ml | 700 ml | 300 ml |

Labeling:Refer to: MB 2.01 Safe Work Practices in MolecularExpiration: 1 year  |
| **5% Extran** |  | Make working solution as follows:

|  |  |  |
| --- | --- | --- |
| Step | Reagent (Extran 300) | Volume |
| 1 | Extran 300 | 50 ml |
| 2 | Water | 950 ml |

Labeling:Refer to: MB 2.01 Safe Work Practices in MolecularExpiration: 1 year |

**REFERENCES**

1. eSensor® Respiratory viral Panel, PI1032 REV:D, December 2013, Clinical Micro Sensors, Inc. dba GenMark Diagnostics, Inc., 5964 La Place Court, Carlsbad, CA 92008, 1-800-373-6767, ww.genmarkdx.com
2. NucliSENS® Lysis Buffer, product circular 14900 E, 200292, September 2009.
3. eSensor XT-8 RVP Control Panel package insert; circular M243 102914.001, Maine Molecular Quality Controls, Inc. [www.mmqci.com](http://www.mmqci.com)

|  |  |
| --- | --- |
| Historical Record |  |
|  | **Version** | **Written/Revised by:** | **Effective Date:** | **Summary of Revisions** |
|  | 1 | P. Ackerman | 07.22.2015 | Initial Version |
|  | 2 | P. Ackerman | 08.27.2016 | Reformatted for CMS upload; changed logo |
|  | 3 | J. Laramie | 11.27.2017 | Changed MMQCI testing schedule: no testing with new lot/shipment. Rotation of panels monthly instead of weekly. |
|  | 4 | J. Laramie | 02.01.2017 | Added Extran 300 to the reagent list. Changed cleaning of hood in room 1 from bleach to 5% Extran. Changed cleaning of hood in room 3 from 70% ethanol to a water rinse twice. Added preparation of 5% extran. Biennial review: 01.12.2018 |
|  | 5 | J. Laramie | 08.17.2018 | Added notes about reagent labeling |
|  | 6 | J. Laramie | 02.17.20 | -Removed Simplexa testing from process control prep-Changed 5 day consecutive testing to day 1 and 5 (or day 2 for hMPV)-Changed hMPV prep from a patient sample pool to cultured material-Added nA target ranges for prep |