|  |
| --- |
|  MC 6.02 AST-XN08 Susceptibility Reporting Guidelines  |
| **Purpose** | This procedure provides instruction and guidance for routine testing and reporting of Antimicrobial Agents / organism groupings for the AST-XN08 Vitek card for patient testing. This procedure also provides instructions for a rolling validation for lowering breakpoints to be consistent with CLSI M100 edition 29 published in 2019. Organisms with intrinsic resistance to antimicrobial agents will not be reported.  |
| **Principal and Clinical Significance** | The decisions for the most appropriate antimicrobial agents to test and report are made with input from Pharmacy, Infectious Disease and the Clinical Laboratory. The goal is to provide clinically relevant information that will decrease the chance of developing antibiotic resistance, harmful effects of inappropriate antimicrobial use and avoid reporting results that could adversely affect patient care. |
| **Policy Statements** | This procedure applies to Microbiologists who perform culture set-up and plate reading |
| **Special Safety Precautions** | Microbiologists are subject to occupational risks associated with specimen handling.1. [*Biohazard Containment*](file:///%5C%5Ckidsnet.childrenshc.org%5Cchcdfs%5Cdept%5CLab%20Procedures%5CMicro%20Procedure%20Manuals%5CMC%20200%20%20%20%20Safety%5CMC%20201%20%20%20Biohazard%20Containment.doc)
2. [*Biohazardous Spills*](file:///%5C%5Ckidsnet.childrenshc.org%5Cchcdfs%5Cdept%5CLab%20Procedures%5CMicro%20Procedure%20Manuals%5CMC%20200%20%20%20%20Safety%5CMC%20204%20%20%20Biohazardous%20spills.doc)
3. [*Safety in the Microbiology Laboratory*](file:///%5C%5Ckidsnet.childrenshc.org%5Cchcdfs%5Cdept%5CLab%20Procedures%5CMicro%20Procedure%20Manuals%5CMC%20200%20%20%20%20Safety%5CMC%20202%20%20%20Safety%20in%20the%20Microbiology%20Lab%20Policy.doc)
 |
|  |  |  |  |  |
| **Procedure** | 1. The AST-XN08 card follows CLSI guidelines and breakpoints.
2. Perform and report KB or MicroScan results for antimicrobial agents that have card

limitations for patient testing following organism tables below. |
| **Lowering Breakpoints Validation** | 1. CLSI M100 edition 29 has updated 2019 breakpoints on Pseudomonas aeruginosa and *Acinetobacter* spp. for Meropenem. To validate the 2019 breakpoints, we will perform KB and enter results into Sunquest for patient testing. Vitek MIC results for Meropenem will be visible on the Vitek printout but will not cross into Sunquest. **Do not manually enter results**.
2. Using the results from the AST-XN08 card and the KB results, we will perform a rolling validation. Enter Vitek and KB results into the spreadsheet to collect data to validate the lower breakpoints.[**GN95 XN08 Rolling Validation.**](file:///G%3A%5CLAB%5CMicrobiology%5CMicro%20Validations%5CGN95%20rolling%20validation.xlsx)
3. If results are discrepant, report the KB results. For validation purposes, repeat both KB and XN08, freeze organism and notify supervisor.
 |
| **Antibiotics reported from card and Product Limitations** |

|  |  |
| --- | --- |
| Antibiotic | Product Limitations |
| Amikacin | Perform alternate method on *Acinetobacter baumannii* |
| Aztreonam | Perform alternate method on *Pseudomonas* spp. |
| Cefoxitin |  |
| Cefpodoxime | Perform alternate method on *Morganella morganii, Serratia spp.* |
| Ceftolozane/Tazobactam | Perform alternate method on *Morganella morganii, Providencia rettgeri, Serratia marcescens*Perform alternate method on resistant *Citrobacter koseri, Proteus mirabilis*,  *Proteus vulgaris and Serratia liquefacians* |
| Cefuroxime |  |
| Meropenem | Perform alternate method on resistant *Aeromonas spp* |
| Tetracycline |  |
| Tigecycline | Perform alternate method on resistant *Citrobacter freundii, Citrobacter koseri, Enterobacter cloacae, Enterobacter aerogenes, Escherichia coli, Klebsiella pneumoniae, Klebsiella oxytoca, and Serratia marcescens.*  |

Table 1For Non-Enterobacteriaceae, perform MicroScan if alternate method is required. There are no CLSI guidelines for KB with Non-Enterobacteriaceae. |
| **Validation for lowering breakpoints** |

|  |  |
| --- | --- |
| Antibiotic | Validation for lowering breakpoints- perform KB |
| Meropenem | Perform KB for validation of breakpoints on *Pseudomonas aeruginosa* and *Acinetobacter spp.* |

Table 2Enter results into the spreadsheet to collect data to validate the lower breakpoints.[**GN95 XN08 Rolling Validation.**](file:///G%3A%5CLAB%5CMicrobiology%5CMicro%20Validations%5CGN95%20rolling%20validation.xlsx) |
| **References** | Vitek AST-XN08 Gram Negative Susceptibility Card bioMerieux 2017/11CLSI M100 edition 29 Performance Standards for Antimicrobial Susceptibility Testing 1/2019 |
| **Training Plan/ Competency Assessment** | **Training Plan** | **Initial Competency Assessment** |
| 1. Employee must read the procedure.
2. Employee will observe trainer performing the procedure.
3. Employee will demonstrate the ability to perform procedure, record results and document corrective action after instruction by the trainer.
 | 1. Direct observation.
 |
|  |  |
| **Historical Record** |  |  |  |  |
|  | **Version** | **Written/Revised by:** | **Effective Date:** | **Summary of Revisions** |
| 1 | Susan DeMeyere | 4/14/2020 | Initial version |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |  |
|  |  |  |  |  |