| Analytical Balance Maintenance Program Policy |  |  |  |  |
| --- | --- | --- | --- | --- |
| Purpose | This policy provides instructions for ANALYTICAL BALANCE MAINTENANCE PROGRAM and provides guidance for the processes and procedures for use and care of analytical balances, scales, and external masses in use throughout Children’s Laboratory, performing the maintenance, and monitoring of the program. |  |  |  |
| Policy Statements | - This procedure applies to all laboratory employees - Balances must be kept clean and in good condition. - Balances will be serviced by qualified personnel according to the schedule that follows - Balances are to be installed where vibrations will not interfere with readings - Balances will be checked for accuracy at intervals appropriate to use and according to departmental procedures as indicated in the schedule that follows. Minimum verification frequency is annually for routine use, and biannually if used for weighing controls or calibrators. - New balances must be checked for accuracy before use. - Balances must be checked for accuracy whenever they are moved. - Records must be maintained showing results and acceptability of accuracy checks. - Analytical weight sets must be stored covered, and handled with gloves or acceptable devices such as padded tweezers - Analytical weight sets must be checked for accuracy annually by Biomed. |  |  |  |
| Historical Record | Version | Written/Revised by: | Effective Date: | Summary of Revisions |
|  | 1 | Linda Lichty | June 21, 2007 | Initial Version |
|  | 2 | D. Helfinstine | April 1, 2011 | New format, renumbered from CH 6.055 |
|  | 3 | Erin Bartos | April 12, 2017 | Add Ohaus Explorer Scale, remove asset 002662, replaced St Paul Transfustion Detecto with Ohaus CS2000 |
|  | 4 | Kelsi Brown | April 22, 2019 | Biennial Review. Updated Logo. |
|  | 4 | Matt Johnson | July 29, 2021 | No change. Biennial review. |
|  | 5 | J. Heimkes | 11/28/2022 | Changed document control number/location of policy; moved from Chemistry to Lab General. |

| Balance | Type | Location | Weighing Range | # Sensitivity | Service Frequency/ Qualification | Accuracy check frequency | Acceptable difference | Weight Class to certify | CHC Asset Tag # |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sartorius BL/500 | Type II | Minneapolis Histology | 1-1000 gms | 0.1 grams | Annual Biomed | Day of Use |  0.1 gm | ANSI/ASTMClass 3 | 008370 |
| Fisher XT | Type II | Minneapolis Morgue | 1 – 10,000 gms- | 0.01 grams | Annual Biomed National Scale | Day of Use |  1 gm | ANSI/ASTMClass 3 | 008371 |
| Sartorius Basic | Type II | Minneapolis Histology | 1 – 3000 gms | .000 grams | Annual Biomed National Scale | Day of Use |  0.1 gm | ANSI/ASTMClass 2 | Phillips M0336 |
| Sartorius B120S | Type II | Minneapolis Lead Room | .0001 – 120.0 gms | .0001 gms | Annual Biomed National Scale | Annual Biomed |  1% | ANSI/ASTMClass 1 | 006571 |
| Ohaus CS2000 | Type II | Minneapolis Sendouts | 0 – 2000 gms | 1 gm | Annual Biomed | Annual Biomed |  1 gm | ANSI/ASTMClass 3 | 012145 |
| Denver D1-8KD | Type II | St Paul Morgue | 1- 8000 gms | .01 gms | Annual Biomed National Scale | Day of Use |  1 gm | ANSI/ASTMClass 3 | 005090 |
| Sartorius TE 2101 | Type II | St. Paul Histology | 1 – 2100 gms | 0.1 gms | Annual Biomed | Day of Use |  1% | ANSI/ASTMClass 3 | 015105 |
| Sartorius PT 1200 | Type II | St. Paul Chemistry | 1 – 1200 gms | 0.1 gms | Annual Biomed | Annual Biomed |  1% | ANSI/ASTMClass 3 | 005082 |
| Ohaus Explorer | Type I | St. Paul Sweat Chloride | 0 – 120 grams | 0.0001 grams | Annual Biomed | Annual Biomed |  0.025% | ANSI/ASTMClass 1 | 0026998 |
| Ohaus Harvard Trip Balance | Triple Beam | Minneapolis Transfusion | 0-2000 gms | 1 gm | N/A | N/A |  |  | 008254 |
| Detecto Scale | Dietary Scale | Minneapolis Transfusion | 0-500 gms | 2 gm | Annual Biomed | Monthly BB Staff | Per Transfusion procedure |  | 008263 |
| Ohaus Harvard Trip Balance | Triple Beam | St. Paul Transfusion | 0-2000 gms | 1 gm | N/A | N/A |  |  | No asset tag |
| Ohaus CS2000 | Type II | St. Paul Transfusion | 0-2000 gms | 1 gm | Annual Biomed | Monthly BB Staff | Per Transfusion procedure | ANSI/ASTMClass 3 | 023130 |
| Weight Set | Class | Location | Range of Weights | Sensitivity | Qualification Frequency | Serial Number |  |  |  |
| Ohaus | Class 2 | Minneapolis Histology | 1 gm – 2000 gm | 1.001 gm | N/A | A58311 |  |  |  |
| Permaf Analytical | Class 1 | St. Paul Chemistry | 10 mg – 100 gms | .01 mg-.001 mg | Annual | A58312 |  |  | 012859 |
| Troemner | Class 1 | St. Paul Transfusion | 100 g – 1 Kg and 50 g | 0.1 mg | 2 years | 02200 02940 |  |  |  |
| Troemner | Class 1 | Minneapolis Transfusion | 100 g – 1 Kg and 50 g | 0.1 mg | 2 years | 02205 02939 |  |  |  |