|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **IH-COM User Management for IH Reader-24** | | | | | |
| **Purpose** | This procedure provides instructions to ensure that each operator of the IH-Reader 24 has a unique user ID and password with assigned user access. | | | | |
| **Policy Statements** | * The IH-Reader 24 software allows for assignment of access right at various levels * This procedure applies to any individual who will use the IH-Reader 24 * Each Operator assigns their own password. | | | | |
| **Procedure** |  | | | | |
|  | **Step** | Action | | | |
| Creating a User Account | 1 | From the IH-Com software, open ‘Configuration’ from the menu bar and the ‘User administration’. | | | |
|  | 2 | Select the ‘New User’ button. | | | |
|  | 3 | Enter the operator’s Unique Identifier (employee tech code) in the ‘User ID” field | | | |
|  | 4 | Enter the operator’s full name in the ‘Full name’ field. | | | |
|  | 5 | Enter the user’s level.   1. Levels in IH-Com:    1. 42-can change or confirm results, cannot accept results    2. 52-Normal User; can change and confirm results and validate    3. 62-Same as 5 above; can also do additional validation and configure QC    4. 90-Limited admin rights without database functions | | | |
|  | 6 | Select the ‘Set New Password Button to enter a password. | | | |
|  | 7 | Enter the password in the ‘New Password’ field and again in the ‘Repeat Password’ field | | | |
|  | 8 | Select the ‘set Password’ button to confirm | | | |
|  | 9 | Select ‘OK’, then ‘Save’ to except the user account | | | |
|  |  | | | | |
|  | **Step** | Action | | | |
| Changing a Password | 1 | In IH-Com, select ‘Configuration from the menu bar and then ‘User Administration’. | | | |
|  | 2 | Select the User ID that requires a password change. | | | |
|  | 3 | Select the ‘Set New Password’ button | | | |
|  | 4 | Enter the new password in the ‘New Password’ field and then again in the ‘Repeat Password’ field. | | | |
|  | 5 | Select the ‘Set Password’ button | | | |
|  | 6 | Select ‘OK’ then ‘Save | | | |
|  |  | | | | |
| **Limitations** | 1. Passwords must be a minimum of 8 characters, include a capital letter and a special character or number. 2. User ID is limited to 20 characters. 3. If a password is not defined when a user is added, the user will be prompted to create a password when logging in for the first time. 4. A user must have an access level of 62 or higher in order to create user accounts. 5. Authorization levels cannot exceed that of the logged in user. A user level of 99 has full administrator access to all menus in IH-Com. | | | | |
| **References** | 1. IH-Com User Manual NA V1.2-02/2021 Chapter D, Sections 1.1 to 1.7 | | | | |
| **Approval**  **Workflow** | Transfusion Service/Lab Director | | | | |
|  |  | | | | |
| **Historical Record** | **Version** | | **Written/Revised by:** | **Effective Date:** | **Summary of Revisions** |
| 1 | | S. Cassidy | 02/17/2023 | Initial Version |