| Inspection and Labeling of Gel Cards |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for the inspection and labeling of all BioRad IH gel cards |  |  |  |  |
| Policy Statements | - Store at 18 to 25°C - Do not use beyond expiry on the label which is expressed YYYY-MM-DD - Store in an upright position - Do not freeze or expose cards to excessive heat. - Do not store near any hear, air-conditioning sources or ventilation outlets - Visually inspect gel cards before each use. - Label cards with indelible marker or ink. Sunquest aliquot labels, donor unit labels, and IH-Com labels are acceptable for labeling cards. - All IH-System reagents and test samples must be brought to room temperature (18-25°C) prior to use |  |  |  |  |
| Materials | - IH gel card - Indelible pen or marker - Patient Sunquest aliquot label - Donor unit label - IH-Com label |  |  |  |  |
| Procedure |  |  |  |  |  |
|  | Step | Action |  |  |  |
|  | 1 | Visually inspect the gel card before each use. - Do not use cards showing signs of drying, discoloration, bubbles, crystals or other artifacts. - Do not use cards with damaged foil strips. - Do not use gel cards if the gel matrix is absent or if the liquid level in the microtube is not at or below the gel matrix. A clear liquid layer should be visible on top of the uniform gel matrix in each microtube. - Cards with dispersed drops observed at the top of the microtube, due to improper storage or shipping conditions, have to be centrifuged with the IH-Centrifuge L with preset time and speed before use. If drops are still observed on top of the microtube after one centrifugation it is recommended to not use the card. |  |  |  |
|  | 2 | Label the gel cards for the testing to be performed> - Patient Identification labeling options: - IH-Com label - Sunquest allocate label - Patient’s first and last initial - Donor Unit identification labeling options: - IH-Com Label - Donor unit label - Last 3 digits of donor number - Testing identification - Antibody Screening cells: S1 and S2 or SI and SII - Direct antiglobulin testing:DAT - Weak D testing: WD |  |  |  |
| References | - Package insert IH-Card AHG Anti-IgG, Current version, Bio-Rad Medical Diagnostics GmgH - Package insert IH-Card AHG Anti-IgG,C3d, Current version, Bio-Rad Medical Diagnostics GmgH - Package insert IH-ABO/RH cards, Current version, Bio-Rad Medical Diagnostics GmgH |  |  |  |  |
| Appendices | Appendix A: Example of prober card labeling Appendix B: Example of gel card failures |  |  |  |  |
| Approval Workflow | Transfusion Service/Lab Director |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: | Effective Date: | Summary of Revisions |
|  | 1 |  | J Wenzel | 02/08/2008 | Initial Version |
|  | 2 |  | J Wenzel | 4/10/2012 | CMS format |
|  | 3 |  | S. Cassidy | 02/17/2023 | Updated for new reagents and format from a PDF |

Appendix A: Example of properly labeled gel card

Appendix B: Examples gel card failures