| IH-500 Maintenance |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for performing maintenance on the IH-500 to ensure that is properly maintained. |  |  |  |  |  |  |
| Policy Statements | - Perform the maintenance work according to the instructions. Carefully read the steps described before carrying out any operation. - When any resource is removed, on-board time management is no longer traced by the system. - IH-500 unloads all reagent vials and sample racks. If reagent vials are reloaded, check their integrity before reusing them. - A Quality Control Check (QC) must be performed after any maintenance operation. |  |  |  |  |  |  |
| Related Documents | TSf 17.43.1 IH-500 Maintenance Form |  |  |  |  |  |  |
| Materials | Equipment |  |  | Reagents |  |  | Supplies |
|  | - IH-500 |  |  | - Microcide SQ - DI Water |  |  | - 70% Ethanol |
| Procedure |  |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  |  |
| Daily | 1 | Check and clean the outer surfaces of the instrument with a suitable disinfectant, if needed. |  |  |  |  |  |
|  | 2 | Check and clean the outer surfaces of the instrument, the hand-held barcode scanned, and touch screen with a link-free cloth soaked in 70% ethanol. |  |  |  |  |  |
|  | 3 | Complete form TSf 17.43.1 IH-500 Maintenance Form |  |  |  |  |  |
|  | Step | Action |  |  |  |  |  |
| Weekly | 1 | When the weekly maintenance alarm appears on the main screen, launch the maintenance procedure as soon as possible. No testing can be performed until maintenance is completed. |  |  |  |  |  |
|  | 2 | Login with required user rights. |  |  |  |  |  |
|  | 3 | Select MENU (header strip). |  |  |  |  |  |
|  | 4 | Select “Stop all” to stop all processes running on the instrument. - All tests in progress are canceled. - Open the solid waste area door. - Remove the cleaning liquid container from its storage area. - Unscrew the cap and remove the black cap assembly. - Empty any remaining solution present following local waste procedures. - Rinse container. - Prepare cleaning liquid solution (if necessary before beginning). Cleaning liquid solution is 2 liters of deionized water and 8 mL of Microcide SQ. This solution is stable for 7 days. - Tighten the cap on the container. 1.Solid waste door 2.Cleaning liquid container 3.Black cap 4.Tubing |  |  |  |  |  |
|  | 5 | Select MAINTENANCE - A maintenance screen is displayed - Verify that there is no leakage around the liquid containers. - The left pane displays the estimated time to complete the weekly maintenance procedure. - Follow the on-screen instructions as follows below: - Open the solid waste area (5) and containers door (1) - Disconnect and remove the right system liquid container (4) with the blue cap. - Empty both LIQUID WASTE containers, rinse with cleaning liquid and reconnect them. - Select next. - Connect the CLEANING LIQUID container in position #1 of the system liquid container. - Select NEXT when done, the screen indicates Step 1 is completed. - Follow onscreen instructions to disconnect the cleaning liquid container and connect it in position #2 on the system liquid - Select NEXT when done. The screen indicated Step 2 is completed. - Follow ON screen instructions: - Disconnect the CLEANING LIQUID container and move to the storing area (5). - Ensure both SYSTEM LIQUID container (4) are full. - Connect both SYSTEM LIQUID containers. - Select NEXT when done. The screen indicates Step 3 is completed. - Select FINISH to close the Screen. - Weekly Maintenance has been completed. - Select MENU - Select INITIALIZE INSTRURMENT - Perform QC |  |  |  |  |  |
|  | 6 | Open up the Maintenance Report folder on the IH-Com desktop and drag a copy of the completed maintenance report into the Weekly Maintenance Reports folder also on the desktop. |  |  |  |  |  |
| As Needed Maintenance | 1 | Login with the required user rights |  |  |  |  |  |
| Needle Cleaning/Change | 2 | Select MENU |  |  |  |  |  |
|  | 3 | Select STOP ALL, then YES to confirm. All tests in progress are canceled. |  |  |  |  |  |
|  | 4 | Select CHANGE PIPETTE NEEDLE as the answer next to “Do you want to change the needle?’ |  |  |  |  |  |
|  | 5 | Select Next to confirm change of the needle. |  |  |  |  |  |
|  | 6 | Select Next. The pipetting area door unlocks |  |  |  |  |  |
|  | 7 | Follow on-screen instructions: - open the pipetting area door (1); - check sample racks are removed or in the locked position prior to open-ing. - place a tissue under the pipette needle tip before disconnecting the tub-ing as it drips; - disconnect the pipette needle (3) from the tubing by unscrewing the red connector (4); - open both locks (5) by turning counter clockwise; - remove the pipette needle from the housing (6); - remove the sampling needle from its packaging; - position the new pipette needle (3) in its housing (2). - tighten the tubing connector (red) (4) on the pipette needle (3). Turn it gently clockwise till you hear a click; - lock the needle by turning the two locks clockwise (5); With the locks closed, the pipette needle is connected and locked in position on its pipettor Z axis support. - clean the pipette needle reference position with a lint-free cloth soaked in 70% ethanol; |  |  |  |  |  |
|  | 8 | Select Next. |  |  |  |  |  |
|  | 9 | Select Next to validate the needle replacement |  |  |  |  |  |
|  | 10 | Select Finish to close the screen. |  |  |  |  |  |
|  | 11 | Select MENU. |  |  |  |  |  |
|  | 12 | Select INITIALIZE INSTRUMENT. IH-500 unloads all reagent vials and sample racks and the system initializes. |  |  |  |  |  |
|  | 13 | Perform a QC run after a change pipette needle. |  |  |  |  |  |
| References | - IH-500 User Manual NA V1.0-07/2017, Chapter 10 |  |  |  |  |  |  |
| Approval Workflow | Transfusion Service/Lab Director |  |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: |  | Effective Date: | Summary of Revisions |  |
|  | 1 |  | S. Cassidy/J.Hudgens |  |  | Initial Version |  |