## Slide 1: Clinical Laboratory Hematology
Hemostasis: Laboratory Testing & Instrumentation
Copyright © 2020, 2015, 2012 Pearson Education, Inc. All Rights Reserved

## Slide 2: Specimen Collection and Processing

## Slide 3: Specimen Collection (1 of 8)
Accuracy of testing
  Relies on proper
    Collecting of specimen
    Processing of specimen
    Storing of specimen

## Slide 4: Specimen Collection (2 of 8)
Blood collection for coagulation testing
  Two-syringe or two-tube technique
    Blood in first syringe or tube was discarded to minimize contamination with T F
    Current guidelines no longer require discard tube for P T and A P T T testing
  Coagulation testing tubes should not be drawn after tubes with additives (e.g., heparin or E D T A)
  Direct venipuncture is preferred

## Slide 5: Specimen Collection (3 of 8)
Blood collection for coagulation testing
  Butterfly needles of 21-22 gauge with polystyrene tubing attached to syringe or vacutainer system for patients with difficult venous access
    Two-tube technique used to eliminate problem of under filling tube due to air in tubing (proper 9:1 blood to anticoagulant ratio)

## Slide 6: Specimen Collection (4 of 8)
Blood collection for coagulation testing
  Indwelling catheter
    Must avoid heparin contamination
    Catheter line should be flushed with saline
    Adults-first 5 milliLitre of blood should be discarded
    Heparin contamination
      Commercial heparinase can be added to the plasma to remove heparin

## Slide 7: Specimen Collection (5 of 8)
Anticoagulant of choice
  3.2% buffered sodium citrate
Ratio of whole blood:anticoagulant-9:1
  >55% hematocrit
    Smaller volume of plasma relative to citrate
    Excess free citrate binds calcium in the test procedure
    Falsely prolonged clotting times
    Adjust citrate concentration
    Citrate tubes are available that draw 4.5 milliLitre, 2.7 milliLitre, 1.8 milliLitre

## Slide 8: Specimen Collection (6 of 8)
Anticoagulant to blood ratio
  9:1 ratio critical for valid results
    If underfilled-↑ citrate, binds calcium in test procedure
      Falsely prolonged test results
    If overfilled-insufficient calcium bound
      Clotting can occur in tube
      Falsely prolonged results

## Slide 9: Figure 36-1
C = Volume of sodium citrate
V = Volume of whole blood drawn
H c t = Patient hematocrit
For example, to draw 5 milliLitre of blood from a patient with a hematocrit of 63% you calculate the amount of sodium citrate to use as follows:
So mix 5 milliLitre of blood with 0.34 milliLitre of sodium citrate
Notes: A correction formula to adjust the citrate/blood ratio is recommended when the hematocrit is greater than 55%. When blood is drawn for coagulation testing in a patient with a hematocrit that is more than 55%, the reduced plasma volume (exceeds the 1:9 ratio of buffered sodium citrate to blood) can affect coagulation testing.

## Slide 10: Specimen Collection (7 of 8)
Accurate labeling of tube is critical
Include
  Appropriate identifiers according to institutional and regulatory agencies guidelines in patient’s presence
  Additional info for further identification
    For example: time of draw, before or after transfusion

## Slide 11: Specimen Collection (8 of 8)
Check patient history of receiving blood products and other medications/supplements
  If testing done within the half-life of administered clotting factor then test may measure transfused component as well as the patient’s component
  Dietary supplements and drugs can affect hemostasis
  Fibrinolytic factors-diurnal variability
  Platelet studies-medication history

## Slide 12: Specimen Processing (1 of 6)
Plasma for coagulation testing
  Routine hemostasis testing
    Performed on citrated platelet-poor plasma (P P P)
  Platelet function testing
    Performed on platelet-rich plasma (P R P) and whole blood

## Slide 13: Specimen Processing (2 of 6)
P P P
  Citrated specimen centrifuged for 15 minutes at 1500-2500g
  Depending on coagulation instrumentation
    Leave plasma on top of packed cells or
    Remove plasma with plastic pipette and place in plastic tube with cap
      Leave small amount of plasma over packed cell layer-no platelet contamination

## Slide 14: Specimen Processing (3 of 6)
Use platelet-poor plasma because platelets contain
  P F 4
    Neutralizes heparin
  Phospholipids
    Affect L A testing and factor assay testing
  Proteases
    Alter results for von Willebrand testing

## Slide 15: Specimen Processing (4 of 6)
Any clot in specimen
  Specimen unacceptable
  After removing P P P, twirl applicator stick in packed cells to detect clots
Fresh platelet-poor plasma best
Separated plasma
  Stored at 18-24°Celsius or 2-8°Celsius for up to 4 hours
  If testing is delayed >4 hours
    Store at −70° Celsius for up to 6 months

## Slide 16: Specimen Processing (5 of 6)
Platelet-poor plasma (P P P)
  Frozen specimens
    Thawed rapidly at 37°Celsius
    Excessive heating can destroy F V and F V I I I
    Never store in self-defrost freezers
  Special testing may require special collection and storage procedures
Platelet-rich plasma (P R P)
  Needed for platelet function studies
  Centrifuge at low spins to yield platelet concentration of

## Slide 17: Specimen Processing (6 of 6)
Citrated whole blood
  Platelet function testing (using P F A)
  Whole blood platelet aggregation
  Global tests of coagulation (T E G, T E M)
  Check for clots
    Platelets can become activated by foreign surfaces
      Wooden applicator stick inserted after testing to detect clots
  All whole blood specimens should be allowed to equilibrate at room temperature for 30 minutes before testing

## Slide 18: Table 36-1. Potential Sources of Error in Coagulation Testing
P P P, platelet poor plasma
Specimen | Reagents | Procedure
Inappropriate ratio of anticoagulant to blood | Incorrect preparation of reagents | Incorrect temperature
Failure to correct citrate volume if hematocrit is > 55% | Failure to properly store reagents | Incorrect incubation times
Clotted, hemolyzed, icteric, or lipemic specimen | Use of reagents beyond reconstituted stability time | Incorrect volumes of sample, reagents, or both
Poor aspiration of plasma off of cells so that sample is not P P P | Use of reagents beyond expiration date | Improperly functioning instrument
Delay in processing or testing | Contaminated reagents | Blank
Inappropriate storage | Blank | Blank

## Slide 19: Laboratory Investigation of Primary Hemostasis

## Slide 20: Laboratory Investigation of Primary Hemostasis (1 of 2)
Evaluation of primary hemostasis includes
  Complete blood count
  Blood smear evaluation
    For platelet morphology
  Platelet function

## Slide 21: Laboratory Investigation of Primary Hemostasis (2 of 2)
Platelet function assays
  Specific drugs can affect platelet function for the life span of platelet (7-10 days)
  Screening tests
    In vivo testing using the B T
      Considered obsolete by most investigators
    In vitro function testing using the P F A
  Definitive/specialized tests
    Platelet aggregation using P R P and/or whole blood
    Flow cytometry

## Slide 22: Bleeding Time (B T)
In vivo measurement of primary hemostasis
  Affected by Plt number, function, and vascular integrity
Measures time required for bleeding to cease from a standardized skin incision
No longer recommended by C L S I

## Slide 23: Platelet Function Analyzer (P F A) (1 of 3)
In vitro method
  Evaluate platelet function
  Standardized technique
  Eliminates variables associated with B T

## Slide 24: Platelet Function Analyzer (P F A) (2 of 3)
P F A-100®, 200®; Siemens Healthcare, Inc.
  Platelet function in whole blood
    Measured at a high shear rate
  Instrument aspirates citrated blood to three reservoirs
    Collagen/epinephrine
    Collagen/adenosine diphosphate (A D P)
    Low concentration of A D P and prostaglandin E1 (P G E1)
      Assessing P2Y12 inhibitors
  Pressure sensor detects formation of platelet plug

## Slide 25: Platelet Function Analyzer (P F A) (3 of 3)
P F A-100®, 200®; Siemens Healthcare, Inc.
  Closure time
    Time to occlude the aperture
    Function of platelet count, platelet activity, V W F activity, and hematocrit
  Sensitive to
    V W D, aspirin-induced platelet dysfunction, clopidogrel interference, and aggregation defects

## Slide 26: Table 36-2 Expected Results with the P F A
Condition or Therapy | Substance Used to Coat Aperature Collagen/Epinephrine | Substance Used to Coat Aperature Collagen/A D P | Substance Used to Coat Aperature P G E1/A D P
Normal | Normal | Normal | Normal
Acetylsalicylic acid (A S A) | Abnormal | Normal | Normal
V W D | Abnormal | Abnormal | Unknown
Glanzmann’s thrombasthenia | Abnormal | Abnormal | Abnormal
Clopidogrel | Variable; not recommended | Variable; not recommended | Abnormal

## Slide 27: P O C Platelet Aggregation
Verify Now Accumetrics Assay System
  Turbidimetric optical detection
    Measures P l t aggregation by light transmission aggregometry (L T A)
    3.2% sodium citrated collection tube
  Blood is dispensed into designated cartridges
    Microbead agglutination technology
    Cartridges contain fibrinogen-coated beads with
      A A-aspirin testing
      A D P and P G E1-clopidogrel testing
      Thrombin receptor-integrilin testing

## Slide 28: Platelet Aggregometry
Comprehensive studies of platelet function
Specimen
  3.2% sodium citrated blood
  Can use whole blood or P R P
Pre-examination variables
  Recent (within 7-10 days) medications
  Food, exercise
  Tourniquet use
  Needle size
  Temperature, equilibration time
  Centrifugation processing

## Slide 29: P R P Light Transmission Aggregometry (1 of 6)
P R P
  Carefully prepared from a citrated blood sample
  Adjust platelet count with patient’s own P P P to a standard number
    Usually 200,000/microLiter
    Some investigators do not recommend adjustment
      Can cause spontaneous aggregation
    >600,000/μ L requires adjustment with the patient’s own P P P
  Sample stirred, warmed to 37°Celcius in an aggregometer

## Slide 30: P R P Light Transmission Aggregometry (2 of 6)
P R P
  Aggregating reagent (agonist) is added
  Platelets aggregate
    Change in O D of the P R P
    Recorded as a graph (curve)
  Depending on the agonist used and its concentration
    Primary and/or secondary wave of aggregation
  Common platelet agonists
    A D P, epinephrine, collagen, ristocetin, arachidonic acid

## Slide 31: P R P Light Transmission Aggregometry (3 of 6)
Primary wave
  Direct response of the platelets to the aggregating reagent
  Represents
    Platelet shape change
    Formation of small aggregates
Secondary wave
  Complete aggregation response
  Result of endogenous A D P being released from the activated platelet dense bodies

## Slide 32: P R P Light Transmission Aggregometry (4 of 6)
Biphasic curve
  Aggregating reagent that results in both primary and secondary wave
Monophasic curve
  Aggregating reagent that results in one wave
Changes in platelet aggregation curves
  Identify qualitative platelet disorders

## Slide 33: P R P Light Transmission Aggregometry (5 of 6)
Ristocetin
  Action depends on
    Interaction of plasma V W F and
    Platelet membrane glycoprotein I b/I X/V
  Represents platelet agglutination
  R I P A abnormal
    V W F
      Corrected by adding normal plasma to patient’s P R P (plasma defect)
    B S S
      Not corrected by adding normal plasma (platelet defect)

## Slide 34: P R P Light Transmission Aggregometry (6 of 6)
Arachidonic acid (A A)
  Screening for patient’s ingestion of aspirin (or aspirin containing products) and N S A I Ds within past 7-10 days
  No secondary wave of aggregation with A D P or epinephrine
Should not be used for assessing patients at risk for thrombosis

## Slide 35: Figure 36-2 Normal Platelet Aggregation Curves
Notes: Normal responses to the commonly used aggregating reagents: ADP, epinephrine, collagen, ristocetin. (a) Baseline before addition of the reagents. (b) Initial increase in absorbance that occurs immediately following the addition of reagent, representing a change in platelet shape. (c) The primary wave of aggregation. (d) Secondary wave of aggregation.

## Slide 36: Table 36-3 Typical Pattern of Response to Aggregating Reagents in Qualitative Platelet Disorders a (1 of 2)
Qualitative Platelet Disorders | Aggregating Reagents Collagen | Aggregating Reagents A D P | Aggregating Reagents Epinephrine | Aggregating Reagents Ristocetin | Aggregating Reagents Arachidonic Acid
Normal | Monophasic curve (secondary wave only): representing secondary platelet aggregation associated with release of endogenous A D P | Biphasic curve with 2 times 10 to the power negative 5 M A D P: demonstrating primary and secondary waves of platelet aggregation | Biphasic curve: demonstrating primary and secondary waves of platelet aggregation | Biphasic curve: demonstrating primary wave resulting from immediate platelet agglutination; secondary wave associated with platelet release induced by platelet agglutination | Monophasic curve (secondary wave only): representing effect of thromboxane A2 on secondary platelet aggregation
von Willebrand disease | Normal | Normal | Normal | No response | b
Bernard-Soulier syndrome | Normal | Normal | Normal | No response | b

## Slide 37: Table 36-3 Typical Pattern of Response to Aggregating Reagents in Qualitative Platelet Disorders a (2 of 2)
a See Chapter 33
b Not applicable. Arachidonic acid evaluates the thromboxane A2 synthetic path-way.
Qualitative Platelet Disorders | Aggregating Reagents Collagen | Aggregating Reagents A D P | Aggregating Reagents Epinephrine | Aggregating Reagents Ristocetin | Aggregating Reagents Arachidonic Acid
Glanzmann’s thrombasthenia | No response | No response | No response | Normal | b
Storage pool disease | No response | Primary wave only | Primary wave only | Normal | Normal
Aspirin ingestion | No response | Primary wave only | Primary wave only | Normal | Suppressed response
Clopidogrel ingestion | Normal | No response | Primary wave only | Normal | Normal

## Slide 38: Figure 36-3 Platelet Aggregation Patterns in Various Disorders
Notes: Typical tracings obtained using concentrations of collagen, ADP, and ristocetin illustrate the differences between platelet aggregation in normal subjects and that in patients with VWD, BSS, TSA, SPD, and aspirin ingestion and ASAs. Impaired release of ADP accounts for the SPD and ASA defects. VWF corrects the defective ristocetin aggregation in VWD but not in BSS. ADP, adenosine diphosphate; VWD, von Willebrand disease; BSS, Bernard-Soulier syndrome; TSA, thrombasthenia; SPD, storage pool disease; ASA, aspirine-like disorders; VWF, von Willebrand factor.

## Slide 39: Whole Blood Aggregation
Allows for a smaller sample size
Assay
  Measures electrical resistance across two metal wires (probes)
  Whole blood sample exposed to small electric current-coats wire with monolayer of platelets
  Upon the addition of an agonist
    Platelets form aggregates on the probes
    Change in impedance measured as function of time
    Citrated sample challenged with ristocetin, A D P, collagen, A A, thrombin

## Slide 40: Platelet Aggregometer (1 of 2)
Multiplate® platelet function analyzer
  Detects effects of antiplatelet inhibitors (aspirin, clopidogrel, and G P I I b/I I I a)
  Uses a disposable cuvette with two independent sensors
  Assesses three parameters
    Velocity of aggregation
    Maximum aggregation
    Area under curve (A U C)

## Slide 41: Platelet Aggregometer (2 of 2)
Multiplate® platelet function analyzer
  Calculated results expressed as mean value of parameters determined
  Abnormal results
    Could indicate that antiplatelet treatment needs to be modified

## Slide 42: Additional Tests Evaluating Platelets
Flow cytometry
  Monoclonal antibodies directed against glycoproteins on platelets
    G P I b/I X (B S S)
    G P I I b/I I I a (Glanzmann’s thrombasthenia)
  Presence and function can be evaluated
  Vasodilator-stimulated phosphoprotein (V A S P) phosphorylation
    Clopidigrel resistance
Clot retraction (rarely performed)
  Function of activated platelets

## Slide 43: Laboratory Investigation of Secondary Hemostasis

## Slide 44: Table 36-4 Laboratory Tests in Preterm and Term Infants as Compared with Those of Adults and Older Children
P T, prothrombin time; A P T T, activated partial hromboplastin time; T T, thrombin time; F D P, fibrin degradation products
Test | Preterm Infant (28-31 weeks) (Day 1) | Preterm Infant (32-36 weeks) (Day 1) | Term Infant (37-41 weeks) (Day 1) | Adults and Older Children | Age Adult Level Reached
Platelet count times 10 cubed per microlitre | 150-430 | 150-430 | 174-456 | 150-450 | Before birth
Platelet aggregation | Abnormal | Abnormal | Abnormal | Normal | One month
P T, sec | 14.6-16.9 | 10.6-16.2 | 10.1-15.9 | 10.8-13.9 | Comparable at birth
A P T T, sec | 80-168 | 27.5-79.4 | 31.3-54.5 | 26.6-40.3 | By 6 months
T T, sec | Blank | 19.0-30.4 | 23.5 ± 2.38 | 19.7-30.3 | At birth
F D P | Blank | Normal | Normal | Normal | Before birth

## Slide 45: Introduction-Laboratory Investigation of Secondary Hemostasis
Evaluate
  Coagulation factors
  Coagulation inhibitors
Starts with screening tests
  Prothrombin time (P T)
  Activated partial thromboplastin time (A P T T)
  Thrombin time (T T)
  Quantitative fibrinogen

## Slide 46: Prothrombin Time (P T) (1 of 4)
Screening test
  For inherited and acquired deficiencies in extrinsic pathway (extrinsic X-ase pathway)
    Translocation of tissue factor (T F) to the cell surface activates the coagulation cascade by formation of T F/F V I I complex

## Slide 47: Prothrombin Time (P T) (2 of 4)
Procedure
  Thromboplastin is added to a citrated patient P P P
  Record time for fibrin formation (clotting time)
Thromboplastin (rabbit brain and recombinant human tissue factor)
  Vary in sensitivity due to differences in thromboplastin source, type, and concentration of phospholipids
  Must establish reference interval for reagent/instrumentation combination

## Slide 48: Prothrombin Time (P T) (3 of 4)
Clot detection
  Optical or electromechanical methods
  Manual, automated, P O C devices
General reference interval
  12-15 seconds
Prolonged P T
  Deficiency in factors
  fibrinogen, or presence of
  inhibitor, heparinoids, D O A Cs
  Proteins induced by vitamin K absence or antagonism (P I V K A)
    lack γ-carboxyglutamic acid residues for binding C a++ and P L

## Slide 49: Prothrombin Time (P T) (4 of 4)
Historically used to monitor oral anticoagulant therapy
  Coumadin or warfarin
  Results are now reported using the international normalized ratio (I N R)

## Slide 50: Activated Partial Thromboplastin Time (A P T T) (1 of 3)
Screen
  Inherited or acquired deficiencies of proteins in the intrinsic pathway (intrinsic X-ase pathway)
  Circulating inhibitors
Sensitivity of commercial reagents
  Varies greatly
Screening tests (P T and A P T T)
  Detect factor deficiency only when factor decreases to a level of 25-40% of normal
Monitor unfractionated heparin therapy

## Slide 51: Activated Partial Thromboplastin Time (A P T T) (2 of 3)
Procedure
  Partial thromboplastin
    Provides phospholipid surfaces
      Simulates activated platelet surfaces
    Activator (kaolin, celite, micronized silica, ellagic acid)
      Negatively charged surface for activation of F X I I and P K

## Slide 52: Activated Partial Thromboplastin Time (A P T T) (3 of 3)
Procedure
  Citrated plasma and A P T T reagent incubated together at 37°Celsius
    Optimal activation of contact factors
  Calcium chloride is added
  Time to fibrin formation is recorded
Clot detection
  Optical or electromechanical methods
  Manual or automated devices

## Slide 53: Thrombin Time (T T) (1 of 2)
Measures conversion of fibrinogen to fibrin when thrombin is added to undiluted citrated P P P
Procedure
  Incubate P P P at 37°Celsius and add thrombin
  Time to fibrin formation is recorded,
General reference interval 15-19 seconds

## Slide 54: Thrombin Time (T T) (2 of 2)
Interference in conversion of fibrinogen to fibrin
  Hypofibrinogenemia or dysfibrinogenemia
  Presence of heparin, D T I s, or F D Ps
  Rare cases
    Antibodies against bovine thrombin, paraproteins (myelomas)
Extremely long T T
  Usually indicates drug effect (heparin or D T I)
  If sample contains heparin, neutralize with
    Hepzyme, protamine sulfate

## Slide 55: Quantitative Fibrinogen Assay (1 of 2)
Reference method
  Clauss assay
    Clot-based functional measurement
    Adds thrombin to various dilutions of known concentrations of fibrinogen
      Clotting times are plotted on a log/log graph
        x-axis-known concentrations
        y-axis-clotting times
    Fibrinogen results for the controls and patient
      Determined from the reference curve using their respective C Ts

## Slide 56: Quantitative Fibrinogen Assay (2 of 2)
Reference method
  Fibrinogen concentration inversely proportional to the clotting time
  Reference intervals-200 to 400milligrams/deciLitres
↓ fibrinogen
  D I C, primary and secondary fibrinolysis, liver disease, congenital disorders
↑ fibrinogen
  Inflammatory disorders, pregnancy, women on oral contraceptives

## Slide 57: Figure 36-4 Fibrinogen Standard Curve
Notes: I, II, and III represent reference plasmas of known fibrinogen concentrations. The TT in seconds is on the y-axis, and the fibrinogen concentration in mg/dL is on the x-axis. TT, thrombin time.

## Slide 58: Tests to Evaluate Specific Factor Deficiency
Further testing performed if
  P T and/or A P T T is prolonged
Mixing studies
Specific coagulation factors
Reptilase time
Prekallikrein screening test
Factor
screening test
V W F tests

## Slide 59: Mixing Studies (1 of 2)
Also known as
  Circulating anticoagulant screen, or
  Screening test for circulating inhibitor
Performed to differentiate
  Factor deficiency from presence of circulating inhibitor
  Repeat abnormal screening tests (P T, A P T T)
    Using several different dilutions of patient’s P P P mixed with normal pooled plasma

## Slide 60: Mixing Studies (2 of 2)
Results of mixing study
  Patient’s plasma corrected after incubation with normal plasma
    Factor deficiency
    Normal plasma replenishes the deficient factor in patient’s plasma
    Next step perform factor assays
  Patient’s plasma not corrected after incubation with normal plasma
    Presence of circulating or specific factor inhibitor
    Next step perform tests to verify type of inhibitor

## Slide 61: Table 36-5 Differentiation of Factor Deficiency from Circulating Inhibitors Using the Mixing Study Procedure
P T, prothrombin time; A P T T, activated partial thromboplastin time.
Deficiency or Inhibitor | Immediate P T or A P T T after Mixing | Mix, Perform P T or A P T T after 2-hr incubation at 37°Celsius
Factor deficiency | Correction | Correction
Lupus anticoagulant or heparin | No correction | No correction
F V I I I inhibitor | Correction | No correction
F V inhibitor | No correction | No correction

## Slide 62: Specific Coagulation Factors (1 of 5)
Performed to:
  Confirm specific factor deficiency
  Determine actual activity of factor
Principle
  Ability of patient’s plasma to correct a prolonged P T or A P T T of a known factor-deficient plasma (substrate)
One-stage assays based on the P T
  Extrinsic (F V I I) and common pathway (F I I, F V, F X)
One-stage assays based on the A P T T
  Intrinsic pathway (F V I I I,F I X, F X I, F X I I)

## Slide 63: Specific Coagulation Factors (2 of 5)
Procedure for factor assays
  Standard curve constructed from
    Clotting times of factor-deficient substrate plasma containing varying dilutions of a reference pooled plasma
  Patient’s clotting time
    1:10 and 1:20 and/or 1:40 dilutions of patient plasma with specific factor-deficient plasma
    Perform P T or A P T T on mixture
    Times are converted to % activity from standard curve
    Results should be linear
      If not-suggests inhibitors

## Slide 64: Specific Coagulation Factors (3 of 5)
Procedure for factor assays
  Normal factor activity reference interval
    ~50-150%
Commercial chromogenic kits
  F V I I I assay
    F V I I I is activated by thrombin
    F V I I I a acts cofactor in conversion of F X to F X a by F I X a in presence of C a++ and P L
    Amount of F X a generated is directly proportional to amount of F V I I I in sample

## Slide 65: Table 36-6 Dilutions for Factor Assays
a The example is based on a reference plasma value of 95.0%. When reference plasma with an assayed value other than 100% is used, the activity would be based on its stated value.
b Additional dilutions are prepared when patient values are less than 1% activity. Clotting times are determined for each dilution and then plotted on the standard curve. The percent activity is plotted on the x-axis and the clotting time on the y-axis (Figure 36-5)
Tube | Buffer (milliLitre) | Plasma (milliLitre) | Activity a (%) | Dilution b
1 | 0.9 | 0.1 | 95.0 | 1:10
2 | 1.9 | 0.1 | 48.0 | 1:20
3 | 0.5 | 0.5 milliLitre from tube 2 | 24.0 | 1:40
4 | 0.5 | 0.5 milliLitre from tube 3 | 12.0 | 1:80
5 | 0.5 | 0.5 milliLitre from tube 4 | 6.0 | 1:160

## Slide 66: Figure 36-5 Factor Activity Curve Using an F V I I I Reference Plasma Activity of 95.0%
Notes: The factor activity curve is prepared by plotting the CT in seconds for each reference plasma dilution on the y-axis and the percent factor activity for each dilution on the x-axis. The CT of at least two patient dilutions is measured, and the percent of activity is interpolated from the factor activity curve. The graph includes the r value and the slope information along with the dilution, CT, and measured values. M/D: M is to record when the operator performs a manual rerun of the calibration point and D is a record for a point deletion if necessary. CT, clotting time. Modified from the STA-R Evolution coagulation analyzer.

## Slide 67: Specific Coagulation Factors (4 of 5)
Commercial chromogenic kits
  Chromogenic F X
    Uses R V V to activate F X to F X a
    F X a hydrolyses S-2765 and liberates chromophore p N A
    Color intensity is proportional to F X concentration
    Useful for managing patients with L A receiving warfarin and patients being treated with D T I s

## Slide 68: Specific Coagulation Factors (5 of 5)
Commercial chromogenic kits
  Less commonly used due to high cost
  Not affected by interfering substances that affect clot-based assays
    Chromogenic factor
    has been shown to be
    falsely low in the presence of D O A Cs
  P C, P S, A T, and P A I-1 assays available and can be performed with same instruments

## Slide 69: Reptilase Time (1 of 2)
Used with T T to detect heparin
  Not affected by heparin
  In the presence of heparin
    R T is normal, T T is increased
Differentiates dysfibrinogenemia from presence of F D Ps and paraproteins
Reptilase is a serine protease
  Found in venom of Bothrops atrox snake
  Cleaves fibrinopeptide A from fibrinogen
    Thrombin cleaves both fibrinopeptide A and B

## Slide 70: Reptilase Time (2 of 2)
Addition of reptilase to P P P
  Initiates clot formation
  Detected by optical or electromechanical methods
  Manual, semi-automated, automated methods
General reference interval
  18-22 seconds
↑ 25 seconds or longer
  Dysfibrinogenemia, hypofibrinogenemia, afibrinogenemia

## Slide 71: Table 36-7 Differentiation of Conditions Associated with a Prolonged Thrombin Time Using the Reptilase Time
Condition | Thrombin Time | Reptilase Time
Heparin contamination | Prolonged | Normal
Dysfibrinogenemia | Prolonged | More prolonged
Presence of fibrin degradation products | More prolonged | Prolonged

## Slide 72: Prekallikrein Screening Test
Patients with P K (Fletcher factor) deficiency
  Have a prolonged A P T T
  Correction of prolonged A P T T to normal after
    10-minute incubation period before adding calcium chloride
    Longer incubation-↑ contact activation of F X I I
  Confirmed with specific factor assay using P K-deficient substrate

## Slide 73: Factor Thirteen (1 of 2)
F X I I I
  Necessary for stable fibrin clot formation
  Occurs by forming covalent bonds between fibrin monomers
  Routine screening does not detect deficiency
Screening test principle
  Fibrin clot has ↑ solubility because of the lack of cross-linking of fibrin polymer in absence of F X I I I

## Slide 74: Factor Thirteen (2 of 2)
Procedure
  Patient’s P P P is clotted with calcium chloride
  Allowed to clot for 1 hour at 37°Celsius
  Clot is placed in urea or 1% monochloroacetic acid in 37°Celsius
  Factor
  deficiency
    Patient’s clot dissolves within 24-hour period
    Patient’s clot dissolves within 30 minutes
      Indicates activity <2%

## Slide 75: Laboratory Testing for V W F (1 of 5)
Diagnosis of V W D requires
  Quantitative V W F by both functional and antigenic methods
The National Heart Lung and Blood Institute (N H L B I)
  Recommends that all initial testing for V W D include
    V W F ristocetin cofactor assay (V W F:R C o) or equivalent
    V W F antigen levels (V W F:A g)
    F V I I I clot-ting activity (F V I I I:C)

## Slide 76: Laboratory Testing for V W F (2 of 5)
Variables that affect accurate determination
  Difference over time in V W F results in patients with V W D
    Sometimes V W F:A g and/or V W F:A levels will be normal
  Results of screening tests can be normal in
  Endogenous release of V W F
    Can result in transient F V I I I and V W F increase
      Patient should be calm during phlebotomy
  Processing of citrated sample
  Reference standard
  Patient’s blood type
  Variety of clinical disorders

## Slide 77: Table 36-8 Relationship between Blood Groups and Von Willebrand Factor Level
Blood Groups | von Willebrand Factor Level (%)
O | 74.8
A | 105.9
B | 116.9
AB | 123.9

## Slide 78: Laboratory Testing for V W F (3 of 5)
Treatment for classic V W D
  Desmopressin (D D A V P)
    Causes two-fold to five-fold ↑ in patients V W F
    plasma level in
  Test dose of D D A V P is given to patient prior to surgery
    Citrated plasma samples are drawn
      Preinfusion
      30 minutes postinfusion
      4 hours postinfusion
    Assayed for F V I I I activity, V W F:A g, V W F:A

## Slide 79: Laboratory Testing for V W F (4 of 5)
V W F:R C o assay (ristocetin cofactor assay)
  In presence of ristocetin
    V W F induces platelet agglutination
    Measured using aggregometer, automated analyzer
    Slope plotted versus percent V W F activity
      Results extrapolated from a calibration curve

## Slide 80: Laboratory Testing for V W F (5 of 5)
V W F:A g
  E L I S A assay
  Rapid methods
    Automated, immunoturbidimetric assays
  Latex immunoassay (L I A) test procedure
    Microlatex particles coated with rabbit antihuman V W F antibodies
    Produces agglutination when mixed with plasma containing V W F
    Extent of agglutination (turbidity) correlates to the level of V W F:A g present in the citrated plasma sample

## Slide 81: V W F Multimer Analysis
Determine correct disease subtype
  Plasma electrophoresis
    Low concentration of agarose (0.65%)
    Staining with radiolabeled A b to V W F or
    S D S-P A G E
  Also useful for diagnosing T T P
    Presence of U L-V W F multimers indicates decreased A D A M T S-13 activity
  Performed in specialized reference labs or coagulation centers

## Slide 82: Collagen-Binding Assays for V W D
Specialized assays
  Differentiate V W D type 2A and 2B from 2M
  Perform both collagen-binding assays and V W F:R C o assays
    Increases ability to differentiate V W D type 2 variants
      Only 2A and 2B have abnormal V W F:C B

## Slide 83: A D A M T S-13
V W F-cleaving protease
  Methodologies
    E L I S A, fluorescence resonance energy transfer (F R E T), and other direct or indirect assays available
    Measure A D A M T S-13 activity, antigen, or autoantibody levels
    Use citrated P P P
    Generally performed in specialty labs

## Slide 84: Identification of Inhibitors
Most common circulating inhibitors
  Lupuslike anticoagulants (L A s) or Antiphospholipid antibodies (a P L s) and
  F V I I I inhibitors
Measurement of the inhibitor to F V I I I
  Determined with the F V I I I inhibitor assay (Bethesda method or Nijmegen-Bethesda)

## Slide 85: L A/Antiphospholipid Antibodies (1 of 6)
Originally described as lupus anticoagulants
  Now referred to as a P L
  Usually I g G
    Directed against the protein component of protein-phospholipid complexes
  In vivo-thrombotic tendency
  In vitro-prolong phospholipid-dependent clotting assays

## Slide 86: L A/Antiphospholipid Antibodies (2 of 6)
Diagnosis requires demonstration of:
  Abnormal P L dependent clotting test
  Presence of and inhibitor (clotting test does not correct in mixing study)
  Phospholipid-dependent inhibitor
Used to evaluate
  Patients with history of recurrent abortions, venous or arterial thrombosis
  As part of the evaluation of autoimmune disorders
  Potential cause for elevated A P T T (with or without elevated P T)

## Slide 87: L A/Antiphospholipid Antibodies (3 of 6)
Sample (P P P) integrity important
Two different screening tests recommended
  Dilute R V V T (d R V V T)
    Recommended as initial screening test
  A P T T (silica as activator, low phospholipid content)
    Second test for screening
If either test is prolonged
  Proceed with designated testing
If patient presents with a positive history of thrombophilia and A P T T is not prolonged
  Should proceed with testing for L A/a P L

## Slide 88: L A/Antiphospholipid Antibodies (4 of 6)
I S T H guidelines for presence of L A
  Double centrifugation to ensure P P P
    If testing delayed
      Quick freeze; thaw at 37°Celsius
  Two clot-based assays using different clotting principles
    d R V T T-most robust
    Low P L A P T T with silica activator
  Perform T T to detect heparin (absorb if necessary)

## Slide 89: L A/Antiphospholipid Antibodies (5 of 6)
I S T H guidelines for presence of L A
  P N P for mixing studies prepared by double centrifugation and test 1:1 mixture without incubation
  Develop cutoff values on at least 40 healthy adult donors <50 years old
  Perform mixing studies on donor plasmas and calculate the index of circulating anticoagulant (I C A) as follows:
  Perform confirmatory test by increasing concentration of P L in screening test

## Slide 90: L A/Antiphospholipid Antibodies (6 of 6)
I S T H guidelines for presence of L A
  Confirmatory testing of healthy donor plasma at low (screen) and high (confirm) P L concentrations to obtain C Ts
    Cutoff is value corresponding to mean of individual % corrections defined as:
    Result is confirmation of L A if percentage correction is above defined cutoff value
  Results should be normalized (ratio of result/P N P) and with accompanying interpretation

## Slide 91: Dilute Russell’s Viper Venom Test (d R V V T) (1 of 2)
Venom from the Russell viper (Daboia)
Screening reagent
  Dilute Russell’s viper venom, calcium chloride, phospholipids (low level)
  Added to patient’s P P P
  Activates F X to produce clot

## Slide 92: Dilute Russell’s Viper Venom Test (d R V V T) (2 of 2)
Determine ratio of patient’s C T to C T of normal control
  L A/a P L is present if patient’s d R V V T is longer than normal control
    Ratio of patient’s C T to normal control C T is <1.2 normally
  Confirmatory test
    Use higher P L concentration R V V
    Report ratio of high and low P L tests

## Slide 93: Hexagonal Phospholipids
Substitutes egg phosphatidylethanolamine in a hexagonal phase configuration (H P P)
  a P L Abs recognize the H P P configuration
  Addition of H P P
    Neutralizes inhibitory effect of the a P L antibodies
    Does not neutralize factor-specific antibodies
  Incubate test plasma with and without H P P
    Perform A P T T on both mixtures
    Test with H P P will have a shortened clotting time

## Slide 94: Specific Factor Inhibitor Assay (1 of 2)
Nijmegen-Bethesda assay
  Developed to measure F V I I I inhibitors
    Develop in 3-52% of patients with severe hemophilia A
  Can be used for other inhibitors to coagulation proteins
Specific inhibitors
  Occur in 10-15% of hemophiliacs at any time after first factor concentrate infusion
  Monitor patient’s response to treatment by ordering F V I I I:C assays
  If bleeding does not stop-could be presence of inhibitor

## Slide 95: Specific Factor Inhibitor Assay (2 of 2)
Nijmegen-Bethesda inhibitor assay
  Mix patient’s plasma with equal volume of P N P of known F V I I I activity
  Perform F V I I I assay on incubation mixture to measure residual activity
  Convert activity to Bethesda units (B U) using standard Bethesda chart
  Testing should be performed on pre-infusion trough sample

## Slide 96: Laboratory Investigation of the Fibrinolytic System

## Slide 97: Introduction-Laboratory Investigation of the Fibrinolytic System
D-dimer
Fibrin degradation products
Euglobulin clot lysis
2-antiplasmin activity

## Slide 98: D-Dimer (1 of 3)
Specific degradation fragment
  Results from plasmin lysis of fibrin clot
    Fibrin clot has been cross-linked by F X I I I a
  Excellent marker for D I C
  Also elevated in
    V T E
      Pulmonary embolism
      Deep vein thrombosis
    Trauma, surgery, sepsis

## Slide 99: D-Dimer (2 of 3)
D-dimer test
  Immunoturbidometric method
  Uses monoclonal antibodies against D-dimer fragment
  Three methodologies
    E L I S A
    Manual latex- or hemagglutination
    Automated immunoturbidimetric methods

## Slide 100: D-Dimer (3 of 3)
D-dimer test
  Most laboratories use automated methods
    Antibody-coated latex particles aggregate in the presence of D-dimer
      Increases turbidity of the sample
    Increase in scattered light is proportional to the amount of D-dimer in the sample.
  D-dimer level is not elevated
    No thrombotic process is ongoing
  D-dimer level is elevated
    V T E or another condition

## Slide 101: Table 36-9 Partial List of Non-Venous Thromboembolism Causes of Elevated D-Dimer
D I C
Trauma
Surgery
Hematoma
Arterial thrombosis
General hospitalization
Pregnancy
Cancer
Diabetes
Thrombolytic therapy
Older age

## Slide 102: Fibrin Degradation Products (1 of 2)
↑ fibrinolytic activity
  ↑ levels of F D Ps
  Sample
    Special collection tube containing
      Thrombin-clots sample and ensures all fibrinogen has been removed
      Fibrinolytic inhibitor-prevents in vitro fibrinolysis
  Patient’s serum mixed with latex particles coated with antibodies to F D P
  Observes reaction mixture macroscopically

## Slide 103: Fibrin Degradation Products (2 of 2)
Normal reference interval varies according to manufacturer and sensitivity
Largely replaced by D-dimer
  F D P may be useful in acute fibrinogenolysis (increased F D P with discordant or negative D-dimer)

## Slide 104: Euglobulin Clot Lysis
Used before specific antibodies to fibrinolytic products
Euglobulin protein fraction is precipitated from plasma
  Fibrinolytic inhibitors remain in supernatant
  Precipitate is re-dissolved, clotted with thrombin
  Observed for clot lysis
Shortened lysis time-indicates ↑ fibrinolytic activity

## Slide 105: 2-Antiplasmin Activity
2-antiplasmin
  Inhibitor, regulates fibrinolytic system
  Congenital deficiencies
    Characterized by bleeding occurring some hours after an initial injury
      Hemostatic plug breaks down prematurely
  Acquired decreases
    Observed in liver disease and D I C
  Chromogenic method
    Can be used with a microtiter plate, test tube, or auto-mated procedures

## Slide 106: Laboratory Investigation of Hemostasis (Primary, Secondary, or Fibrinolytic Pathway) Using Global Assays

## Slide 107: Global Testing (1 of 4)
Global testing instruments-P O C hemostasis testing
  Analyze entire hemostasis process
    Coagulation
    Anticoagulation effects
    Fibrin formation and stabilization
    Clot retraction (platelet function)
    Fibrinolysis

## Slide 108: Global Testing (2 of 4)
Thromboelastography (T E G)
  Near-patient testing method, assesses blood coagulability
  Evaluates torque created from blood clotting onto a stationary pin in an oscillating cup
  Whole blood or citrated whole blood may be used
    Manufacturer provides activators
    Citrated samples require calcium replacement
  T E G® platelet mapping assay
    Assess patients who received platelet-inhibiting drugs
      Aspirin, N S A I D S, clopidogrel, other P 2 Y12 inhibitors

## Slide 109: Global Testing (3 of 4)
R O T E M® (T E M) Haemostasis Whole Blood Analyzer
  Computerized device based on T E G
    Evaluates the torque created from blood clotted within a stationary cup onto an oscillating pin
  Evaluates
    Intrinsic pathway (I N T E M, contact factor activation)
    Extrinsic pathway (E X T E M, tissue factor activation)
    Fibrinogen assessment (F I B T E M, tissue factor activation with platelet inhibition)
    Heparin anticoagulation (H E P T E M, contact activation with heparinase)
    Whole blood or citrated whole blood may be used

## Slide 110: Global Testing (4 of 4)
Calibrated automated thrombogram (C A T)
  Thrombin generation test
    Monitors the rise and fall of thrombin in clotting plasma
  C A T
    T G T that can measure both hypercoagulability and hypocoagulability
    Sensitive to action of most antithrombotic drugs

## Slide 111: Figure 36-6 Principle of Thromboelastography
Notes: A blood sample is placed in a cup that has a pin carefully placed within and that is connected to a torsion wire. As the cup rotates in a back and forth movement, the aggregates formed within it cause the wire to become more rigidly placed and reflects the strength of the aggregates formed within the cup. The movement or lack of movement is reflected via either an optical or magnetic detector that produces a graphic presentation.

## Slide 112: Figure 36-7 Thromboelastography Results
Notes: The graphic presentation shows the lag phase (r-time) where the blood is still fluid, and therefore the wire hangs freely and is not affected by the rotation of the cup. As thrombin is formed, a thrombus forms in the rotating cup, which transmits its movement onto the suspended sensor. The kinetics time (k-time) is the rate of clot formation and the time necessary for the clot amplitude to increase from 2 to 20 mm. The maximum amplitude reveals the stabilization of clot by platelets and FXIII. Finaly, the lysis onset time measures the degree of fibrinolysis after a set time frame from maximum amplitude completion.

## Slide 113: Laboratory Investigation of Hereditary and Acquired Thrombophilia (Hypercoagulable States)

## Slide 114: Laboratory Investigation of Hypercoagulable States (1 of 2)
Hypercoagulable state
  Patients will have repeated thrombotic episodes
  Laboratory tests assist in determining if thrombosis is due to
    Hereditary abnormalities
    Acquired abnormalities
Methodologies for hypercoagulable state can be
  Clotting assays
  Chromogenic assays
  E L I S A assays
  Latex immunoassays

## Slide 115: Laboratory Investigation of Hypercoagulable States (2 of 2)
Five clinical categories when to test for these conditions
  Individuals with provoked venous thromboembolism
  Individuals with venous thromboembolism that is unprovoked
  Relatives of subjects who have experienced a thrombotic event
  Female relatives of a subject who has experienced thrombosis, and who may be considering the use of estrogen medications
  Female relatives of patients who have had a thrombotic event and who are considering becoming pregnant

## Slide 116: Antithrombin (A T) (1 of 2)
Previously called antithrombin
  Serine protease inhibitor
  With heparin, has powerful and immediate antiprotease action
  If patient has ↓ A T, poor response to heparin
  Chromogenic two-part assay
    Measuring A T activity in the presence of heparin
    First-plasma incubated with excess thrombin or F x a in the presence of heparin
    Second-residual thrombin or F x a activity is measured
      Inversely proportional to the test sample A T level

## Slide 117: Antithrombin (A T) (2 of 2)
Other tests for A T
  Clot-based method
  Immunological methods
Congenital deficiencies are rare
Acquired deficiencies
  D I C, liver disease, nephrotic syndrome
Heparin decreases plasma A T levels

## Slide 118: Protein C (1 of 3)
Activated protein C (A P C)
  Anticoagulant that inactivates F V a and F V I I I a
Functional assays of protein C (P C)
  Analyzed by measuring
    Amidolytic activity using a chromogenic substrate or
    Anticoagulant activity in a clot-based assay

## Slide 119: Protein C (2 of 3)
Clot-based assay
  P T T reagent containing protac (P C activator derived from snake venom)
    Activates P C and contact factors of intrinsic pathway
  Patient’s plasma incubated with P C-deficient plasma and P T T reagent containing the activator followed by calcium chloride
  Time for clot formation is measured
  Activation of P C, results in inactivation of F V a and F V I I I a
  Prolongation is proportional to plasma P C activity

## Slide 120: Protein C (3 of 3)
Chromogenic assay
  Widely used
  Patient’s plasma is incubated with a specific P C activator (protac)
  Chromogenic substrate is added
  A P C is measured based by assessing the amount of colored product produced

## Slide 121: Protein S (1 of 2)
Circulating protein S (P S)
  Exists in two forms
    Free (40%)
      Serves as a cofactor for A P C
      Enhancing anticoagulant activity
    Bound to C4b binding protein (60%)

## Slide 122: Protein S (2 of 2)
Clot-based assays-free protein S
  Based on the ability of P S to serve as a cofactor for the anticoagulant effect of A P C
  Patient P P P is mixed with P S-deficient plasma
  A P C and F V a are added to mixture and incubated followed by addition of calcium chloride
  C T is measured
  C T is proportional to P S activity in sample

## Slide 123: Activated Protein C Resistance (A P C R)
One of the most common risk factors for thrombosis
  Presence of the factor V Leiden (F V L) mutation
    Results in resistance of F V a to degradation by A P C
Clot-based assay principle
  Addition of A P C to plasma sample induces a prolongation of A P T T mediated by inactivation of F V a and F V I I I a in plasma sample
  Screening test-not a diagnostic for F V L
Diagnosis requires clotting and molecular (P C R) testing for F V Leiden mutation
  10% of patients with A P C R have mutations other than F V Leiden

## Slide 124: Prothrombin G20210A
Prothrombin G20210A mutation
  Patients heterozygous for mutation can have prothrombin levels up to 30% higher than noncarriers
  Clot-based tests unreliable
  Genetic screening

## Slide 125: Additional Testing for Thrombosis
Other testing for thrombosis
  Plasminogen
  α2 antiplasmin activity (α2-antiplasmin)
  Plasmin-α2-antiplasmin (P A P) Complex
  Plasminogen activator inhibitor (P A I-1)
  Tissue plasminogen activator (t P A)
  Thrombin activatable fibrinolysis Inhibitor (T A F I)

## Slide 126: Molecular Markers of Hemostatic Activation

## Slide 127: Molecular Markers of Hemostatic Activation (1 of 2)
Molecular markers of:
  Platelet activation
  Generation of coagulation enzymes
  Products of intravascular fibrin formation or dissolution
Useful for early detection of:
  Platelet-driven and coagulation-driven disorders
Improve medical care for:
  Thrombotic and fibrinolytic disease
Testing is not routine

## Slide 128: Molecular Markers of Hemostatic Activation (2 of 2)
Process of activating coagulation
  Converts a number of zymogens to active serine proteases
  Direct measurements of most hemostatic enzymes in vivo not possible
    Naturally occurring protease inhibitors
  Activation peptides released on activation can be measured by immunochemical assays (R I A, E L I S A)
  Enzyme-inhibitor complex can be measured by similar immunochemical methods
  Routine laboratories do not offer these due to
    High cost and low volume of test requests

## Slide 129: Table 36-10 Immunochemical Markers of Hemostatic Activation
Coagulation | Blank
Activation peptides | Factor I X activation peptide
Blank | Factor X activation peptide
Blank | Prothrombin activation fragment F1 + 2
Blank | Protein C activation peptide
Blank | Fibrinopeptide A, fibrinopeptide B
Enzyme-inhibitor complexes | Thrombin-antithrombin complex (T A T)
Blank | Factor X a-antithrombin complex
Blank | Factor I X a-antithrombin complex
Blank | Activated protein C-protein C inhibitor complex
Fibrinolysis | Bβ1-42 fragment (plasmin cleavage of fibrin and fibrinogen)
Blank | Plasmin-antiplasmin complex (P A P)
Platelet activation (flow cytometry) | Activated G P I I b/I I I a complex P-selectin (C D62-P) G P I b/I X/V (decrease)

## Slide 130: Laboratory Evaluation: Assessment for Pharmaceutical Intervention of Hemostasis

## Slide 131: Oral Vitamin K Antagonist Therapy (1 of 5)
Monitoring oral anticoagulant therapy (Coumadin, warfarin)
  International normalized ratio (I N R)
    Determined from the P T
    Used to standardize P T results for various reagents and instruments
  W H O international reference standard for thromboplastin reagent
    Allows for calibration of commercial thromboplastins

## Slide 132: Oral Vitamin K Antagonist Therapy (2 of 5)
International sensitivity index (I S I)
  Each manufacturer has an in-house standard
    A specific lot of its own thromboplastin
    Calibrated against the W H O standard
  Calculated as a specific correction factor
  I S I for a particular reagent is instrument dependent
  Value is listed on each package (lot number) of thromboplastin
    Different I S I for different instruments

## Slide 133: Oral Vitamin K Antagonist Therapy (3 of 5)
International normalized ratio (I N R)
  Most laboratories report both the P T and I N R
  I N R is the standard for monitoring long-term anticoagulant therapy
  Is independent of reagents and methods
  To perform calculations need
    Laboratory’s mean value for the P T in at least 20-30 normal control subjects using the specific lot of thromboplastin with the instrument
    I S I of reagent used

## Slide 134: Figure 36-8 International Normalized Ratio (I N R) Calculation Formula
Where
R = P T ratio obtained with the working thromboplastin
P T ratio = Patient’s P T/Mean normal P T
I S I = International sensitivity index (provided by manufacturer)
Example: Patient’s P T = 21.5 seconds
Mean normal P T = 12.0 seconds
I S I = 1.35

## Slide 135: Oral Vitamin K Antagonist Therapy (4 of 5)
Therapeutic ranges
  Hypercoagulable state (arterial or venous thromboembolism)
    I N R between 2.0-3.0
  Mechanical valve replacement
    I N R may be higher
Oral anticoagulant treatment influenced by:
  Body mass, Vitamin K content in diet, drug interactions and supplements, wellness, hepatic function

## Slide 136: Oral Vitamin K Antagonist Therapy (5 of 5)
P O C instruments
  Determine P T/I N R from capillary puncture
  No C a++ required
    Blood sample has not been citrated
  Used by patients, anticoagulation clinics, and physician office laboratories

## Slide 137: Direct Oral Anticoagulants (D O A Cs)
Direct thrombin inhibitor
F X a inhibitors
  Stroke prevention, treatment, and prophylaxis of V T E
No known dietary restrictions
Generally do not require routine continuous monitoring
Screening tests (P T, A P T T) have variable sensitivity to D O A Cs
I N R should not be used
  As a gauge for anticoagulation
  As an indicator for comparing D O A C to oral vitamin K antagonist anticoagulation

## Slide 138: Heparin Therapy Monitoring
Two types of heparin
  Unfractionated heparin (U H or U F H)
    Traditionally monitored by A P T T; can also use T T
  Low-molecular-weight heparin (L M W H)
    Common treatment for patients with/at risk for thrombosis
    Enoxaparin, dalteparin, tinzaparin, nadroparin, and fondaparinux (pentasa)
    Cannot be monitored with A P T T
    Monitor with X a inhibition assay if necessary

## Slide 139: A P T T and Heparin Monitoring
Monitoring U F H therapy
  No standardization of the A P T T to normalize differences in reagents and test systems
  Weight-based heparin dosing nomogram should be provided to clinicians
    Used to determine appropriate rate of heparin infusion to reach and/or maintain a therapeutic range
  Great variation in patients’ response to heparin
    Many factors can impact heparin monitoring
  Anti-X a chromogenic assay
    Avoids problems with variables that affect the A P T T
    Can be used to monitor U F H and L M W H

## Slide 140: Table 36-11 Factors Affecting the Monitoring of Heparin
Conditions that influence the pharmacokinetics or general bioavailability of heparin: heparin binds to various plasma proteins (e.g., platelet factor 4, fibrinogen, V W F, fibronectin, and vitronectin), macrophages, and endothelial cells, resulting in decreased bioavailability
Conditions that alter the characteristics of the A P T T dose-response to heparin
Conditions that cause an abnormal baseline A P T T (factor deficiency, lupus anticoagulant)
Patient’s body weight
Antithrombin level
Increased levels of Fibrinogen and/or F V I I I (heparin “resistance”)
Specimen collection and processing problems
Time of collection (minimum of 4 hours needed to achieve steady state with heparin)
Delayed testing (can result in innacurate results due to factor lability and P F4 release)
Failure to achieve platelet-poor plasma (can result in release of P F4 from activation of residual platelets resulting in underestimation of heparin in the sample)

## Slide 141: Heparin Therapy-T T
Less commonly used
Advantage
  Not influenced by plasma factor deficiencies
Can be used to detect
  Presence of heparin
  Exquisitely sensitive to dabigatran, the I I a D O A C
May be useful in patients with elevated baseline A P T Ts
May serve as a surrogate for assessing residual U F H when transitioning to other A P T T monitored therapy

## Slide 142: Heparin Therapy-Anti-X a Assay
Anti-X a chromogenic assay
  Excess F X a is added to patient’s P P P and incubated
    Heparin with A T in patient’s P P P inhibits F X a
  Chromogenic substrate is added to mixture
  Residual F X a cleaves chromogen and produces yellow color
    Measured spectrophotometrically

## Slide 143: Hemostasis Instrumentation

## Slide 144: Electromechanical Clot Detection Systems (1 of 2)
Electromechanical systems
  Measure time for clot to form in plasma
  Detect changes in reaction mixture
  Are insensitive to lipemia, hemolysis, bilirubin, turbidity
Fibrometer
  Detects the completion of an electrical circuit between two electrodes when a clot forms
  Largely replaced with more advanced instruments

## Slide 145: Electromechanical Clot Detection Systems (2 of 2)
Diagnostica Stago analyzers
  Monitors amplitude of steel ball in a magnetic field as plasma viscosity changes during clot formation
    Amplitude decreases as viscosity increases and timer stops
Electromechanical systems are insensitive to lipemia, hemolysis, bilirubin, turbidity

## Slide 146: Optical/Photometric/Turbidimetric Clot Detection Systems
Measure increase in turbidity when fibrinogen is converted to fibrin
  Change in O D and end point of clot formation are measured
A C L T O P C T S analyzer
  Beckman Coulter Inc./Instrument Laboratory
  671 nm wavelength range L E D light
    Significantly eliminates lipemia, hemolysis, and bilirubin interference
Sysmex C As and B C S analyzers
  Utilize optical/turbidimetric clot detection methods

## Slide 147: Chromogenic Detection Methods
Use chromophore (color-specific generating substance) (e.g., p N A) which has maximum absorbance of 405 nm
Enzyme of interest cleaves chromogenic substrate and releases p N A
Color intensity directly proportional to amount of p N A released
Directly or indirectly related to concentration of the measured substance

## Slide 148: Immunologic Detection Methods
Latex microparticles coated with specific antibody against antigen being measured are mixed with test plasma
Agglutination or turbidity results in increased light absorbance
Antigen level is proportional to amount of light absorbed

## Slide 149: Selecting a Hemostasis Analyzer (1 of 2)
When selecting an analyzer, consider
  Methodologies
  Random access testing
  Reagent systems
  Level sensing
  Barcode reading
  Tube sampling
  Computer and software features

## Slide 150: Selecting a Hemostasis Analyzer (2 of 2)
When selecting an analyzer, consider
  Reflex testing
  Throughput
  S T A T processing
  Cost savings
  Downtime
  Footprint
  Training and support

## Slide 151: Point-of-Care Hemostasis Instrumentation
Bedside or self-monitoring hemostasis testing systems
  Portable or handheld
  Require very little sample
Features and functions
  Multiple specimen types
  Integrated Q C with each analysis
  Automatic Q C lockout
  Onboard system for automatic error detection