Clinical Laboratory

| SECTION: URINALYSIS |  |
| --- | --- |
| DOCUMENT TYPE: Maintenance |  |
| TITLE: Maintenance: Clinitek Status |  |
| APPROVED BY: Medical Director / Administrative Laboratory Director | DOCUMENT #: 12.0343.3 |
| WRITTEN BY: Charlotte Yuza, MLT(ASCP) | EFFECTIVE: 12/1/11 |
| SCOPE: Hospital, NW, Branch Clinic & FastCare Laboratories | COMPLEXITY: NA |

Purpose:

To outline preventative maintenance for Clinitek Status.

Policy:

Daily cleaning of the Clinitek consists of cleaning the Test Table.  Disinfecting and cleaning of the calibration bar

is performed only if necessary.

Supplies:

- Cotton tipped swab

- Lint free tissue or soft cloth

- Bleach solution 5% or Isopropyl Alcohol 70%

Procedure:

Daily

Cleaning of the Test Table

- Remove test table and rinse both sides of the table under running water.

- Wet a cotton-tipped stick with water and carefully clean test table (except for the white calibration bar).

- Air dry table or dry with a soft cloth or lint free tissue.

- Re-insert table into the analyzer by holding the table at the end opposite the white calibration bar, with the white calibration bar facing upwards.

- Push the test table firmly but slowly, just over halfway into the analyzer.

- Replace the test table insert.

As Needed

Disinfecting the Test Table and Insert

- Prepare cleaning solution (bleach or alcohol) in a tall, narrow container to a depth of about 4 inches.

- Place the insert and/or test table into the solution, making sure the white calibration bar on the test table remains above the liquid level.

- Soak the table and insert for a minimum of 2 minutes and a maximum of 10 minutes.

- Do not soak longer than 10 minutes.

- Rinse the test table and insert thoroughly with water.

- Dry with a damp cloth and replace the test table and table insert in the analyzer.

Cleaning the White Calibration Bar

- Remove the insert from the test table.

- Remove the test table by pulling it slowly out of the analyzer.

- Check the white calibration bar on the test table for dirt or discoloration.

- If the calibration bar is clean and unmarked, replace the table into the analyzer.

- If the white calibration bar is dirty or discolored, gently wipe and clean it with a new cotton-tipped stick or lint-free cloth wetted with distilled water.

- Allow the calibration bar to air dry and then inspect the surface for dust, foreign material, scratches, or scuffs.

- Reinsert the test table.

Cleaning the outside of the analyzer

- Turn the analyzer off by pressing the on/off button for 2 seconds.

- Wipe the outside (including the display) with a damp (not wet) cloth and a mild detergent.

  - Do not use any type of solvent, oil, grease, silicone spray, or lubrication on the analyzer.

  - Do not spray the glass cleaner directly onto the screen.  Don not use laboratory wipes such as kimwipes, since they may scratch the screen.

  - Care should be taken to avoid liquid from entering the printer compartment.

- The display may be disinfected using the same solutions on and allow to remain for 10 minutes.  Wipe clean using a clean cloth dampened with water, then dry.

References:

Clinitek Status Operator’s manual revision B 2004-10

CAP/COLA Standards:

URN.26400	Routine maintenance

| # Approved By | # Date | # Reviewed/Revised By | # Date | # Revision |
| --- | --- | --- | --- | --- |
| Medical Director | 3/31/09 | C Yuza | 3/31/09 | New |
| Dr Vege | 6/30/2010 | C Yuza | 4/23/10 | Added Hospital to scope and CAP standards, updated format |
| Dr Vege | 6/16/11 | C Yuza | 6/14/11 | Annual review |
| J DeClercq Dr Vege 1/9/12 | 11/29/11 | C Yuza | 11/22/11 | Removed Clinitek 50, added cleaning of outside of analyzer instructions. |

Similar Procedures (Specify title of procedure and other OMC locations used if applicable)

Attachments: