Clinical Laboratory

| SECTION: Equipment |  |
| --- | --- |
| DOCUMENT TYPE: Policy |  |
| TITLE: Preventative Maintenance Program |  |
| APPROVED BY: Laboratory Supervisor | DOCUMENT #: 13.0007.1 |
| WRITTEN BY: Charlotte Yuza, MLT(ASCP) | EFFECTIVE: 8/1/12 |
| SCOPE: FastCare and Skyway Laboratories | COMPLEXITY: NA |

Purpose:

To describe the preventative maintenance program for FastCare and Skyway Laboratories.

Policy:

- Safety / Electrical checks are performed on all equipment prior to being put into service and at least annually thereafter by the OMC Maintenance staff.

- Preventative maintenance and cleaning is performed per manufacturer’s recommendations.

- Cleaning is performed using manufacturer’s recommended cleaners.

Procedure:

Definitions:

General Cleaning is defined as the cleaning of the outside of the instrument or item of splashes, spills, dust, etc.

Internal Cleaning is defined as a manufacturer specified cleaning procedure; refer to individual maintenance procedures for details.

| Maintenance Schedules |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Instrument | Cleaning Schedule | Cleaning Performed by: | Preventative Maintenance Schedule | Preventative Maintenance Performed by: | Scheduled Replacement (Life span) |
| Clinitek Status | Daily- general cleaning Weekly- clean table | Lab staff | NA | NA | 6 - 10 years |
| Incubator LabLine | Daily- general cleaning | Lab staff | NA | Lab staff | 6 - 10 years |
| Thermometers (digital only) | Daily- general cleaning | Lab staff | Annually- replace batteries | Lab staff | 6 - 8 years |
| Refrigerator | Daily- general cleaning Semiannually- internal cleaning | Lab staff | Only as needed- vacuum coils | Maintenance Dept | 6 - 10 years |
| Label Printer | Monthly- internal clean platen | Lab staff | NA | NA | 8 - 10 years |
| Laser Printer | Semiannually- internal clean with canned air | Lab staff | NA | NA | 8 - 10 years |
| Counter tops | Daily- general cleaning | Lab staff | NA | NA | 20+ years |
| Specimen Coolers | Daily- general cleaning Semiannually- clean all surfaces | Lab staff | NA | NA | 8 - 10 years |
| Humidifiers | Daily- general cleaning Monthly (while in use) - internal cleaning change filters. | Lab staff | NA | NA | 3 – 5 years |

COLA Standards:

## COLA MA 20	Manufacturer’s Maintenance Instructions

COLA MA 21	Preventative Maintenance Performed

| # Approved By | # Date | # Reviewed/ Revised By | # Date | # Revision |
| --- | --- | --- | --- | --- |
| C Yuza | 7/31/12 | C Yuza | 7/31/12 | New locations |

Similar Procedures (Specify title of procedure and other OMC locations used if applicable)

..\..\URINALYSIS & MICROSCOPY\MAINTENANCE\Maintenance- Clinitek Status.doc

..\PROCEDURES\Humidifier Usage and Preventative Maintenance.doc

Attachments: