Clinical Laboratory

| SECTION: Equipment |  |
| --- | --- |
| DOCUMENT TYPE: Procedure |  |
| TITLE: Min-Max Thermometers |  |
| APPROVED BY: Medical Director | DOCUMENT #: 13.0323.5 |
| WRITTEN BY: Judy Vaughan, MT(ASCP) | ### EFFECTIVE: 2/1/2012 |
| SCOPE: Branch & Southeast Clinic Laboratories | COMPLEXITY: NA |

Purpose:

To describe the procedure for reading min-max thermometers.

Policy:

Min-Max thermometers are used to monitor refrigerator temperatures in Laboratories not staffed 24/7 and not monitored by TempTrak.

Supplies:

Min-Max thermometers

Procedure:

Southeast Clinic Thermometers:

HealthCare Logistics Thermometers:

- The unit has two operating modes:

    - Normal Display Mode

    - Alarm Display Mode

    - To change from one mode to another, press MODE button once.

- Normal display mode shows the current temperature and the maximum and minimum temperature at the same time.

- Resetting Minimum and Maximum Memories:

    - Reset both memories: Press (MEMORY CLEAR) button once.

    - Reset individual Max or Min memories: Press either (MAX) or (MIN) button once to reset the

respective memory. The other memory will be unaffected.

    - Resetting a memory will result in current temperature being displayed in the max or min

memory display.

- Minus 50ºC to +70ºC is the temperature range for this thermometer.

Fisher thermometers:

1.	Normal display mode shows the current temperature.

2.	Press MAX button to display maximum temperature record in memory.

3.	Press MIN button to display minimum temperature record in memory.

2.	Clearing Maximum Temperature Memory:

- Press MAX and hold, and then only MEM at the same time for memory clear.

3.	Clearing Minimum Temperature Memory:

- Press MIN and hold, and then only MEM at the same time for memory clear.

Branch Clinic and Rochester NW Thermometers:

- Branch Clinic Min/Max thermometers display the current temperature and min / max alarm settings on the display screen.

  - If the display becomes dim then change the battery with the same type of battery.

  - After changing the battery the temperature alarm settings will be reset to default values.

- To reset the minimum and maximum temperatures:

  - Press MAX/MIN to display maximum measured reading.

  - Press the key again to return to display minimum measured reading.

  - Press the key again to return to normal display.

  - To clear memory, press and hold the key for about 3 seconds then release or switch thermometer off then on. Before taking the Max/Min reading memory, clear the memory once.

- Temperature Alarm Setting

  - Press Set to enable alarm setting and the HI value (default is OFF) will be flashing.

  - Use ↓ or ↑ key to set the desired alarm high limit.

  - Press Set again and the LO value (default is OFF) will be flashing.

  - Use ↓ or ↑ key to set the desired alarm low limit.

  - Press Set to finish setting.

    - During the limit setting, holding down ↓ or ↑ will advance the value the automatically and using the ON/OFF key can turn the limit off or on.

    - If the temperature value is out of the set limit them the alarm will sound.

    - Press any key will stop the alarm sounding but the HI or LO icon will keep flashing until the temperature value is within the set limit.

Procedure Notes:

- Min-Max thermometers are reset daily to detect any temperature fluctuations when the Laboratory is not staffed.

- Any temperature outside of acceptable ranges must be investigated, corrective action taken, and contents evaluated for any adverse effects. Document on appropriate log.

References:

Min-Max Thermometer instruction inserts.

CAP Standards:

CHM.24700		Temperature Checks

COLA Standards:

MA7, MA8, and MA16

| # Approved By | # Date | # Reviewed/Revised By | # Date | # Revision |
| --- | --- | --- | --- | --- |
| Dr Bradley Karon | 4/2007 | J Vaughan | 4/2007 | New format |
| Dr Durga Vege | 2/2008 | J Vaughan | 2/2008 | Updated Main Clinic instructions |
| Jennifer De Clercq | 2/2009 | J Vaughan | 2/2009 | Removed Hospital |
| Dr Durga Vege | 2/1/2010 | J Vaughan | 1/29/2010 | Updated format, Policy and SE Clinic Instructions |
| Dr Durga Vege | 2/16/2011 | J Vaughan | 2/1/2011 | Annual review |
| Dr Durga Vege | 2/13/2012 | J Vaughan, C Yuza | 2/1/2012 | Updated Branch Clinic information for new thermometers. |

Similar Procedures (Specify title of procedure and other OMC locations used if applicable)

## ..\POLICIES\TEMPERATURE & HUMIDITY MONITORING.doc

Attachments: