Clinical Laboratory

| SECTION: Equipment |  |
| --- | --- |
| DOCUMENT TYPE: Policy |  |
| TITLE: Temperature and Humidity Monitoring |  |
| APPROVED BY: Medical Director | DOCUMENT #: 13.0341.5 |
| WRITTEN BY: Judy Vaughan, MT(ASCP) | EFFECTIVE: 5/2/2011 |
| SCOPE: All Laboratories | COMPLEXITY: NA |

Purpose:

To describe the procedure for monitoring temperatures of Laboratory freezers, refrigerators, and incubators.

Policy:

Temperatures of all refrigerators, freezers, and incubators are monitored daily to ensure the integrity of stored supplies.

Supplies:

Liquid in Glass Thermometer

Min/Max thermometer

TempTrak monitoring devices

Slant tubes

Procedure:

- The temperature of all incubators, refrigerators, and freezers is monitored and recorded daily.

- The room temperature of critical testing / storage areas and humidity are monitored and recorded daily.

- Min/Max thermometers are located in refrigerators, freezers, and incubators in locations without 24/7 staffing or remote monitoring.

- Min/Max temperatures are recorded and reset daily.

- Specified SE Clinic and Hospital refrigerators and freezers are remotely monitored 24/7 by TempTrak.

- Slant tubes are located in freezers without 24/7 remote monitoring or staffing to detect thawing which may occur over weekends or holidays.

a.	Slant tubes are checked for evidence of thawing and/or refreezing.

b.	Thawing is indicated if the liquid is at a straight level versus the slant at which the tube was originally frozen.

c.	Slant tubes are documented as “OK” or “Unacceptable” on the designated log sheet.

- Unacceptable results must be resolved and corrective action documented.

- Acceptable temperature limits are listed on the log sheets.

- If the temperature is out of range, one or more of the following events must occur:

a.	Check contents for physical damage (broken bottles, discoloration, thawed contents, etc). If present, document and contact supervisor for evaluation or replacement instructions.

b.	Incubators:

- If the temperature is more than 2º C above or 5º C below the acceptable limits, immediately move the contents to another suitable unit.

- Adjust the temperature control as appropriate and recheck the temperature after 1 hour.

- If still out of range, contact Maintenance.

c.	Refrigerators:

- If the temperature is 0º C or lower, immediately remove the contents and place in another suitable unit.

- Recheck temperature after 1 hour. If still out of range, adjust temperature control and check again in 1 hour.

- If still out of range, contact Maintenance.

d.	Others:

- If the temperature is more than 5º C above or below the acceptable limit, immediately move the contents to another unit.

- Recheck temperature after 1 hour. If still out of range, adjust temperature control and check again in 1 hour.

- If still out of range, contact Maintenance.

e.	Frost-Free Freezer:

- Re-read the temperature every hour (up to 4 hours) before attempting to adjust the temperature control.

f.	TempTrak remote monitoring: reset alarm and document corrective action on-line. See TempTrak

procedure for more information.

g.	Document all corrective action on appropriate log or on-line (TempTrak).

- Document length of time temperature was out of range.

- Any adverse effects to reagents, controls, or calibrators stored in the unit.

- Corrective action and follow-up.

h.	All affected reagents, controls, calibrators, etc. must be validated before use using previously run patient samples, non-affected calibrator or control materials, or correlations with a reference laboratory.

- Document all results.

- If the humidity is out of range contact maintenance for possible mechanical adjustments.

  - If humidity is below range perform manual dip testing as necessary for Urinalysis.

  - Delay RPR testing until humidity is with acceptable range.

- All thermometers are checked against a NIST-certified thermometer before initial use and at least semi-annually.

References:

Saunders, Tietz Textbook of Clinical Chemistry and Molecular Diagnostics, 4th Edition, 2006

TempTrak Operator’s Manual

CAP Standards:

CHM.24500		NIST-certified thermometer available

CHM.24600		Check before initial use

CHM.24700		Temperature checks

CHM.24900		Corrective action

COLA Standards:

MA 6			Humidity monitoring and correction actions

MA 7 			Temperature Monitoring- Refrigerators

MA 8 			Temperature Monitoring- Freezers

MA 9 			Temperature Monitoring- Room Temperature

MA 10			Temperature Monitoring- Incubators

MA 13 			Temperature Monitoring- temperature dependant equipment

MA 14			Acceptable temperature range limits

MA 16			Documentation of corrective actions regarding temperature and humidity

| # Approved By | # Date | # Reviewed/Revised By | # Date | # Revision |
| --- | --- | --- | --- | --- |
| Dr Bradley Karon | 1/2007 | J. Vaughan | 1/2007 | New format |
| Dr Durga Vege | 1/2008 | J Vaughan | 1/2008 | Added humidity |
| Dr Durga Vege | 1/2009 | J Vaughan | 1/2009 | Annual Review added humidity in title, Step # 7 and COLA standards |
| Dr Durga Vege | 1/4/2010 | J Vaughan | 12/30/2009 | Updated format; added TempTrak info; edited #7 |
| Dr Durga Vege | 5/11/2010 | J Vaughan | 4/28/2010 | Updated g |
| Dr Durga Vege | 5/11/2011 | J Vaughan | 5/2/2011 | Updated #2 |
| Dr Durga Vege | 5/24/2012 | J Vaughan | 5/1/2012 | Annual review |

Similar Procedures (Specify title of procedure and other OMC location(s) used if applicable)

..\MAINTENANCE\THERMOMETER CHECKS.doc

..\PROCEDURES\Temptrak Temperature Monitoring.doc

Attachments:

..\FORMS & LOGS\LOG-DAILY TEMPERATURES-SE CLINIC.xls

..\FORMS & LOGS\LOG-TEMP ERRORS-SE CLINIC.doc

Temperature Log- Branch Clinic 13.0641-c.1