Clinical Laboratory

| Section: Equipment |  |
| --- | --- |
| Document Type: Procedure |  |
| Title: Vortex Specimen Mixer Usage and Preventative Maintenance |  |
| Approved by: Laboratory Administrative Director or Laboratory Supervisor | Document Number: 13.0342.1 |
| Written By: Charlotte Yuza | Effective Date: 1/3/2013 |
| Scope: Rochester NW and Branch Laboratories | Complexity: NA |

Purpose:

To outline the intended use and preventative maintenance of vortex specimen mixers.

Policy:

Properly trained staff with documented preventative maintenance and usage competence may perform this task.

- Intended use: mixing of quality control materials such as Hematology and ESR controls upon receipt to re-suspend RBCs from the bottom of the control vials.

- Optional use: re-suspension of urine sediment after centrifugation and removal of supernatant from the KOVA system.

Supplies:

Vortex mixer

Warnings and Precautions:

- Do not use tubes with a larger diameter that that of the vortexer cap (20mm).

- Do not start vortexing at maximum speed.

- Place the unit on a solid work surface or laboratory bench.

- Follow manufacturer’s handling instructions for quality control materials.

Procedure:

- For continuous vortexing mode:

  - Push the switch into position (switch on the left front side).

  - Set the speed dial on the right front side to the set desired intensity.

  - Control the intensity of vortexing by varying the pressure of a tube in the mixing cup.

- For intermittent vortexing mode:

  - Push the switch into position.

  - Vortexer will be activated when touching the vortexing cup with a test tube.

  - Set a desired level of vortexing using the speed control dial.

  - Control the intensity by varying the pressure of a tube in the mixing cup.

- Clean the unit only with a damp wet cloth; do not use chemical cleaning agents.

References:

Tap Dancer 270100 Operating Instructions, Feb 2007 Rev.0

| # Approved by | # Date | # Reviewed by / Revised by | # Date | # Revision |
| --- | --- | --- | --- | --- |
| J DeClercq | 1.4.13 | C Yuza | 1/3/2013 | New document |

Similar Procedures (Specify title of procedure and other OMC location(s) used if applicable)

Attachments: