Clinical Laboratory

| SECTION: Urinalysis |  |
| --- | --- |
| DOCUMENT TYPE: Procedure |  |
| TITLE: Routine Urinalysis |  |
| APPROVED BY: Medical Director | DOCUMENT #: 12.0432.15 |
| WRITTEN BY: Judy Vaughan, MT(ASCP) | EFFECTIVE: 3/1/2013 |
| SCOPE: All Laboratories | COMPLEXITY: Moderate |

Purpose:

To describe the procedure for performing routine urinalysis.

Clinical Utility:

Urinalysis results assist in the diagnosis of kidney function, urinary tract infections, carbohydrate metabolism (e.g. diabetes mellitus), and liver function. They also measure physical characteristics, including acid-base balance and urine concentration. Test results can be used with other diagnostic information to rule out certain disease states.

Policy:

Staff with documented training and competency may perform this task.

Principle:

Reagent strips for urinalysis contain reagent test pads that react with protein, blood, leukocytes, nitrite, glucose, ketone, pH, specific gravity, bilirubin, and urobilinogen. The strips may be read visually or instrumentally.

Reagents and Materials:

1.	Kova System tubes, caps, pipettes, and slides

2.	Biorad Quantify Plus Controls (see package insert for additional information)

- Initial and date the control bottles when opened.

- Store control solutions at 2-8C.

- Controls are stable for 31 days stored tightly capped at 2-25C.

- Allow controls to come to room temperature prior to testing.

- Mix gently by inversion; avoid foaming.

- After use, promptly recap and return to refrigerator.

3.	Siemens Multistix 10SG

- Store at room temperature, 15 - 30C.

- Initial and date the reagent bottle when opened.

- Do not remove the strip from the bottle until immediately before use.  Replace cap immediately and tightly after removing the strip.

- Avoid touching the test areas (pads) of the reagent strip.

- Refer to package insert for chemical principles of procedures and strip ingredients and performance characteristics.

4.	BD Vacutainer Urine Collection tubes

5.	2% Acetic Acid

6.	Clinitest tablets

Warnings and Precautions:

Use of Universal Precautions is required.

| Specimen Requirements and Storage |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Parameter | Min amt required | Patient Prep | Acceptable specimen types | Stability Room temp | Stability Refrig temp | Stability Frozen temp | Reject due to: |
| UADIP UA-R UA7 | 1 mL (12 mL preferred) | Clean Catch protocol; Catheter | Random Urine, Catheter, Pedibag | 1 hour (30 minutes if culture also ordered) | 8 hours | Not acceptable | Improper collection container, fecal contamination, improper specimen handling, urine preservatives** |

**Important: Some urine specimens may have been collected during a critical procedure or by means of an invasive procedure; therefore, it is important to never dispose of an unacceptable specimen until the caregiver has been notified. Follow rejected specimen protocol.

SPECIMEN VOLUMES:

| Preferred Specimen Volume | 12 mL |
| --- | --- |
| Minimum Specimen Volume | 1 mL |
| If less than 12 mL urine | Enter comment: “Microscopic exam performed on less than 12 mL of urine.” |
| If less than 2 mL urine | Do not centrifuge before performing microscopic exam. Enter comment: “Microscopic exam performed on uncentrifuged specimen.” |

Quality Control:

- Biorad Quantify Plus Macroscopic and Microscopic controls are run each day of testing.

- SE Clinic: record microscopic results on log sheet.

- Biorad Quantify Plus Macroscopic controls are run when a new bottle of strips is opened.

- At least two patient correlations are performed when changing reagent strip lot numbers.

Procedure:

1.	Check the patient sample identification carefully.

2.	Label the Kova tube using 2 identifiers.

3.	Swirl specimen cup to evenly distribute the sediment.

4.	Transfer the urine specimen from the specimen cup filling the Kova tube to the 12 mL graduation.

5.	Note urine color and clarity.

6.	Using Multistix 10SG reagent strips, perform chemical analysis using the Clinitek analyzer or by visual observation. (See Clinitek instrument operating procedure for directions on automated readers)

a.	Completely immerse all reagent areas in fresh, well-mixed uncentrifuged urine.

b.	Remove reagent strip immediately.

c.	While removing strip, run the edge of the reagent strip against the side of the urine container to remove

excess urine.

d.	Compare each test pad to the corresponding row of color blocks on the bottle label or place strip in the

automated reagent strip reader.

e.	If read manually, hold the strip close to the color blocks and match carefully, reading test results at the

specified times using good light (See reagent strip bottle and/or package insert for specific read times).

7.	Perform Clinitest on children under 2 year of age (see Clinitest procedure).

8.	Clear, yellow urines on adults, with a negative dipstick do not require a microscopic exam except when the provider has specifically ordered one (See Microscopic Urinalysis Protocol).

9.	Centrifuge 12 mL of urine in the capped Kova tube at 2000 rpm for 5 minutes.

10.	Remove tube from centrifuge being careful not to disturb or dislodge the sediment.

11.	Insert Kova pipette into tube.

12.	Decant supernatant to a volume of 1 mL (12:1 ratio of urine to sediment).

13.	Mix sediment into a homogenous mixture by gentle agitation of the tube.

14.	Withdraw a small sample using the Kova pipette.

15.	Transfer sediment sample to a Kova slide.

16.	Examine microscopically.

Specimens for Possible Urine Culture: (See “Urine Collection Tube” procedure more information)

- At the Hospital, the primary sample from ER, Urgent Care and in-patients are retained for 24 hours.

- Retain a BD urine tube for culture and sensitivity (gray-topped) on the following samples (see Urine Culture and Sensitivity-BD Vacutainer Urine System” procedure for more information):

- Southeast and Branch Clinics

- Specimens that reflex a microscopic (UA-M)

- Pediatric specimens (<5 years)

- Cath specimens

- Hospital

- Pediatric specimens (<5 years)

- Cath specimens

- Outpatients (i.e. nursing home specimens)

- Label appropriately and store at room temperature. Stable for 48 hours.

- If there is insufficient volume of urine available to save for a possible urine culture order, enter canned, chartable comment under Misc test results, “QNS- no urine culture sample retained in lab for 48 hours”

- If microscopic results indicate the need for a culture but no sample is retained, contact the provider with this information.

Procedural Notes:

- This procedure follows manufacturer’s instructions without alteration.

- For automated reagent strip reading-Clinitek Urine analyzer: see specific procedure for operating instructions

3.	UA Reflex (UA-R): Microscopic testing performed under the following circumstances:

- Children under 5 years

- Catheter specimens

- Positive leukocyte esterase, blood, nitrites, and/or protein

- Presence of color that may interfere with dipstick readings

- Any clarity other than clear.

- Specific provider request

4.	Urinalysis-Dip (UADIP): No microscopic, dipstick only.

- Clinics: UA-M will reflex for positive dipstick results. Cancel UA-M on UADIP orders.

5.	Occupational Medicine Department UA7 order for Health Maintenance Exams (insurance, sports physicals, or pre-employment physicals): If there are any positive results for blood, nitrite, or leukocytes, urine microscopic testing will be reflexed in the LIS.

6.	Retain all urine samples for at least 24 hours (discard 2 days post-collection).

Result Interpretation:

1.	Correlate WBCs, RBCs, and casts found microscopically with the presence of leukocytes, blood, and/or

protein. Check patient and sample identification carefully. Reconcile discrepancies before reporting results. Consult package inserts for interfering substances.

- Normal sediment is not free of cells or casts, but contains a limited number of formed elements. The presence of one or two red or white cells/hpf or a few epithelial cells or hyaline casts is not necessarily abnormal.

- Notify department whenever WBC or RBC casts or other unexpected results are reported.

Limitations:

1.	For interfering substances on Multistix 10 SG, see package insert.

2.	All urines are observed for abnormal color and other specimen limitations causing erroneous results on the dipstick.

3.	There may be instances when one or more of the chemistry test strip results cannot be reported due to

interfering substances and/or color (i.e. pyridium).

- Enter “*Cannot report” for urine macroscopic results with @U52 comment: “Cannot report dipstick results due to chemical interference”.

- Report microscopic results only.

- 4+ (packed field) red cells and/or differentiation of red cells from yeast:

- To “clear” red cells to allow for identification and quantification of other cellular elements, add 1-2 drops of 2% acetic acid to the urine sediment. RBCs will lyse but WBCs and yeast will stay intact.

- Scan whole slide prior to the addition of acetic acid, otherwise structures such as red cell casts or certain crystals may dissolve.

- Note: Add 0.5 ml 50% acetic acid to 12.5 mls water to make 2% acetic acid solution. Label appropriately with 30-day outdate.

- If humidity is outside of acceptable range (20-80%), perform manual reading of dipstick.

Reporting Results:

Macroscopic:

Urine Type

Clean catch Foley indwelling Cath Foley out Supra Pediatric bag Non-specified Collection hat DNR Other-see below Foley insert

| Color | Clarity | Glucose (mg/dL) | Bilirubin | Ketones (mg/dL) | Nitrites | Leukocyte | Specific Gravity | Blood | Protein (mg/dL) | Uro-bilinogen (EU/dL) | pH |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Colorless | Clear | Negative | Negative | Negative | Negative | Negative | 1.005 | Negative | Negative | 0.2 | 5.0 |
| Straw | Hazy | 100/Trace | Trace | Trace | Positive | Trace | 1.010 | Trace | Trace | 1.0 | 5.5 |
| Yellow | Cloudy | 250 | Small | 15 |  | Small | 1.015 | Small | 30 | 2.0 | 6.0 |
| Amber | Turbid | 500 | Moderate | 40 |  | Moderate | 1.020 | Moderate | 100 | 4.0 | 6.5 |
| Orange |  | >=1000 | Large | >=80 |  | Large | 1.025 | Large | >=300 | >=8.0 | 7.0 |
| Red |  |  |  |  |  |  | >=1.030 |  |  |  | 7.5 |
| Green |  |  |  |  |  |  |  |  |  |  | 8.0 |
| Black |  |  |  |  |  |  |  |  |  |  | 8.5 |
| Brown |  |  |  |  |  |  |  |  |  |  | >=9 |

Microscopic Examination:

Examine 10-12 fields utilizing a battlement track to review and report average number of cellular components.

| Element | Reporting Field | Report Format |
| --- | --- | --- |
| Crystals* | HPF (40X) High Power Field | Enter Number and Type of crystal identified Rare 1+: Few in every microscopic field. 2+: Many in every microscopic field. 3+: Each microscopic field is almost packed. 4+: Packed field. Select “See Below” and open comment box to result. Canned comments available (F5) |
| Bacteria | HPF (40X) High Power Field | Few: Present in some but not all microscopic fields. None Seen 1+: Few in every microscopic field. 2+: Many in every microscopic field. 3+: Each microscopic field is almost packed. 4+: Packed field. |
| Casts | LPF (10X) Low Power Field | None Seen See Below: Open Result Comment box to enter number and type of cast identified using Canned Comments (F5) choices. Number: Rare 0-2 3-5 6-10 11-20 21-50 51-100 >100 Packed Field Type: Hyaline, Fine Granular, Coarse Granular, Red Blood Cell, White Blood Cell, Epithelial, Waxy, Mixed Cellular |
| Element | Reporting Field | Report Format |
| Renal Tubular Cells (URTUB) | HPF (40X) High Power Field | Number (choose from keypad): None Seen 0-1 2-10 11-20 21-50 51-100 >100 Packed Field Renal Tubular Cell Result Comment (@U10) attached to result: “Cells suggestive of renal tubular cells present.” SE Clinic: Have second tech verify the presence of renal tubular cells. |
| Transitional Epithelial Cells (UTEPI) | HPF (40X) High Power Field | Number (choose from keypad): None Seen 0-3 4-10 11-20 21-50 51-100 >100 Packed Field |
| Squamous Epithelial Cells (USEPI) | LPF (10X) Low Power Field | None Seen Few: Present in some but not all microscopic fields. 1+: Few in every microscopic field. 2+: Many in every microscopic field. 3+: Each microscopic field is almost packed. 4+: Packed field. |
| WBCs & RBCs | HPF (40X) High Power Field | Number (choose from keypad): None Seen 0-3 4-10 11-20 21-50 51-100 >100 Packed Field |
| Oval Fat Bodies* | LPF (10X) Low Power Field | Rare 1+: Few in every microscopic field. 2+: Many in every microscopic field. 3+: Each microscopic field is almost packed. 4+: Packed field. |
| Yeast*, spermatozoa*, trichomonas* etc | HPF (40X) High Power Field | Report as present. |
| Mucous |  | Do not report |
| >100 cells |  | Full field |
| Nothing else visible |  | Packed field |

*To report Misc (crystals, yeast, trichomonas, etc):

Enter under Misc. Select “See Below” Open comment field. Enter number and type of elements seen F5 for canned comments Select DNR if none seen.

2.	Crystals:

| Crystals-Acid Urine | Crystals-Alkaline Urine | Crystals-Acid, Neutral, or Slightly Alkaline Urine | Crystals- Alkaline or Neutral |
| --- | --- | --- | --- |
| - Amorphous urates. - Uric acid. - Sodium urate. - Bilirubin. - Cystine. - Cholesterol. - Leucine. - Tyrosine. | - Calcium carbonate. - Ammonium biurate. - Amorphous phosphate. - Calcium phosphate. | - Calcium oxalate. - Hippuric acid. | - Triple phosphate. |

- Uric acid crystals are soluble with heat (60°C). Warm Kova tube in warm tap water to clear specimen.

- Amorphous phosphate, calcium phosphate, and triple phosphate crystals are soluble in acetic acid.

Reporting Template:

| Hospital | HUA |  | Chatfield | CHAT |  | Plainview | PLAIN |  | Spring Valley | SV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SE Clinic | GUA |  | NW | NW |  | Preston | PSTN |  | Stewartville | STEW |
| Byron | BYRON |  | Pine lsland | PINE |  | St Charles | SC |  | Wanamingo | WNMG |
| FastCare N | FCN |  | FastCare S | FCS |  | Skyway | SKY |  |  |  |

Expected Values:

- Color: Colorless/Straw/Yellow

- Clarity: Clear

- Glucose, Bilirubin, Ketones, Blood, Nitrite, Leukocytes: Negative

- Protein: Negative-Trace

- Specific Gravity: 1.000-1.035 (>1.030)

- pH: 5.0-7.0

- Urobilinogen: 0.2-1.0 E.U./dL

- RBCs: <3/hpf (male & female)

- WBCs: <3/hpf (male)  <10/hpf (female)

- Casts: 0-2 hyaline

- Crystals: None seen

- Squamous Epithelial Cells: Few squamous/lpf

- Transitional Epithelial Cells: 0-3 transitional/hpf

- Renal Tubular Cells: 0-1 renal tubular/hpf

- Bacteria: None seen

Alternate Method / Referral of Specimens:

Perform visual reading of dipstick if automated method unavailable.

References:

Urinalysis and Body Fluids, Mosby, 1995

A Handbook of Routine Urinalysis, Lippincott, 1983

Wet Urinalysis, ASCP, 2003

Biorad Quantify Plus Controls Package Insert, 3/2012

Siemens Multistix 10SG Reagent Strip Package Insert 6/2010

CAP/COLA Standards:

URN.22300		Urine preservation

URN.23100		Reference intervals

URN.30200		Dipstick UA

QC10			QC manufacturers instructions

QC15			Quantitative QC

APM11			Directions for microscopic testing

| # Approved By | # Date | # Reviewed/Revised By | # Date | # Revision |
| --- | --- | --- | --- | --- |
| Dr Bradley Karon | 3/2007 | J Vaughan | 3/2007 | New format; new controls |
| Dr Bradley Karon | 5/2007 | J Vaughan | 5/2007 | Change to Kova System & new controls |
| Dr Bradley Karon | 7/2007 | J Vaughan | 7/2007 | Added LIS entry of microscopic results and color/clarity values. Updated reference range for WBCs |
| Dr Durga Vege | 2/2008 | J Vaughan | 2/2008 | Updated pH reporting & CAP standards |
| Dr Durga Vege | 5/2008 | J Vaughan | 5/2008 | Added NC for UA7 reflex UA-M |
| Dr Durga Vege | 7/2008 | J Vaughan | 7/2008 | Added black & brown UA color & trace bilirubin; Added pH reference range |
| Dr Durga Vege | 3/2009 | J Vaughan | 3/2009 | Removed NC for UA-M reflexed from UA7; added calling RBC/WBC casts |
| Dr Durga Vege | 4/28/2009 | J Vaughan | 4/28/2009 | Added specimens for possible urine culture information; added COLA standards |
| Dr Durga Vege | 6/8/2009 | A Petersen/J Vaughan | 6/1/09 | Updated format; Added information for hospital specimens retained for possible culture; Updated references; changed Bayer to Siemen; updated urine culture storage temp |
| Dr. Durga Vege | 8/3/2009 | J Vaughan | 7/27/2009 | Updated format; Added Clinical Utility; updated Limitations #4, Renal cell ID & Crystals under Microscopic |
| Dr. Durga Vege | 12/21/2009 | J Vaughan | 12/2/2009 | Updated QC; added Color, Clarity, Casts and Crystals to Expected Values; updated Result Interpretation #3; updated reporting of Casts and Misc, added Limitations #5 |
| Dr. Durga Vege | 3/16/2010 | J Vaughan | 3/11/2010 | Added Procedure Note #5; updated References |
| Dr Durga Vege | 1/28/2011 | J Vaughan | 1/14/2011 | Added Renal cell comment & urine types; Updated Expected Values; Updated CAP Standards |
| Dr Durga Vege | 3/14/2011 | J Vaughan | 3/7/2011 | Updated Transitional Epi and Renal Tubular reporting & Reference Ranges |
| Dr Durga Vege | 3/12/2012 | J Vaughan, C Yuza | 3/1/2012 | Updated Epi test codes, Microscopic Examination, Procedure Notes #4, & References |
| Dr Jeremy Gibson | 3/27/2013 | C Yuza / J Vaughan | 3/1/2013 | Added FastCares and Skyway to Reporting Template, and microscopic battlement review path; updated Specimen Requirements, Procedure #2 and References |

Similar Procedures (Specify title of procedure and other OMC location(s) used if applicable)

URINALYSIS TESTING- CLINITEK 100.doc

URINALYSIS CLINITEK STATUS.doc

..\MAINTENANCE\MAINTENANCE- CLINITEK STATUS OR CLINITEK 50.doc

..\POLICIES\MICROSCOPIC URINALYSIS POLICY.doc

CLINITEST- REDUCING SUGARS.doc

..\..\SPECIMEN COLLECTION & PROCESSING\PROCEDURES\URINE COLLECTION TUBES.doc

Attachments: