- Purpose

  - The ABL 825 analyzers are designed for in vitro diagnostic use on undiluted heparinized whole human blood. The following parameters are measured:

| pH/blood gas | pH (acidity) |
| --- | --- |
|  | pCO2 (carbon dioxide tension) |
|  | pO2 (oxygen tension) |
| Oximetry | ctHb (total hemoglobin concentration) |
|  | sO2 (oxygen concentration) |
|  | FCOHb (fraction of carboxyhemoglobin in total haemoglobin) |
|  | FMetHb (fraction of methemoglobin in total hemoglobin) |
| Electrolytes | cK+ (potassium ion concentration) |
|  | cNa+ (sodium ion concentration) |
|  | cCa2+ (calcium ion concentration) |
|  | cCl- (chloride ion concentration) |
| Metabolites | cGlu (D-glucose concentration) |
|  | cLac (L(+)-lactate concentration) |

  - Abnormal levels of these parameters may be the cause or consequence of a variety of disorders such as an acid-base imbalance, which may indicate respiratory, metabolic, or kidney disorders.

- Performance Specifications

  - Performance characteristics for individual parameters are documented in section 5 of the ABL 800 series Reference Manual

  - Traceability to the primary standards at Radiometer can be found in booklet AS117 in the front of the Reference Manual.

  - Measured Parameters:

| Parameter | Unit | Measuring Range |
| --- | --- | --- |
| pH | pH scale | 6.300-8.000 |
| pCO2 | mmHg | 5.0-250 |
| pO2 | mmHg | 0.0-800 |
| ctHb | g/L | 0.0-277 |
| sO2 | % | 0.0-100.0 |
| FCOHb | % | 0.0-100.0 |
| FMetHb | % | 0.00-100.0 |
| cK+ | mmol/L | 0.5-25.0 |
| cNa+ | mmol/L | 7-350 |
| cCa2+ | mmol/L | 0.20-9.99 |
| cCl- | mmol/L | 7-350 |
| cGlucose | mmol/L | 0.0-60 |
| cLactate | mmol/L | 0.0-30 |

    - The Measuring Range for a parameter is the range within which the analyzer is physically capable of measuring. The measuring range corresponds to the “range of indication” as defined in the “International vocabulary of basic and general terms in metrology” (VIM)

  - Derived Parameters: cHCO3-, Base Excess

Equations for derived parameters can be found in section 6 of the ABL 800 series Reference Manual.

  - Precision Limits for Duplicate Results

| pH | 0.02 |
| --- | --- |
| pCO2 | 4 mmol/L |
| pO2 | 6 mmol/L |

- Specimen Type

  - Preparation Prior to Sampling

    - Correct sampling device should be prepared before sampling

    - Arterial and Venous blood gas samples: Heparinized blood gas syringe with separate cap for capping after sampling.

    - Capillary samples: puncture device such as lancet , capillary tube(s) and caps

    - Patients are informed and prepared in order to avoid unnecessary anxiety which could cause hyperventilation

  - Sampling

    - Arterial

      - Blood Gas Syringe- Arterial blood gas samples are drawn by the Physician, Respiratory Technologist or trained nursing staff.

    - Venous

      - Blood Gas Syringe- Venous blood gas samples are usually drawn by the Physician, Respiratory Technologist or trained nursing staff but can also be collected by trained MLA or MLT.

    - Capillary

      - A minimum of one full 100 uL capillary tube is required for testing.

    - Ionized Calcium

      - Whole Blood in blood gas syringe.

      - Serum or plasma

  - Specimen Transport

    - Whenever possible all samples should be transported to the Laboratory immediately.

    - Arterial samples drawn on Nursing units are transported on ice in zipper lock bags.

    - Capillary samples are transported in ice water.

  - Specimen Storage and Stability

    - Blood Gas

      - Storage time should be kept to a minimum- preferably less than 10 minutes at room temperature- to obtain the most accurate results.

      - Plastic syringes should not be iced, but kept at room temperature as long as the blood is analyzed within 30 minutes of collection. Oxygen and carbon dioxide levels in blood kept at room temperature for 30 minutes or less are minimally affected.

      - If a prolonged time delay before analysis is anticipated (more than 30 minutes), samples should be stored in ice water.

      - Storage time in ice water should not exceed 1 hour.

    - Ionized Calcium

      - Specimens stored centrifuged and unopen at 4C are stable up to 70 hours.

- Equipment and Reagents

  - Equipment: ABL 825 FLEX

  - Reagents and Consumables

    - Refer to individual packaging for expiration dating. Do not use any reagent or consumable past the manufacturer’s expiration date.

    - Electrodes and Membranes

| Electrode | Corresponding Membrane (4 units/box) |
| --- | --- |
| Reference Electrode | Reference Membrane |
| pH Electrode | NA |
| pO2 Electrode | pO2 Membrane |
| pCO2 Electrode | pCO2 Membrane |
| K Electrode | K Membrane |
| Na Electrode | Na Membrane |
| Cl Electrode | Cl Membrane |
| Ca Electrode | Ca Membrane |
| Glucose Electrode | Glucose Membrane |
| Lactate Electrode | Lactate Membrane |
| Dummy Electrode | NA |

    - Storage

      - Glucose Electrode: 2-25 oC

      - Lactate Electrode: 2-10 oC

      - All other Electrodes: 2-32 oC

    - Stability

      - Expiration date and Lot No. are printed on a label on the side of the membrane box.

  - Reagents

    - Reagent Storage and Stability

| Reagent | Storage | Stability |
| --- | --- | --- |
| Cal1 Solution | 2-32oC | Expiration date printed on label Stability in use: 2 weeks |
| Cal2 Solution | 2-32 oC | Expiration date printed on label Stability in use: 2 weeks |
| S7770 tHb Cal Solution | 2-25 oC In a dark place | After opening the solution must be used at once |
| S4980 Rinse Solution | 2-32 oC | Expiration date printed on label |
| S8370 Cleaning Solution | 2-32 oC | Expiration date printed on label |
| S5362 Hypochlorite Sol’n | 2-8 oC In a dark place | Expiration date printed on label, Store tightly capped to avoid contamination |
| Gas 1 | 2-32 oC | Stable for 25 months from date of filling |
| Gas 2 | 2-32 oC | Stable for 25 months from date of filling |

  - Other Materials Required

    - Absorbent pads

    - Magnet

- Calibration: see separate procedure IHL-CHM-ABL825-Calibration

- Procedure

  - Blood Gas Syringe

    - Thoroughly mix the sample by inverting the syringe several times and rolling it between palms of your hands.

    - Check that the analyzer is in Ready mode.

    - Remove the syringe cap and dispel a drop of blood onto an absorbent pad. This will allow for visual inspection for presence of clots and to remove blood from the neck of the syringe which has not been properly mixed. If clots are present, request a new sample.

    - Lift the syringe inlet flap and place the syringe onto the sampling port.

    - Select the syringe sampling mode on the screen. 195 uL is the default selection to perform all tests. If necessary, touch a different volume mode.

    - Touch Start button to begin aspiration of the sample.

    - When prompted by analyzer, remove the syringe and close the inlet flap.

    - Scan the patient’s barcode when the Patient Identification screen is displayed. The information can be manually entered via the numeric key pad displayed and the alpha keyboard accessed by pressing the keyboard button.

  - Capillary

    - Thoroughly mix the sample by moving the mixing wire up and down the full length of the capillary using a magnet.

    - Check that the analyzer is in Ready mode.

    - Do NOT remove the mixing wire before the sample is aspirated. Slide the mixing wire to the end of the capillary opposite from the end to be aspirated.

    - Remove both capillary end caps.

    - Lift the capillary inlet flap and insert the capillary into the inlet.

    - Select the desired mode by touching the appropriate button.

| Capillary 85 uL | Blood gas |
| --- | --- |
| Capillary 35 uL | pH only |
| Capillary 35 uL | Oxi only |

    - Touch Start.

    - When prompted by analyzer, remove the capillary and close the inlet flap.

    - Scan the patient’s barcode when the Patient Identification screen is displayed. The information can be manually entered via the numeric key pad displayed and the alpha keyboard accessed by pressing the keyboard button.

- Quality Control: see separate procedure IHL-CHM-ABL825-Quality Control

- Interferences /Cross-Reactions:

| Substance | Interference |
| --- | --- |
| Halothane (anesthetic) | May give unreliable pO2 results |
| Lipid therapy, protamine sulphate (treatment) | In OXI measurements. After measurement on blood from a patient who has received lipid therapy or protamine sulphate it may be necessary to clean the analyzer using the Cleaning program. |
| Methylene Blue, HiCN (medication) | In OXI measurements. |
| Anions:Br-, I-, S2-and ClO4- (drugs) | Erroneously high cCl- |
| Anticoagulants (sampling) | Radiometer recommends only the use of heparin as anticoagulant. Other anticoagulants can cause erroneous electrolyte, calcium, glucose or lactate results. |
| Thiocyanic acid (degradation product from treatment with nitroprussid. Also produced in thiosulphate treatment of cyanide poisoning). | Erroneously high glucose and lactate measurements. |
| Glycolic acid (ethylieneglycol degradation product) | Erroneously high lactate measurements. |
| Insufficiently stabilized blood. Caustic fluids(e.g. strong acids or bases, detergents, etc.). Fluids that precipitate. Fluids that affect the sensor enzymes. Fluids that form complexes with the analyzer solutions (calcium). High viscosity fluids. Hydrophobic fluids. Reactive fluids. | Other fluids mode allows measurement on fluids other than heparinised human blood. Warning/caution: Be aware that some substances, such as listed in the left column, measured in the Other Fluids mode may damage the analyzer or the electrodes. This can affect the subsequent measurement of patient samples or quality controls. |
| Galactose, glucosamine, maltose, mannose, xylose | Erroneously high glucose measurements. |

- Calculating Results: Not applicable

- Reference Intervals

| Test | Age | Reference Interval |
| --- | --- | --- |
| Arterial, Neonate |  |  |
| pH | <1 day | 7.33-7.49 |
|  | >/=1 day-<2 days | 7.25-7.49 |
|  | >/=2 days | 7.35-7.45 |
| pCO2 | <2 days | 26-41 mmHg |
|  | >/=2 days | 35-45 mmHg |
| pO2 | <2 days | 65-76 mmHg |
|  | >/=2 days | 70-104 mmHg |
| HCO3- | <2 days | 16-25 mmol/L |
|  | >/=2 days | 21-28 mmol/L |
| Base Excess |  | -2 to +2 |
| Arterial, >1 year |  |  |
| pH | >1 year | 7.35-7.45 |
| pCO2 |  | 35-45 mmHg |
| pO2 |  | 80-100 mmHg |
| HCO3- |  | 22-28 mmol/L |
| O2 Sat |  | 95-98 % |
| Base Excess |  | -3.0 to +3.0 |
| Arterial, Cord |  |  |
| pH | Newborn | 7.22-7.41 |
| pCO2 |  | 32-64 mmHg |
| pO2 |  | 14-22 mmHg |
| Base Excess |  | -7.8 to -2.2 |
| Venous, Cord |  |  |
| pH | Newborn | 7.29-7.41 |
| pCO2 |  | 25-53 mmHg |
| pO2 |  | 23-35 mmHg |
| Base Excess |  | -6.2 to -1.8 |
| Venous |  |  |
| pH |  | 7.31-7.41 |
| pCO2 |  | 40-52 mmHg |
| pO2 |  | 30-50 mmHg |
| HCO3- |  | 22-28 mmol/L |
| O2 Sat |  | 60-85 % |
| Base Excess |  | -3 to +3 |
| Hemoglobin | </= 6 days | 150-220 g/L |
|  | 6-28 days | 140-200 g/L |
|  | 29-58 days | 115-180 g/L |
|  | 59-88 days | 90-135 g/L |
|  | 89 days- <1 year | 100-140g/L |
|  | >/= 1-<5 years | 110-140g/L |
|  | >/= 5-<14 years | 120-160g/L |
|  | Female >/=14 years | 115-150g/L |
|  | Male >/=14 years | 135-165g/L |
| Ionized Calcium | </= 3 days | 1.30-1.60 mmol/L |
|  | 4-30 days | 1.20-1.48 mmol/L |
|  | 31 days-13 years | 1.20-1.38 mmol/L |
|  | >/= 13 years | 1.15-1.29 mmol/L |
| Sodium | </=7 days | 133-146 mmol/L |
|  | 8-30 days | 134-144 mmol/L |
|  | 31-179 days | 134-142 mmol/L |
|  | 180 days to <1 year | 133-142 mmol/L |
|  | </=1 to <15 years | 134-143 mmol/L |
|  | >/= 15 years | 135-145 mmol/L |
| Potassium | <30 days | 3.2-6.0 mmol/L |
|  | 30 days to <1 year | 3.5-6.0 mmol/L |
|  | >/=1 to <6 years | 3.5-5.0 mmol/L |
|  | >/=6 years | 3.5-5.1 mmol/L |
| Chloride | </=7 days | 96-111 mmol/L |
|  | 8-179 days | 96-110 mmol/L |
|  | >/=180 days to <1 year | 96-108 mmol/L |
|  | >/= 1 to <20 years | 96-109 mmol/L |
|  | >/= 20 years | 95-107 mmol/L |
| CO2 | </=7 days | 17-26 mmol/L |
|  | 8-30 days | 17-27 mmol/L |
|  | 31-179 days | 17-29 mmol/L |
|  | 180 days to</= 1 year | 18-29 mmol/L |
|  | 1-</= 19 years | 20-31 mmol/L |
|  | >/= 20 years | 21-32 mmol/L |
| Glucose- Random | </= 1 day | 2.2-5.0 mmol/L |
|  | >1 to</= 29 days | 2.7-5.0 mmol/L |
|  | >/=30 days to <1 year | 3.3-5.6 mmol/L |
|  | >/=1 year | 3.6-8.0 mmol/L |
| Lactate | </=61 days | 1.0-3.5 mmol/L |
|  | 62 days- </=1 year | 1.0-3.3 mmol/L |
|  | > 1 to </= 19 years | 1.0-2.4 mmol/L |
|  | >/= 20 years | 0.4-2.0 mmol/L |
| Capillary |  |  |
| pH | </=1 day | 7.33-7.49 |
|  | 2 days | 7.25-7.43 |
|  | 3-30 days | 7.32-7.43 |
|  | >/=31 days | 7.34-7.43 |
| pCO2 | </=60 days | 27-40 mmHg |
|  | >/=61 days | 26-41 mmHg |
| pO2 | </=7 days | 60-75 mmHg |
|  | >/=8 days | 80-100 mmHg |
| HCO3- |  | 16-24 mmol/L |
| Base Excess | </=60 days | -10 to -2.0 |
|  | >/= | -6.6 to -0.2 |
| O2 Sat | </=1 day | 30-80 % |
|  | >/=2 days | 60-85 % |

- Critical Intervals

|  | Low Critical | High Critical |
| --- | --- | --- |
| ctHb</=29 days | 50 g/L | 250 g/L |
| ctHb >29 days | 50 g/L | 200 g/L |
| cK+ </=1 year | 3.0 mmol/L | 8.0 mmol/L |
| cK+ >1 year | 2.8 mmol/L | 6.2 mmol/L |
| cNa+ | 120 mmol/L | 160 mmol/L |
| cCa2+</=30 days | 1.00 mmol/L | 1.60 mmol/L |
| cCa2+>30 days | 1.00 mmol/L | 1.45 mmol/L |
| cGlucose </=1 year | 1.80 mmol/L | 17.00 mmol/L |
| cGlucose >1 year | 2.20 mmol/L | 25.00 mmol/L |

- Laboratory Interpretation Results

  - Examine the analyzer printout

| ? | Indicates an error in the last quality control or calibration, or due to sample problems |
| --- | --- |
|  or  | Indicates the parameter is outside the reference range but inside critical limits |
| Double or  | Indicates the parameter is outside the upper or lower critical limit, but inside the measuring range. Call results exceeding the critical limit. |
| Triple  or  | Indicates the parameter is outside the measuring range. Question nursing regarding possible IV contamination. |
| ..... | Indicates the parameter cannot be calculated or exceeds the numeric limit of the analyzer. |

  - Duplicate Testing

    - At the start of each shift a minimum of 2 patient samples shall be measured in duplicate on two different analyzers when available.

    - Further patient samples may be analyzed only once unless the technologist deems it prudent to measure a sample in duplicate.

      - Internally inconsistent (e.g. sum of pO2 + pCO2 > 150 mmHg for patients breathing room air) or

      - Extremes of the range of expected values (outside Critical limits)

  - Check duplicate test results between instruments for precision.

    - Precision limits:

| pH | 0.02 |
| --- | --- |
| pCO2 | 4 mmol/L |
| pO2 | 6 mmol/L |

    - If results are acceptable ,report the results through the LIS

    - If precision replication is not present, repeat the testing on the same sample.

  - Reporting

    - Comments, such as the quality of the specimen as received, transportation delays, improper labelling of an irreplaceable specimen must be documented in the LIS before accepting results.

    - Results are accepted in I-Net and automatically reported through Ultra.

- Safety Precautions:

Blood sampling should be performed by authorized personnel only. Always handle blood and the collection devices with care. Avoid direct contact with the sample by using gloves. Sterile techniques are required at all times to avoid infection.

- Potential Sources of Variability

Always follow the sampling procedures described in Specimen Collection Guidelines and in section 12 of the ABL 800 series Operator’s Manual. Failure to follow these procedures may introduce clots or air bubbles in the sample and yield erroneous results.

- References

ABL 800 series Operator’s Manual, Radiometer Medical ApS., Denmark 2008

ABL 800 series Reference Manual, Radiometer Medical ApS., Denmark 2008

NCCLS Publication C46-A, Blood Gas and pH Analysis and Related Measurements; Approved Guideline, NCCLS, USA 2001

NCCLS Publication C31-A2, Ionized Calcium Determinations: Precollection Variables, Specimen Choice, Collection, and Handling; Approved Guideline- Second Edition