Principle of the Test

The ADVIA Centaur PSA assay is a two-site sandwich immunoassay using direct chemiluminometric technology, which uses constant amounts of two antibodies. The first antibody, in the Lite Reagent, is a polyclonal goat anti-PSA antibody labeled with acridinium ester. The second antibody, in the Solid Phase, is a monoclonal mouse anti-PSA antibody, which is covalently coupled to paramagnetic particles.

The system automatically performs the following steps:

- dispenses 35 µL of sample into a cuvette

- dispenses 250 µL of Solid Phase and 100 µL of Lite Reagent and incubates for 7.5 minutes at 37°C

- separates, aspirates, and washes the cuvettes with reagent water

- dispenses 300 µL each of Acid Reagent and Base Reagent to initiate the chemiluminescent reaction

- reports results according to the selected option, as described in the system operating instructions or in the online help system

A direct relationship exists between the amount of PSA present in the patient sample and the amount of relative light units (RLUs) detected by the system.

Clinical Application and Usefulness

For in vitro diagnostic use in the quantitative measurement of prostate-specific antigen (PSA) in human serum using the ADVIA Centaur and ADVIA Centaur XP systems. This assay is indicated for the measurement of serum PSA in conjunction with Digital Rectal Exam (DRE) as an aid in the detection of prostate cancer in men aged 50 years and older. This assay is further indicated as an aid in the management (monitoring) of patients with prostate cancer.

PREANALYTIC PROCESS

Specimen Collection and Handling

Specimen Collection

|  |  | BIOHAZARD All products or objects that come in contact with human or animal body fluids should be handled, before and after cleaning, as if capable of transmitting infectious diseases. Wear facial protection, gloves, and protective clothing. |
| --- | --- | --- |

- Serum is the recommended sample type for this assay.

- This assay requires 35 µL of sample for a single determination. Additional volume is required for onboard dilutions.

- Before placing samples on the system, ensure that samples are free of fibrin or other particulate matter and that samples are free of bubbles. Remove particulates by filtration or centrifugation at 1000 x g for 15 to 20 minutes.

Specimen Storage and Stability

- Allow samples to clot adequately before centrifugation.

- Keep tubes stoppered and upright at all times.

- Do not use samples that have been stored at room temperature for longer than 8 hours.

- Tightly cap and refrigerate specimens at 2 to 8°C if the assay is not completed within 8 hours.

- Freeze samples at or below -20°C if the sample is not assayed within 48 hours.

- Freeze samples only once and mix thoroughly after thawing.

- Specimen Rejection Criteria

- See IHL-OPS-III-Unacceptable Specimen Protocol

- Specimen Referral Criteria

- Specimens can be referred out if reagents are unavailable or on the decision of the Quality Management Officer.

REAGENTS

Storage and Stability

- Store the reagents upright at 2 – 8°C.

- Lite Reagent and Solid Phase stable until the expiration date on the pack label, or for 28 days onboard the system.

- Multi-Diluent 2 stable until the expiration date on the pack label, or for 28 consecutive days after accessing the ancillary reagent pack.

CAUTION:

- Discard the primary reagent packs at the end of the onboard stability interval.

- Do not use reagents beyond the expiration date.

Ingredients

Reagent ingredients for the ADVIA Centaur PSA assay are as follows:

| Reagent | Volume | Ingredients |
| --- | --- | --- |
| Lite Reagent | 10.0 mL/reagent pack | polyclonal goat anti-PSA antibody (~77 ng/mL) labeled with acridinium ester in buffered saline with bovine serum albumin and preservatives |
| Solid Phase | 25.0 mL/reagent pack | monoclonal mouse anti-PSA antibody (~25 µg/mL) covalently coupled to paramagnetic particles in buffered saline with bovine serum albumin and preservatives |
| Multi-Diluent 2 | 10.0 mL/reagent pack | goat serum with sodium azide (0.1%) and preservatives |

CAUTION: This device contains material of animal origin and should be handled as a potential carrier and transmitter of disease.

NOTE: Sodium azide can react with copper and lead plumbing to form explosive metal azides. On disposal, flush reagents with a large volume of water to prevent the buildup of azides, if disposal into a drain is in compliance with federal, state, and local requirements.

Reagents Special Preparation

No special preparation of reagents is required.

Calibration

For detailed information about scheduling a calibration, refer to Scheduling Calibrators.

Two-point Calibration Interval

Use Calibrator Q to perform two-point calibrations.

Perform a two-point calibration every 28 days and:

- when changing lot numbers of primary reagent packs

- when replacing system components

- when quality control results are repeatedly out of range

Defining Calibrator Values for Two-point Calibration

Use the barcode reader to enter the calibrator values from the Calibrator Assigned Value card onto the system.

- At the workspace, select Calibration.

- Select Calibrator Definition.

- Select Scan Data.

- Use the handheld barcode reader to scan the barcodes (from top to bottom) on the Calibrator Assigned Value card.

- Ensure that the calibrator values are correct. After you select Save, you cannot add or delete a test from a calibrator definition.

- Select Save.

Defining the Master Curve

The ADVIA Centaur PSA assay requires a Master Curve calibration when using a new lot number of Lite Reagent and Solid Phase. Use the barcode reader to enter the Master Curve values from the Master Curve card onto the system.

- At the workspace, select Calibration.

- Select Master Curve Definition.

- Select Scan Data.

- Use the handheld barcode reader to scan the barcodes on the Master Curve card.

IMPORTANT: Ensure you are scanning the side of the Master Curve card labeled ADVIA Centaur.

NOTE: The table on the Master Curve card represents the fields as they are displayed in the ADVIA Centaur System only. Use the illustrations in the Using the Master Curve Card and Calibration – Master Curve Definition Window Customer Bulletin, part number 078D0560-01 Rev. A, to match the fields in the Master Curve card with the fields as they are displayed in the ADVIA Centaur XP System.

- Select Save.

Quality Control (QC)

For detailed QC procedural information, refer to the Operator's Guide.

QC Materials

Use Ligand Plus, Tumor Marker Plus, or an equivalent quality control material to monitor assay performance.

QC Frequency

Analyze all levels of quality control material on each day that samples are analyzed.

Analyze all levels of quality control material each time a two-point calibration is performed.

Troubleshooting Out-of-Range QC Values

If the quality control results do not fall within the Expected Values or within the laboratory’s established values, do not report results. Take the following actions:

- Verify that the materials are not expired.

- Verify that required maintenance was performed.

- Verify that the assay was performed according to the instructions for use.

- Rerun the assay with fresh quality control samples.

- If necessary, contact your local technical support provider or distributor for assistance.

ANALYTIC PROCESS

See IHL-CHM-Centaur Instrument Operation

POSTANALYTIC PROCESS

Reporting Results

As with all in vitro diagnostic assays, each laboratory should determine its own reference range(s) for the diagnostic evaluation of patient results.

Reference Interval

< 4.00 µg/L

Critical Values

NA

Units for Reporting Results

The system reports serum total PSA results in ng/mL (mass units) or µg/L (SI units), depending on the units defined when setting up the assay. The conversion formula is 1 ng/mL = 1 µg/L. You can define the units and the number of decimal places for test results using the Worklist – Test Summary window

Acceptable Results

Patient test results are acceptable and may be reported when:

- QC results are acceptable.

- All instrument flags attached to the result have been investigated and resolved.

Corrective Action

- Patient test results must be repeated and corrective action taken when:

- Instrument flags indicate.

Procedure Notes

Calculations

For detailed information about how the system calculates results, refer to the Operator's Guide.

High Dose Hook Effect

Patient samples with high total PSA levels can cause a paradoxical decrease in the RLUs (high dose hook effect). In this assay, patient samples with total PSA levels as high as 10,000 ng/mL (10,000 µg/L) will assay greater than 100 ng/mL (100 µg/L).

Disposal

Dispose of hazardous or biologically contaminated materials according to the practices of your institution. Discard all materials in a safe and acceptable manner, and in compliance with all federal, state, and local requirements.

Method Limitations

Reportable Range

The reportable range of the ADVIA Centaur PSA assay is 0.01 – 100 ng/mL (µg/L)

Dilutions

- Serum samples with PSA levels greater than 100 ng/mL (100 µg/L) must be diluted and retested to obtain accurate results.

- Patient samples can be automatically diluted by the system or prepared manually.

- For automatic dilutions, ensure that ADVIA Centaur Multi-Diluent 2 is loaded and set the system parameters as follows:

- Dilution point: < 100 ng/mL (100 µg/L)

- Dilution factor: 2, 5, 10, 50, 100

- For detailed information about automatic dilutions, refer to the Operator's Guide.

- Manually dilute the patient samples when patient results exceed the linearity of the assay using automatic dilution, or when laboratory protocol requires manual dilution.

- Use Multi-Diluent 2 to manually dilute patient samples, and then load the diluted sample in the sample rack, replacing the undiluted sample.

- Ensure that results are mathematically corrected for dilution. If a dilution factor is entered when scheduling the test, the system automatically calculates the result.

Other Limitations

WARNING: The concentration of total PSA in a given specimen, as determined by assays from different manufacturers, can vary due to differences in assay methods and reagent specificity. The results reported by the laboratory to the physician must include the identity of the assay for total PSA used. Values obtained with different assay methods cannot be used interchangeably. If, in the course of monitoring a patient, the assay method used for determining serial levels of total PSA is changed, the laboratory must perform additional testing to confirm baseline values.

NOTE: Do not interpret levels of PSA as absolute evidence of the presence or the absence of malignant disease. Before treatment, patients with confirmed prostate carcinoma frequently have levels of PSA within the range observed in healthy individuals. Elevated levels of PSA can be observed in patients with nonmalignant diseases. Measurements of PSA should always be used in conjunction with other diagnostic procedures, including information from the patient’s clinical evaluation.

WARNING: Do not predict disease recurrence solely on serial PSA values.

Specimens obtained from patients undergoing prostate manipulation, especially needle biopsy and transurethral resection, may show erroneously high results. Care should be taken that PSA samples are drawn before these procedures are performed.

Prostate cancer patients under treatment with anti-androgens and LHRH agonists may exhibit markedly reduced levels of PSA. Also, men treated for benign prostatic hyperplasia with inhibitors of 5α-reductase (finasteride) may demonstrate a significant reduction in PSA levels compared to values prior to treatment. Care should be taken when interpreting values from these individuals.

Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays. Patients routinely exposed to animals or to animal serum products can be prone to this interference and anomalous values may be observed. Additional information may be required for diagnosis.

| Specimens that are . . . | Demonstrate < 5% change in results up to . . . |
| --- | --- |
| hemolyzed | 500 mg/dL of hemoglobin |
| lipemic | 1000 mg/dL of triglycerides |
| icteric | 40 mg/dL of bilirubin |

Interference testing was determined according to CLSI Document EP07-A2.

For additional information on performance characteristics, see the product information in the ADVIA Centaur PSA product insert.

Equipment and Supplies

- ADVIA Centaur PSA ReadyPack

- ADVIA Centaur Calibrator Q

- ADVIA Centaur Multi-Diluent 2 (optional)

- ADVIA Centaur Acid Reagent (0.5% H202, 0.1N HNO3)

- ADVIA Centaur Base Reagent (0.25N NaOH and surfactant)

- ADVIA Centaur Cleaning Solution Concentrate (~52.5 g/L sodium hypochlorite)

- ADVIA Centaur Sample Cups and Caps

- ADVIA Centaur Cuvettes

- ADVIA Centaur Tips

- Reagent Wate

References

- Siemens Healthcare Diagnostics ADVIA Centaur PSA Product Insert.

- Siemens Healthcare Diagnostics ADVIA Centaur Reference Manual.

- Siemens Healthcare Diagnostics ADVIA Centaur XP Operator's Guide.

- Clinical and Laboratory Standards Institute (CLSI). Clinical Laboratory Technical Procedure Manuals; Approved Guideline, GP02-A5, 2006.

- Clinical and Laboratory Standards Institute (CLSI). Interference Testing in Clinical Chemistry; Approved Guideline, EP07-A2, 2005.

Technical Assistance

Siemens Healthcare Diagnostics Technical Care Center: 1-877-229-3711

Customer Service: 1-800-255-3232

Serial Number: IRL30361209