- PRINCIPLE/PURPOSE:

  - The LI method is an in vitro diagnostic test for the quantitative measurement of lithium in human serum and plasma on the Dimension Vista® System. Measurements of lithium are used to assure that the proper drug dosage is administered in the treatment of patients with mental disturbances, such as manic-depressive illness (bipolar disorder).

  - Lithium, the third element in the periodic table, is used primarily to treat the manic phase of affective disorders, mania, and manic-depressive illness. The precise mechanism of action of lithium as a mood-stabilizing agent is not known. Lithium is administered in capsule, syrup, or tablet form as salts of either carbonate or citrate. It is readily absorbed from the gastrointestinal tract and does not bind appreciably to plasma proteins. Peak plasma concentrations are reached 2 to 4 hours after oral administration. Approximately 95% of a single dose of lithium is excreted in the urine within 6 to 12 hours, with the remainder being slowly excreted over the next 10 to 14 days.  Lithium concentrations are monitored to ensure patient compliance and prevent toxicity. Because there is a narrow therapeutic range of about 0.60 to1.20 mmol/L, with significant risk of toxicity occurring above 1.5 mmol/L, determination of lithium concentration is crucial in the management of patients on lithium therapy. Since plasma values vary relative to time of last dose, a standardized 12-hour post-dose serum lithium concentration has been recommended to assess adequate therapy.1, 2

  - The LI method employs a patented compound, 7-Nitro-2,12-dicarboxyl-16,17 dihydro-5H,
15H-dibenzo[b,i] [1,11,4,5,7,8] dioxatetraazacyclotetradecine* (lithium dye) that reacts with Li+ ion in an alkaline mixture 
of water and dimethylsulfoxide to form a non-covalent binary complex.

| Li+ + Lithium Dye |  | Dye-Li Complex (absorbs at 540 nm) |
| --- | --- | --- |

  - The change in absorbance at 540 nm due to the formation of this non-covalent binary complex is directly proportional to the lithium concentration in the sample and is measured using a bichromatic (540 and 700 nm) endpoint technique.*U.S. Patent #5,049,666

- PERFORMANCE SPECIFICATIONS:

  - Test Conditions:

| Sample Volume (delivered to the cuvette) | 4.0 µL |
| --- | --- |
| Reagent 1 | 96.0 µL |
| Reagent 2 | 40.0 µL |
| Reagent 3 | 96.0 µL |
| Temperature | 37.0 ºC |
| Reaction Time | 29 seconds |
| Wavelength | 540 and 700 nm |
| Type of Measurement | Bichromatic endpoint |

- SPECIMEN TYPE:

  - Recommended specimen types: serum and sodium heparin plasma.

  - Serum and plasma can be collected using recommended procedures for collection of diagnostic blood specimens by venipuncture.2, 3

  - Follow the instructions provided with your specimen collection device for use and processing.6

  - Serum or plasma should be separated from the cells if analysis is not performed within 2 hours of collection.2 
Once separated, lithium is stable for at least 24 hours at room temperature, 7 days if refrigerated at 4 °C.4,5 
For longer storage, samples may be frozen at -18 °C for at least 6 months.4

- EQUIPMENT AND REAGENTS:

  - Reagents:

| Wells | Form | Ingredient | Concentration |
| --- | --- | --- | --- |
| 1–4 | Liquid | Lithium Dye Reagent |  |
|  |  | Dye | 1.53 x 10-4 M |
|  |  | Dimethylsulfoxide (DMSO) | 38% v/v |
|  |  | K2HPO4 | 0.65 mM |
|  |  | Preservative |  |
| 5–8 | Liquid | Lithium Alkaline Salt Reagent |  |
|  |  | NaCl | 8 mM |
|  |  | Dimethylsulfoxide (DMSO) | 38% v/v |
|  |  | KOH | 0.074 M |
| 9–12 | Liquid | Lithium Blank Reagent |  |
|  |  | Dimethylsulfoxide (DMSO) | 38% v/v |
|  |  | K2HPO4 | 0.65 mM |
|  |  | Preservative |  |

    - Wells are numbered consecutively from the wide end of the cartridge.

    - Nominal value per test in final reaction mixture.

  - Reagent Preparation: All reagents are liquid and ready to use.

  - Store at: 2–8 °C

  - Expiration: Refer to carton for expiration date of individual unopened reagent cartridges. Sealed wells on the instrument are stable for 30 days.

  - Open Well Stability: 7 days for wells 1–12

  - Materials Provided:		LI Flex® reagent cartridge, Cat. No. K4132

  - Materials Required But Not Provided:		DRUG 1 CAL, Cat. No. KC410

Quality Control Materials

- CALIBRATION:

| 5.1 Calibration Material | DRUG 1 CAL, Cat. No. KC410 |
| --- | --- |
| 5.2 Calibration Scheme | 5 levels, n=5 |
| 5.3 Units | mmol/L |
| 5.4 Typical Calibration Levels | Level 1 (Calibrator A): 0 mmol/L Intermediate levels automatically prepared by the instrument: 0.9, 1.8, 3.6 mmol/L Level 5 (Calibrator B): 5.5 mmol/L |
| 5.5 Calibration Frequency | Every 60 days for any one lot Calibration interval may be extended based on acceptable verification of calibration. |
| 5.6 A new calibration is required: | • For each new lot of Flex® reagent cartridges • After major maintenance or service, if indicated by quality control results • As indicated in laboratory quality control procedures • When required by government regulations |

- PROCEDURE:

  - Sampling, reagent delivery, mixing, and processing are automatically performed by the Dimension Vista® System. For details of this processing, refer to your Dimension Vista® Operator’s Guide

- QUALITY CONTROL:

  - At least once each day of use, analyze two levels of a Quality Control (QC) material with known lithium concentrations. Follow your laboratory internal QC procedures if the results obtained are outside acceptable limits.

- INTERFERENCES/CROSS-REACTIONS:

  - Interfering Substances ‡

    - An increase in serum sodium concentration of 23 mmol/L from a nominal serum concentration of 140 mmol/L will produce biases of +0.06 and +0.05 mmol/L at lithium concentrations of 0.70 and 2.10 mmol/L, respectively.

    - A decrease in serum sodium concentration of 23 mmol/L from a nominal serum concentration of 140 mmol/L will produce biases of -0.06 and -0.05 mmol/L at lithium concentrations of 0.70 and 2.10 mmol/L, respectively.

    - Sodium heparin, present in blood collection tubes, does not interfere with the LI method.

  - Maximum Observed Repeatability:  The expected maximum observed standard deviations for repeatability (within-run precision) using n=5 replicates at the following lithium concentrations are:

| LI concentration | Acceptable SD Maximum |
| --- | --- |
| 1.05 mmol/L | 0.13 mmol/L |
| 2.26 mmol/L | 0.17 mmol/L |

    - A system malfunction may exist if the acceptable SD maximum is exceeded.

  - Specific Performance Characteristics;  The following data represent typical performance for the Dimension Vista® System.

    - Precision8

| Material | Mean mmol/L | Standard Deviation (% CV) |  |
| --- | --- | --- | --- |
|  |  | Repeatability | Within-Lab |
| Multiqual® Control |  |  |  |
| Level 1 | 0.54 | 0.02 (4.6) | 0.03 (5.9) |
| Level 2 | 1.09 | 0.03 (2.5) | 0.04 (3.4) |
| Level 3 | 2.26 | 0.04 (1.8) | 0.07 (3.2) |

      - CLSI/NCCLS EP5-A2 was used. During each day of testing, two separate runs, with two test samples, for each test material, were analyzed for 20 days.

      - Multiqual® is a registered trademark of Bio-Rad Laboratories, Irvine, CA 92618, USA.

8.6	Method Comparison”:9

| 8.6.1 Regression Statistics: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Comparative Method | Slope | Intercept mmol/L | Correlation Coefficient | n |
| Dimension® RxL system | 0.99 | 0.14 | 0.998 | 108 |

    - CLSI/NCCLS EP9-A2 was used. The method used to fit the linear regression line was ordinary least squares.

    - The range of serum/plasma lithium values in the correlation study was 0.3–4.8 mmol/L.

## Specificity:

## HIL Interference:  The LI method was evaluated for interference according to CLSI/NCCLS EP7-A2.10 Bias is the difference in the results between the control sample (without the interferent) and the test sample (contains the interferent) expressed in percent. Bias exceeding 10% is considered interference.

| Substance Tested | Substance Concentration | Lithium mmol/L | Bias* % |
| --- | --- | --- | --- |
| Hemoglobin (hemolysate) | Hemoglobin (monomer) 1000 mg/dL [0.62 mmol/L] | 0.9 | <10 |
| Bilirubin (unconjugated) | 60 mg/dL [1026 µmol/L] | 0.9 | <10 |
| Bilirubin (conjugated) | 60 mg/dL [1026 µmol/L] | 0.9 | <10 |
| Lipemia (Intralipid®) | 3000 mg/dL [33.9 mmol/L] | 0.9 | <10 |

8.7.2	Intralipid® is a registered trademark of Fresenius Kabi AG, Bad Homburg, Germany.

* NOTE:  Analyte results should not be corrected based on this bias.

8.7.3	Système International d’Unités [SI units] are in brackets.

### 8.8	Non Interfering Substances:  The following substances do not interfere with the LI method when present in serum and plasma in the amounts indicated. Inaccuracies (biases) due to these substances are less than 10% at lithium concentration of 0.9 mmol/L.

| Substance | Test Concentration | SI Units |
| --- | --- | --- |
| Acetaminophen | 0.025 mg/dL | 1.66 µmol/L |
| Amikacin | 15 mg/dL | 256 µmol/L |
| Ampicillin | 5.3 mg/dL | 152 µmol/L |
| Ascorbic Acid | 5 mg/dL | 227 µmol/L |
| Caffeine | 6 mg/dL | 308 µmol/L |
| Carbamazepine | 3 mg/dL | 127 µmol/L |
| Chloramphenicol | 5 mg/dL | 155 µmol/L |
| Chlordiazepoxide | 1 mg/dL | 33.3 µmol/L |
| Chlorpromazine | 0.2 mg/dL | 6.27 µmol/L |
| Cholesterol | 500 mg/dL | 12.9 mmol/L |
| Cimetidine | 2 mg/dL | 79.2 µmol/L |
| Creatinine | 30 mg/dL | 2652 µmol/L |
| Dextran 40 | 6000 mg/dL | 1500 µmol/L |
| Diazepam | 0.5 mg/dL | 17.6 µmol/L |
| Digoxin | 5 ng/mL | 6.15 nmol/L |
| Erythromycin | 6 mg/dL | 81.6 µmol/L |
| Ethanol | 400 mg/dL | 86.8 mmol/L |
| Ethosuximide | 25 mg/dL | 1770 µmol/L |
| Furosemide | 6 mg/dL | 181 µmol/L |
| Gentamicin | 12 mg/dL | 251 µmol/L |
| Heparin | 3 U/mL | 3000 U/L |
| Ibuprofen | 50 mg/dL | 2425 µmol/L |
| Immunoglobulin G (IgG) | 5 g/dL | 50 g/L |
| Lidocaine | 1.2 mg/dL | 51.2 µmol/L |
| Nicotine | 0.1 mg/dL | 6.2 µmol/L |
| Penicillin G | 25 U/mL | 25000 U/L |
| Pentobarbital | 8 mg/dL | 354 µmol/L |
| Phenobarbital | 10 mg/dL | 431 µmol/L |
| Phenytoin | 5 mg/dL | 198 µmol/L |
| Primidone | 4 mg/dL | 183 µmol/L |
| Propoxyphene | 0.2 mg/dL | 4.91 µmol/L |
| Protein: Albumin | 6 g/dL | 60 g/L |
| Protein: Total | 12 g/dL | 120 g/L |
| Rheumatoid Factor | 500 IU/mL | 500 IU/mL |
| Salicylic Acid | 60 mg/dL | 4.34 mmol/L |
| Theophylline | 4 mg/dL | 222 µmol/L |
| Triglycerides | 3000 mg/dL | 33.9 mmol/L |
| Urea | 500 mg/dL | 83.3 mmol/L |
| Uric Acid | 20 mg/dL | 1190 µmol/L |
| Valproic Acid | 50 mg/dL | 3467 µmol/L |

  - Analytical Sensitivity: 0.03 mmol/L

    - The analytical sensitivity represents the lowest concentration of lithium that can be distinguished from zero. This sensitivity is defined as the mean value (n=20) plus two standard deviations of the level 1 (0 mmol/L) DRUG 1 Calibrator.

    - ‡ The Dimension Vista® LI method (REF K4132) and the Dimension® LI (REF DF132) method utilize the same reagents under equivalent reaction conditions. Interfering substances were tested using Dimension® LI (REF DF132) and the results are representative of both methods.

- CALCULATING RESULTS:

  - Results: The instrument calculates the concentration of lithium in mmol/L using the calculation scheme described in your Dimension Vista® Operator’s Guide.

  - Results of this test should always be interpreted in conjunction with the patient’s medical history, clinical presentation and other findings.

  - Analytical Measurement Range (AMR): 0.2–5.0 mmol/L

    - This is the range of analyte values that can be measured directly on the specimen without any dilution or pretreatment that is not part of the usual analytical process and is equivalent to the assay range.

    - Samples with results in excess of 5.0 mmol/L should be repeated on dilution.

  - Manual Dilution: Dilute with lithium free serum to obtain results within reportable range. Enter dilution factor on the instrument. Reassay. Resulting readout is corrected for dilution.

  - Autodilution (AD): Not recommended

    - Samples with results less than 0.2 mmol/L will be reported as “less than 0.2 mmol/L” by the instrument.

- REFERENCE INTERVALS:

  - 0.60 – 1.20 mmol/L

- CRITICAL INTERVALS:

11.1	N/A

- LABORATORY INTERPRETATION OF RESULTS:

12.1	N/A

- SAFETY PRECAUTIONS:

  - Risk and safety:

    - Irritant. Contains mixture of 5-chloro-2-methyl-2H-isothiazol-3-one and 2-methyl-2H-isothiazol-3-one (3:1).

    - Irritating to eyes and skin.

    - May cause sensitization by skin contact.

    - Avoid contact with skin.

    - Wear suitable gloves.

    - Safety data sheets (MSDS/SDS) available on www.siemens.com/diagnostics

13.2	Precautions: Used cuvettes contain human body fluids; handle with appropriate care to avoid skin contact or ingestion.

13.2.1For in vitro diagnostic use.

- POTENTIAL SOURCES OF VARIABILITY:

  - Limitations of Procedure:  The instrument reporting system contains flags and comments to provide the user with information regarding instrument processing errors, instrument status information and potential errors in lithium results. Refer to your Dimension Vista® Operator’s Guide for the meaning of report flags and comments. Any report containing flags and/or comments should be addressed according to your laboratory’s procedure manual and not reported.

- REFERENCES:

  - Hardman JG, Goodman Gilman A, Limbird LE. Goodman and Gilman’s The Pharmacological Basis of Therapeutics, Ninth Edition, The McGraw-Hill Companies, Inc., New York, NY 1996; pp. 446-453.

  - Burtis CA, Ashwood ER. Tietz Fundamentals of Clinical Chemistry, Fifth Edition, W.B. Saunders Company, Philadelphia, PA 2001; pp. 30-53 (specimen collection), pp. 631 (clinical significance), pp. 1023 (reference values).

  - Clinical and Laboratory Standards Institute/NCCLS. Procedures for the Collection of Diagnostic Blood Specimens by Venipuncture; Approved Standard—Fifth Edition. CLSI/NCCLS document H3-A5 [ISBN 1-56238-515-1]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA, 2003.

  - Kaplan LA, Pesce AJ. Clinical Chemistry: Theory, analysis, and correlation, Third Edition, Mosby, Inc., St. Louis, MO 1996; pp. 1108 (specimen storage and reference values).

  - Young DS. Effects of Preanalytical Variables on Clinical Laboratory Tests, Second Edition, AACC Press, Washington, DC 1997; pp. 4-500 (specimen storage).

  - Clinical and Laboratory Standards Institute /NCCLS. Tubes and Additives for Venous Blood Specimen Collection; Approved Standard—Fifth Edition. CLSI/NCCLS document H1-A5 [ISBN 1-56238-519- 4]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA, 2003.

  - Baselt RC, Cravey RH. Disposition of Toxic Drugs and Chemicals in Man, Fourth Edition, Chemical Toxicology Institute, Foster City, CA 1995; pp. 430 (reference values).

  - Clinical and Laboratory Standards Institute /NCCLS. Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline—Second Edition. CLSI/NCCLS document EP5-A2 [ISBN 1-56238-542-9]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA, 2004.

  - Clinical and Laboratory Standards Institute/NCCLS. Method Comparison and Bias Estimation Using Patient Samples; Approved Guideline—Second Edition. CLSI/NCCLS document EP9-A2 [ISBN 1-56238-472-4]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA, 2002.

  - Clinical and Laboratory Standards Institute/NCCLS. Interference Testing in Clinical Chemistry; Approved Guideline—Second Edition. CLSI/NCCLS document EP7-A2 [ISBN 1-56238-584-4]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA, 2005.

Dimension Vista®, Dimension® and Flex® are trademarks of Siemens Healthcare Diagnostics Inc.

©2008 Siemens Healthcare Diagnostics Inc.

All rights reserved.