- Principle/Purpose:

  - To maintain consistent performance of the Vista system a monthly System Check is required to assess the reagent probe performance.

- Performance Specifications:

  - This process may correct the following issues:

| Assay | Observations |
| --- | --- |
| Calcium | Flagged or low results |
| Creatinine | Flagged results |
| FT4 | High results |
| Glucose | Flagged results |
| Lipase | High results |
| Lithium | Flagged results |
| LDL | Flagged results |
| Microalbumin | High results |
| Phosphorous | High results |
| Total Bilirubin | High results |
| Total Protein | High or low results |
| Uric acid | Low results |
| UCFP | High results |
| Vancomycin | Flagged results |

- Specimen Type:

  - NA

- Equipment and Reagents:

  - Dimension Vista System

  - CHK Flex (part# KS996)

  - Surplus Sample Rack

  - Sample cup filled with one Flex well worth of CHK fluid

- Calibration:

  - NA

- Procedure:

  - Press the Advanced icon, then Diagnostics icon.

  - Select System Check from the Diagnostics menu.

  - Load or verify 2 CHK Flexes on the instrument. Use a 3rd CHK Flex to add solution to sample cup (see 6.6 below)..

  - Scan the sample rack barcode and verify the correct numbers are displayed in the sample rack barcode field.

  - Verify that the correct position is displayed in the sample rack position field.

  - Place a sample cup filled with RED CHK fluid (found in wells 1, 3, 5 and 7 only) onto the instrument.

  - Select Order System Check from the screen.

  - Place the rack in the sample lane and wait for tests to complete.

  - Tests results can viewed on the System Check screen when completed (also viewable from the Results screen: Advanced > Results). These results can be printed.

  - Any failures are displayed in red on the System Check screen. Report failures to Siemens Technical Support.

- Quality Control:

  - NA

- Interferences/Cross-Reactions:

  - NA

- Calculating Results:

  - Once the System Check has been run, evaluate the accuracy and precision of reagent delivery for each combination of probes by determining if the percent difference is greater than 4% using this calculation:

  - [(Highest CR#BS mean- Lowest CR#BS mean)/ Highest CR#BS mean] x 100 = % Difference

NOTE: # = reagent probe number

Server 1: CR1BS compared to CR2BS

Server 2: CR4BS compared to CR3BS

Server 3: CR5BS compared to CR2BS

  - For each of the 3 combinations in step 2, if the percent is greater than 4%, replace the associated reagent probe with the lowest Result MEAN from that combination (see CR1BS/CR2BS example below:

| CR1BS MEAN | CR2BS MEAN | % DIFFERENCE | PASS/FAIL | REPLACE R1 PROBE | REPLACE R2 PROBE |
| --- | --- | --- | --- | --- | --- |
| 384 | 400 | 4 | PASS | NO | NO |
| 360 | 400 | 10 | FAIL | YES | NO |
| 346 | 353 | 2 | PASS | NO | NO |
| 353 | 326 | 8 | FAIL | NO | YES |

  - See Dimension Vista Operator’s iGuide- Chapter 7- Other Maintenance- Maintaining Probes or Dimension Vista Operator’s Guide Section 7-64- Maintaining Probes for instructions to change probe.

- Reference Intervals:

  - NA

- Critical Intervals:

  - NA

- Laboratory Interpretation of Results:

  - See 9.3 above.

- Safety Precautions:

  - Follow all Hospital and IHL Safety Policies

- Potential Sources of Variability:

  - NA

- References:

  - Siemens Urgent Medical Device Correction 14-53 November 2014