PURPOSE

Reagent lot crossovers should be performed on patient samples prior to patient results being reported using new lots of reagent.  Patient samples can be processed using all onboard reagent lots.

PROCEDURE

- Press the Advanced icon, then the Samples icon.  Select Batch Test Mode Test Setup from the menu.

- Select Modify Batch Mode Configuration from the Actions menu.

- Enter Batch ID in the Batch ID field.  (The method code can be typed in here or if testing multiple analytes, a batch number can be entered)

- Select  ‘Sample Fluid Type’ from the drop-down menu.

- Select the method(s) to be processed.

- If multiple replicates are needed select the method of choice and the number of times needed to be run.

- Scan the sample rack barcode.  Use the blue rack and sample cups.  The barcodes are listed in the display box for multiple sample racks.  The sample rack barcodes can also be manually entered in the space provided.

- Press Save Changes in the Actions menu.

- Select Lot Crossover in the Batch Mode menu.

- Load the sample racks on the instrument.

- When the batch mode is processed the sample ID is generated by Rack+Position+Sequence number.  The same sequence number is added to the Batch ID to form the Patient ID field.

- As the results are completed, the Dimension Vista System displays the results on the Results History screen.  The results can also be viewed in the Test Status screen.

- Once the testing is complete, print the results.  Each result will be printed with the reagent Lot number beside it.

- Using the form IHL-QAC-II Reagent and Calibration Validation, fill in the test results.  The results from the new reagent should be within  +/- 10% of the old (current) reagent results.

If the results do not fall within the acceptable range  See IHL-QAC-I Setting up QC Parameters for further instruction.

- Initial the results and give to Dr. Keys for review.

Note:  Once all of the samples in the racks designated to process the Reagent Crossover have completed, those sample racks used in the study will automatically process as normal sample racks for subsequent patient orders.

*For IMT analytes: sodium, potassium, and chloride, patient comparisons will not be performed at this time.*

**For low volume tests try, if possible, to perform the comparison. If unable to perform comparison testing fill out the log sheet with reason not performed.**

REFERENCES

- Siemens Dimension Vista System Operator’s Guide

- Siemens Dimension Vista System iGuide