1.0 General QC Troubleshooting

The following steps can be followed to aid with troubleshooting QC problems.

- Verify the following:

  - The QC product is correct.

  - There is sufficient QC volume in the onboard Dimension Vista® vial.  See IHL-CHM-VISTA-II QC Quick Reference Guide – Met Campus

  - The QC onboard is within the QC vendor’s open stability claim for the analyte.

  - The outlier is within 3 SD.

  - There is sufficient reagent volume in the Flex® reagent cartridge.

- Examine the calibration more closely:

  - Verify coefficients/calibrator signals and compare with previous acceptable calibration.

  - Proceed to Step 3 below if calibration is acceptable.

  - If calibration is unacceptable, proceed to recalibrating the existing or a new Flex® reagent cartridge lot.  Rerun QC.

- Rerun QC onboard Dimension Vista® instrument in duplicate to rule out random error.

- If QC is still unacceptable, run QC in duplicate from a sample cup.

- If QC is now acceptable, remove onboard QC vial.  Fill a new Dimension Vista® vial with fresh QC.  Reload on instrument.  Continue to run and report patient results.

- If QC is still unacceptable, remove and reload the same Flex® reagent cartridge.  Rerun QC and patient samples if desired, using a new well set.  Run some patient samples that were previously tested when the QC was acceptable. If QC and patient results are acceptable, continue to run and report patient samples.

- If QC is still unacceptable, review recent QC data for shifts and gradual trending.  Check for the following:

  - Outlier is outside 3SD limits (1-3s)

  - 2 or more consecutive values are on the same side of the mean

  - Range of two consecutive points is greater the 2SD (2-2s)

  - Both QC levels shift is in the same direction

  - Check peer values

  - If new lot of reagent means may need to be adjusted.  Consult with the Quality Control Officer.

- Install a new Flex® reagent cartridge (same lot number).  Run QC.  If QC is acceptable then continue to run and report patient results.

- If QC is still out, recalibrate and rerun QC.  If both calibration and QC are acceptable, rerun 5 patient samples for comparison.  If acceptable continue to run and report patient results.

- If QC is still unacceptable, do not recalibrate.  Contact the Technical Resource Center for further troubleshooting guidance and log the issue with Siemens.  Enter all troubleshooting performed in the troubleshooting logs.

See IHL-CHM-VISTA-II Quality Control Procedure for additional information

2.0 Off-Peak Activities

To optimize work on the Vistas the following off-peak activities can be performed:

Before the QC is set to run the reagent, calibrator and QC vial supplies should be checked.

Vista 1 at 11:00

Vista 2 at 16:00

Note:  The minimum amount of time to perform the off-peak activities is 30 minutes prior to the QC beginning.

Procedure:

- From the Home screen select  ‘Setup’

- Click on Needs (at bottom of screen)

- From the column on the right hand side of the screen select each in turn:

Reagent Cartridges

Calibrator Vials

QC Vials

Sample Diluent Vials

- From the information provided by checking the above, add any reagents, QC, or Calibrators if necessary.

- REFERENCES

3.1	Dimension Vista® System Operator’s Guide

CE 98/79/EC, Siemens Healthcare Diagnostics Inc. Sir William Siemens Sq. Frimley, Camberley, UK GU168QD

  - Bio-Rad Quality Control Package inserts

  - Laboratory Quality Control manual