The Coagulation laboratory initiated auto-verification on 10/6/01.  It is the department policy that all normal results and selected abnormal results from the IL ACL-Top coagulation analyzers will be auto-verified in the laboratory’s LIS.  The following criteria will be used for autoverification:

| Autoverification Criteria IL ACL-Top Coagulation Analyzers When all of the following criteria evaluate as “True”, then the LIS system will auto-verify INR, Activated Partial Thromboplastin Time (aPTT), Fibrinogen (FIB), Thrombin Time (TT), or D-dimer. |  |
| --- | --- |
| Criteria Name | Range |
| INR | 0.9 – 3.0 |
| INR (AMS) | 0.9 – 3.0 |
| aPTT | 25 - 95 sec |
| FIB | 151 - 449 mg/dL |
| TT | 16 - 25 sec |
| D-dimer | 250 – 10,000 ng/mL FEU |

| Delta Check Rules In addition, an assay not passing the following delta check rules assigned into the LIS will not be auto-verified: |  |
| --- | --- |
| Criteria Name | ## Range |
| All INR’s | Difference Between Previous INR and Current INR > 2.0 |

Validation of auto-verification will be performed annually and whenever there is a change to the system that could affect the autoverification logic.  This is done by using previously tested patient samples, which are selected in order to challenge each of the previously stated criteria as listed below.

# Validation Samples

| Assay |  | Criteria |
| --- | --- | --- |
| INR/INR AMS | 0.9-3.0 sec, passed delta check, autoverified |  |
| INR/INR AMS | 0.9-3.0 sec failed delta check, did not autoverify |  |
| INR/INR AMS | > 3.0 sec, did not autoverify |  |
| aPTT | < 25 sec, did not autoverify |  |
| aPTT | 25-95 sec, autoverified |  |
| aPTT | > 95 sec, did not autoverify |  |
| FIB | 0-150 mg/dL, did not autoverify |  |
| FIB | 151-449 mg/dL, autoverified |  |
| FIB | 449-1000 mg/dL, did not autoverify |  |
| TT | <16 sec, did not autoverify |  |
| TT | 16-25 sec, autoverified |  |
| TT | > 25 sec, did not autoverify |  |
| D-Dimer | <250 ng/mL FEU, did not autoverify |  |
| D-Dimer | 250 – 10,000 ng/mL FEU, autoverified |  |
| D-Dimer | >10,000 ng/mL FEU, did not autoverify |  |

This is a zero-tolerance system for errors in the LIS software application.  Any detected failure of the software system to release or not release sample results for auto-verification based on these criteria will be reported to IT for correction / modification of the software.

Authorized Reviewers

Medical Director, Coagulation

##### Document Control

##### Location of Master: Coagulation Procedure Manual

Master electronic file stored on the Clinical Pathology server:

S:\HEMACOAG\Document Control\Coagulation\Policy\Master Document\Autoverification policy.doc

Number of Controlled Copies posted for educational purposes:  0

Number of circulating Controlled Copies:  0

Location of circulating Controlled Copies: NA

##### Document History

| Signature | Date |  |  |
| --- | --- | --- | --- |
| Prepared by: Karri Henderson MT (ASCP) | 03/28/2007 |  |  |
| Approved by: Marc Smith MD | 03/28/2007 |  |  |
| Reviewed by: (Signature) | Date | Revision # | Modification |
| Marc Smith, MD | 12/05/2007 | 00 | New procedure format |
| Marc Smith, MD | 02/13/2008 | 01 | Changed PT, PT-NH, PT-AMS and TT autoverification ranges on pg 1. Changed PT, PT-NH, PT-AMS and TT autoverification criteria on pg 2. Deleted signature lines on pg 2. |
| Marc Smith, MD | 02/11/2009 | 02 | Changed PT, PT-NH, PT-AMS, aPTT, Heparin aPTT and TT autoverification ranges and criteria. |
| Marc Smith, MD | 08/18/2009 | 03 | Corrected PT-AMS delta check, did not autoverify range. |
| Marc Smith, MD | 02/23/2010 | 04 | Change autoverification ranges for PT, PT-NH and PT-AMS. |
| Marc Smith, MD | 07/13/2011 | 05 | Changed PT autoverification ranges to INR autoverification ranges. Changed Mysis to LIS or IT. Added selected abnormal results. |
| Mark Kolins, MD | 10/20/2011 |  | No change |
| Marc Smith, MD | 10/25/2013 |  | No change |
| Marc Smith, MD | 07/24/2015 | 06 | Deleted autoverification for PT-NH |
| Marc Smith, MD | 04/04/2017 | 07 | Changed TT autoverification to 16-20 second |
| Elizabeth Sykes, MD | 02/22/2018 |  |  |
|  | 10/02/2018 | 08 | Changed instrumentation from Sysmex to IL ACL-Top. Changed autoverification range for INR and INR/AMS to 3.0, aPTT to 25-95, FIB to 151-449, TT to 16-25, and D-dimer to 250-10,000. |
| Peter Millward, MD | 03/13/2019 |  |  |