Purpose

The Iricell 2000 may require a system restart or shutdown to correct errors encountered 	while analyzing specimens. This procedure describes the appropriate way to restart or 	shutdown.  It has the steps for the Velocity and the iQ200

Procedure

Velocity and iQ 200 Shutdown

When it is necessary to restart the instrument, shut down the instrument in the following 	order:

1. Select Instrument on the top right of the main screen

2. Select Go Off Line. The system will prompt: Continue going off line?

3. Select Yes

4. Select Maintenance located at the bottom of the Instrument Screen

5. Select Shutdown.  The system will prompt: Do you want to shut down the 			    instrument?

6. Select Yes. Windows will close

7. Turn OFF the iChemVELOCITY and/or iQ200 by pressing the ON/OFF button

NOTE:  Restart the Velocity and iQ200 Modules as soon as possible.  Do 				  not allow the instrument to cool down or the measurements will 					  be affected.

8. To restart the computer, open the flap on the front of the main computer

9. Press the Power button.

10. The software will automatically load after about 1 minute

11. Log onto the system.

12. Select Go online.

Velocity and iQ200 Restart

This button is used to shut down and to restart only the user interface software, without 	any effect on the systems. This function must be used after software settings have been 	modified, communication settings, or error message requiring a restart of the instrument

1. Select Instrument on the top right of the main screen

2. Select Go Off Line.

3. Select Maintenance located at the bottom of the Instrument Screen

4. From the Maintenance screen, select Restart. The system will prompt: Do you 		    want to restart the instrument?

5. Select Yes

5. Log onto the system.

6. Select Go online.

##### Document Control

##### Location of Master: Master electronic file stored on the Beaumont Laboratory server under

S:/Master printed document stored in Chemistry Laboratory.

Number of Controlled Copies posted for educational purposes:  0

Number of circulating Controlled Copies:  0

Location of circulating Controlled Copies: n/a

##### Document History

| Signature | Date | Revision # |  | Related Documents Reviewed/ Updated |
| --- | --- | --- | --- | --- |
| Prepared by: Myrna Harbar MT(ASCP) | 11/12/2019 | r00 |  |  |
| Approved by: |  |  |  |  |
| Reviewed by: (Signature) | Date | Revision # | Modification | Related Documents Reviewed/ Updated |