Purpose

The purpose of this procedure is to provide a document to demonstrate the operation and maintenance of the Tissue-Tek coverslipper. Technicians, technologists and lab assistants using this coverslipper should have been trained and be familiar with its operation.

Information can be found in the operation manual located in histology.

Forms

AP-HISTO-FORMS-13 v.1

Safety

Always wear personal protective equipment (gloves, gown, goggles) when handling chemicals and solutions.

Supplies and

Reagents:

- Xylene

- Coverslipping film

- Film cutting blade (Tissue-Tek)

Procedure

See the operating manual located in histology for operation

- Maintenance

- Daily:

    - wipe down surfaces to remove any dirt/stain

    - ensure xylene is up to the designated level in the holding tray

    - ensure the xylene waste bottle is empty

    - ensure the clean xylene bottle is filled

    - check that 2 new rolls of coverslipping film are open

    - dump holding tray xylene, wipe clean, replace with fresh xylene

- use a xylene soaked cotton tipped applicator to remove “sticky” residue from all surfaces surrounding the coverslipping zone to eliminate errors due to baskets or slides “jamming”

- Monthly (or as needed)

- Replace cutting blade

- Replace charcoal filter every 6 months.

Results

The coverslipper will be properly operated and maintained

Authorized Reviewers

Director, Anatomic Pathology

Chief, Surgical Pathology

##### Document Control

##### Location of Master: Master electronic file stored on the Beaumont Laboratory server under

##### S:\Anatomic Pathology\Royal Oak\Histology\Document Control\Master Documents\Histology Procedures\Equipment

Number of Controlled Copies posted for educational purposes:  0

Number of circulating Controlled Copies:  1

Location of circulating Controlled Copies: Histology Procedure Manual bench copy

##### Document History

| Signature | Date | Revision # |  | Related Documents Reviewed/ Updated |
| --- | --- | --- | --- | --- |
| Prepared by: S.Scalise | 10-28-2013 | 00 |  |  |
| Approved by: Dr. Armin | 10-28-2013 | 00 |  |  |
| Reviewed by: (Signature) | Date | Revision # | Modification | Related Documents Reviewed/ Updated |
| Dr. Amin |  |  |  |  |
| Dr. Bernacki | 10-31-2017 |  |  |  |
| Dr. Bernacki | 1/8/2018 | 01 | Added procedure – Replace charcoal filter every 6 months |  |
| Dr. Bernacki | 10/22/2019 | 01 |  |  |
| Dr. Bernacki |  | 02 | Removed type of coverslipping film Removed weekly maintenance and added to daily maintenance |  |