Purpose

The purpose of this document is to provide a procedure to demonstrate all processing programs.

PROGRAM #1

SMALL BIOPSIES AND SKINS

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | OFF |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 2 | OFF |  |  |  |  |
| 3 | OFF |  |  |  |  |
| 4 | 10 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 7 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 7 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 10 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 7 | XYLENE | 40 | ON | ON |
| 9 | 7 | XYLENE | 40 | ON | ON |
| 10 | 10 | XYLENE | 40 | ON | ON |
| 11 | 5 | PARAFFIN | 60 | ON | ON |
| 12 | 5 | PARAFFIN | 60 | ON | ON |
| 13 | 10 | PARAFFIN | 60 | ON | ON |
| 14 | 10 | PARAFFIN | 60 | ON | ON |

***Only small specimens that are easily penetrated can be run on this program.

PROGRAM #2

8 HOUR PROCESS

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 3 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 4 | 30 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 30 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 30 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 45 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 30 | XYLENE | 40 | ON | ON |
| 9 | 30 | XYLENE | 40 | ON | ON |
| 10 | 45 | XYLENE | 40 | ON | ON |
| 11 | 30 | PARAFFIN | 60 | ON | ON |
| 12 | 30 | PARAFFIN | 60 | ON | ON |
| 13 | 45 | PARAFFIN | 60 | ON | ON |
| 14 | 45 | PARAFFIN | 60 | ON | ON |

Any surgical and/or autopsy specimen may be placed on this run.

PROGRAM #3

6 HOUR PROCESS

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | 20 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | 20 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 3 | 20 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 4 | 20 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 15 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 15 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 25 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 15 | XYLENE | 40 | ON | ON |
| 9 | 15 | XYLENE | 40 | ON | ON |
| 10 | 25 | XYLENE | 40 | ON | ON |
| 11 | 15 | PARAFFIN | 60 | ON | ON |
| 12 | 15 | PARAFFIN | 60 | ON | ON |
| 13 | 15 | PARAFFIN | 60 | ON | ON |
| 14 | 20 | PARAFFIN | 60 | ON | ON |

***No breast, no uterus, no placenta or other fatty, dense or bloody specimens should be put on this process.

PROGRAM #4

FRIDAY SURGICALS TO COME OFF EARLY SATURDAY MORNING

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 3 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 4 | 30 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 30 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 30 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 45 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 30 | XYLENE | 40 | ON | ON |
| 9 | 30 | XYLENE | 40 | ON | ON |
| 10 | 45 | XYLENE | 40 | ON | ON |
| 11 | 30 | PARAFFIN | 60 | ON | ON |
| 12 | 30 | PARAFFIN | 60 | ON | ON |
| 13 | 45 | PARAFFIN | 60 | ON | ON |
| 14 | 45 | PARAFFIN | 60 | ON | ON |

Any surgical specimen may be placed on this run.

PROCESSING SHOULD BE STARTED BY 5:00P.M. ON FRIDAY ONLY IF REQUESTED

PROGRAM #5

PROGRAM #5

SUNDAY SURGICALS TO FINISH MONDAY MORNING

# “DELAYED START”

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 3 | 30 | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 4 | 30 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 30 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 30 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 45 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 30 | XYLENE | 40 | ON | ON |
| 9 | 30 | XYLENE | 40 | ON | ON |
| 10 | 45 | XYLENE | 40 | ON | ON |
| 11 | 30 | PARAFFIN | 60 | ON | ON |
| 12 | 30 | PARAFFIN | 60 | ON | ON |
| 13 | 45 | PARAFFIN | 60 | ON | ON |
| 14 | 45 | PARAFFIN | 60 | ON | ON |

Any surgical and/or autopsy specimen may be placed on this run.

PROCESSING SHOULD BE STARTED ANY TIME BEFORE 6:00 P.M. SUNDAY

PROGRAM #7

STAT KIDNEY BIOPSY PROCESS

# “IMMEDIATE START”

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | OFF | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | OFF | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 3 | OFF | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 4 | 10 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 7 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 10 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 10 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 7 | XYLENE | 40 | ON | ON |
| 9 | 10 | XYLENE | 40 | ON | ON |
| 10 | 10 | XYLENE | 40 | ON | ON |
| 11 | 7 | PARAFFIN | 60 | ON | ON |
| 12 | 7 | PARAFFIN | 60 | ON | ON |
| 13 | 7 | PARAFFIN | 60 | ON | ON |
| 14 | 10 | PARAFFIN | 60 | ON | ON |

All STAT kidney biopsies, any other small well-fixed biopsies needing STAT processing.

PROCESSING CAN BE STARTED ON IMMEDIATE START AT ANY TIME

PROGRAM #8

EYE/PROSTATE MACRO CASSETTES

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | OFF | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | OFF | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 3 | OFF | 80% ALCOHOLIC FORMALIN | 40 | ON | ON |
| 4 | 30 | 95% ALCOHOL | 40 | ON | ON |
| 5 | 60 | 100% ALCOHOL | 40 | ON | ON |
| 6 | 60 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 90 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 45 | XYLENE | 40 | ON | ON |
| 9 | 45 | XYLENE | 40 | ON | ON |
| 10 | 60 | XYLENE | 40 | ON | ON |
| 11 | 45 | PARAFFIN | 60 | ON | ON |
| 12 | 60 | PARAFFIN | 60 | ON | ON |
| 13 | 60 | PARAFFIN | 60 | ON | ON |
| 14 | 90 | PARAFFIN | 60 | ON | ON |

Any well fixed prostate macro cassettes.

PROCESSING CAN BE STARTED ON IMMEDIATE START AT ANY TIME

PROGRAM #10

Breast Program

STATION       TIME        SOLUTION                                   HEAT         VACUUM   PRESSURE

| 1 | 15 | 10% NEUTRAL BUFFERED FORMALIN | 40 | ON | ON |
| --- | --- | --- | --- | --- | --- |
| 2 | 45 | 10% NEUTRAL BUFFERED FORMALIN | 40 | ON | ON |
| 3 | 60 | 10% NEUTRAL BUFFERED FORMALIN | 40 | ON | ON |
| 4 | 30 | 70% ALCOHOL | 40 | ON | ON |
| 5 | 30 | 95% ALCOHOL | 40 | ON | ON |
| 6 | 60 | 100% ALCOHOL | 40 | ON | ON |
| 7 | 60 | 100% ALCOHOL | 40 | ON | ON |
| 8 | 90 | 100% ALCOHOL | 40 | ON | ON |
| 9 | 60 | XYLENE | 40 | ON | ON |
| 10 | 60 | XYLENE | 40 | ON | ON |
| 11 | 30 | PARAFFIN | 60 | ON | ON |
| 12 | 30 | PARAFFIN | 60 | ON | ON |
| 13 | 45 | PARAFFIN | 60 | ON | ON |
| 14 | 45 | PARAFFIN | 60 | ON | ON |

Authorized Reviewers

Medical Director, Anatomic Pathology

Chief, Surgical Pathology

##### Document Control

##### Location of Master: Master electronic file stored on the Beaumont Laboratory server under

S:\Anatomic Pathology\Royal Oak\Histology\Document Control\Master Documents\Histology Procedures\Processing

Number of Controlled Copies posted for educational purposes:  0

Number of circulating Controlled Copies:  1, Located in the Histology Procedure Manual bench copy

##### Document History

| Signature | Date | Revision # |  | Related Documents Reviewed/ Updated |
| --- | --- | --- | --- | --- |
| Prepared by: | 10/28/13 | 00 | Sharon Scalise |  |
| Approved by: | 10/28/13 | 00 | Dr. Armin |  |
| Reviewed by: (Signature) | Date | Revision # | Modification | Related Documents Reviewed/ Updated |
| Dr. Mitual Amin | 2/14/15 | 00 |  |  |
| Dr. Kurt Bernacki | 9/14/17 | 01 | 1. Modified the Prostate processing protocol to the eye protocol. Steps #8, 9, 10. 2. Added VIP to the title 3. Renamed program #12 to Sunday Holiday Breast Program. Changed station #4 from 12 hours to 26 hours in 95% Alcohol. Should be started at 5am to come off on Tuesday by 3am. |  |
| Dr. Kurt Bernacki | 11-6-17 | 01 |  |  |
| Dr. Kurt Bernacki | 10-22-19 | 01 |  |  |
|  | 11/19/2020 | 02 | Revise Breast program #10 and removed protocol #12 |  |