## Slide 1: TEG® 6s Hemostasis Analyzer
Operator Training

## Slide 2: TEG® 6s Analyzer
The TEG 6s hemostasis analyzer is a non-invasive diagnostic instrument designed to monitor and analyze the hemostasis state of a whole blood sample in order to assist in the assessment of patient clinical hemostasis conditions.
The TEG analyzer is indicated for use in adult patients where an evaluation of their blood hemostasis properties is desired. Hemostasis evaluation with the TEG 6s system is used to assess clinical conditions:
    In cardiovascular surgery and cardiology procedures to assess hemorrhage or thrombosis conditions (item codes 07-601-US, 07-614-US, and 07-615-US).
    In a trauma setting to assess bleeding or clotting conditions (item code 07-605-US).
    In patients where heparin/heparinoids may be present and who are at an increased risk of coagulopathy, in cardiovascular surgery, cardiovascular procedures (e.g. minimally invasive valve replacement or repairs) and liver transplantation (item code 07-604-US).
Please refer to the TEG 6s User Manual (115191-US) for Indications for Use, Contraindications, Warnings, Precautions, and Potential Adverse Events.
For technical support, please call 1-800-GET-A-TEG (1-800-438-2834) Option 3 or your local Haemonetics representative

## Slide 3: Topics
System overview
Cartridges
Principle
Starting the analyzer
Running a patient sample
Test completion
Results display
Running a quality control sample
Troubleshooting and maintenance

## Slide 4
TEG® 6s System
Overview
Set up
Analyzer

## Slide 5: TEG® 6s System Overview
Analyzer
    All-in-one, stand-alone analyzer
    Moves, measures, mixes and tests sample
    Displays and stores results
  Cartridge
    Contains up to 4 test channels and associated reagents
  TEG Manager® software (Optional)
    Browser-based remote viewing and administration software
Cartridge
Analyzer
System configuration by technician via PC Ethernet connection

## Slide 6: Set Up
Location
Reasonably flat and level
Protected from accidental damage
30 cm (12 in) clearance at back for air flow
Isolate the analyzer from all sources of heat
  For example, laptops, heaters or other heat-admitting equipment
Power
Connect AC power adapter to grounded outlet (100 – 240V, 50/60 Hz)
Power cord is firmly screwed into power jack on back side of analyzer
Consider use of UPS during operation
Optional
Barcode scanner attached to USB port Thermal printer attached to USB port
Ethernet cable connecting network port to analyzer for TEG Manager® software

## Slide 7: Analyzer Front
LCD touch screen
User interface
16.5 cm (6.5 inch) color screen
Prompts, directions, selections, and results are displayed
Cartridge slot
Accepts analyzer cartridge
Locks cartridge in place during test
Lighted strip
LED lights up for cartridge insertion and flashes for cartridge removal
Touch screen
Lighted strip
Cartridge slot

## Slide 8: Analyzer Back Panel
Carrying handle
Ethernet
port
Power jack
On/off
switch
Fuse
USB
connections
Cooling fan with filter
Serial Number
Vent

## Slide 9
Cartridges
Packaging
Cartridge parts
Storage
Cartridge types
Blood samples

## Slide 10: Patient Cartridge Package
What’s in the box?
Instructions for Use
Test cartridges (10)
Transfer pipettes (12)
x10
x12
Insert

## Slide 11: Cartridge Storage
Store in sealed pouches at 2 – 8 ⁰C until use
Ideally cartridges should come to room temperature prior to use
If needed cartridges may be run immediately after removal from the refrigerator
However, if error occurs, remove cartridge, rub barcode to remove condensation, wait a minute to warm and then re-run
Unused cartridges in sealed pouch may be returned to the refrigerator

## Slide 12: Cartridge Parts
Citrated: K, KH, RT, FF
EXP 25NOV15
Minimum fill level indicator
Reservoir window
Cartridge type
Sample port
Barcode read by analyzer
Expiration date for cartridge

## Slide 13: Global Hemostasis Cartridge (07-601-US*)
Kaolin Heparinase (CKH)
Functional Fibrinogen (CFF)
CFF (MA
parameter) Assesses fibrin clot
strength
CK (R parameter)
Assesses clot rate
Kaolin (CK)
CKH (R parameter) Assesses clot rate (heparin neutralized)
RapidTEG™ (CRT)
CRT (MA parameter) Assesses overall clot strength
* Indicated for use in cardiovascular surgery and cardiology procedures

## Slide 14: Global Hemostasis – HN Cartridge (07-604-US*)
Kaolin Heparinase (CKH)
Functional Fibrinogen (CFFH)
CK (R parameter)
Assesses clot rate
Kaolin (CK)
CKH (R & LY30 parameter) Assesses clot rate and stability (heparin neutralized)
RapidTEG™ CRTH)
CRTH (MA parameter) Assesses overall clot strength (heparin neutralized)
CFFH
(MA parameter) Assess fibrin clot strength (heparin neutralized)
* Indicated for use in cardiac surgery, liver transplant surgery and cardiovascular procedures

## Slide 15: Global Hemostasis with Lysis Cartridge (07-605-US*)
CFF (MA parameter) Identifies fibrin clot strength
RapidTEG™ (CRT)
Functional Fibrinogen (CFF)
CK (R & LY30 parameter) Assesses clot rate and clot stability
Kaolin (CK)
CRT (MA parameter) Assesses overall clot strength
* Indicated for use in trauma settings

## Slide 16: PlateletMapping® Cartridge (07-614-US/07-615-US*)
ActivatorF (ActF)
Arachidonic Acid (AA)
AA (MA parameter) Assesses AA stimulated clot strength to determine aspirin response
HKH (MA parameter) Assesses clot strength
Kaolin Heparinase (HKH)
ActF (MA parameter) Assesses fibrin – only clot strength
Adenosine Diphosphate (ADP)
ADP (MA parameter)
Assesses ADP stimulated clot strength to determine response to ADP inhibitors eg. clopidogrel
* Indicated for use in cardiovascular surgery and cardiology procedures

## Slide 17: Blood Sample Types
Citrated (Global Hemostasis) Cartridges
3.2% Sodium Citrate Vacutainer – prevents clotting in the tube by binding with calcium
Global Hemostasis (07-601-US)
10 mins (incubation) – 4 hours
Global Hemostasis – HN (07-604-US)
10 mins (incubation) – 2 hours
Global Hemostasis with Lysis (07-605-US)
10 mins (incubation) – 2 hours
PlateletMapping® Cartridges
Non-Gel Sodium or Lithium Heparin Vacutainer – heparin prevents thrombin generation activating platelets
≥ 14.5 IU but ≤ 20 IU Heparin per mL of Blood
ADP & Arachidonic Acid (AA) (07-614-US)
30 mins (incubation) – 2 hours
Adenosine Diphosphate (ADP) only (07-615-US)
15 mins (incubation) – 3 hours
Note: If red-top tube is used for waste, ensure the tube does not contain any additives as this can affect results.
3ml waste for phlebotomy
3 x dead space for A- Line or similar
No under-filling, completely fill under vacuum
Invert x 5-7 prior to testing
Note: In urgent situations, samples may be tested immediately

## Slide 18
Principle
Sample measurement and mixing
Resonant frequency technology
Test result – coagulation components

## Slide 19: Sample Moved to Mixing Chamber
TEG 6s automatically draws a metered amount of blood into the Reagent addition and mixing area.
Metered amount is the exact amount required for performing up to 4 simultaneous tests from one blood sample.
Test area
Reagent addition and mixing area
Sample reservoir

## Slide 20: Sample Moved to Test Ring
Test area
Reagent addition and mixing area
Sample reservoir

## Slide 21: Sample Testing
By subjecting the pendant drop of blood to external vibration through resonant frequency, the device is able to monitor the harmonic motion of the blood drop to detect clotting and lysis.

## Slide 22: Resonant Frequency
X X X X X
Hz
X
X X
XX
X X XX X X X X
Liquid State Gel-Like State
Modulus of elasticity and resonant frequency increase as blood clots and is plotted graphically to create the tracing

## Slide 23: TEG® 6s Test Results
R | Reaction time. The amount of time between the start of the test and the beginning of coagulation.
MA | Maximum Amplitude. The ultimate strength of the clot.
LY30 | Percent lysis at 30 minutes after MA is achieved. The rate of clot breakdown.
Not all parameters are reported on all cartridges. Refer to the TEG 6s User Manual and/or the Package Inserts for completed information.

## Slide 24
Starting the analyzer
Power on
Login
Home
Settings

## Slide 25: Power On
If OFF, move power switch at rear of analyzer to “I” position
Wait for Power On Self Test to complete

## Slide 26: Login
Enter Username first, then press return key or touch Password field
Enter password, then press login
Return key

## Slide 27: Home Screen
View list of stored
tests
Run a new
test
Set date and time (Administrator only)
Logout

## Slide 28: Settings*
If using TEG Manager software, do not update password on device. Only update password in TEG Manager.
*Note: All Settings options only available to Administrators only.
QC settings are grayed out and cannot be configured via the analyzer if TEG Manager® software is configured, in which case use TEG Manager to change them.

## Slide 29: Settings- QC Settings*
Cartridge Specific QC- When enabled, the analyzer prevents patient tests using cartridges that have either not passed QC or that have expired based on the QC Frequency setting.
Lot Specific QC- When enabled, a cartridge manufacturing lot must pass all applicable QC levels prior to allowing patient tests to be run using cartridges from that lot.
*Note: Option only available to Administrators.

## Slide 30: Settings- QC Status*
Indicates QC Status is failed, expired, or incomplete.
Status of each cartridge regardless of the QC mode or product code enabled.
*Note: Option only available to Administrators.

## Slide 31: Settings- Session Configuration*
*Note: Option only available to Administrators.
When enabled, an alert is displayed when the analyzer has exceeded 90% of its storage capacity for patient tests, QC tests, or a combination of both.

## Slide 32
Running a patient sample

## Slide 33: New Test
From the “Home” screen, press New Test

## Slide 34: Select Patient
to add a new patient.
Select the patient ID from the list or press Then press Next.
Add new patient
Scroll bar to view patient list
Select correct
patient

## Slide 35: Add Patient
For a new patient, enter a patient ID or scan patient barcode. Press Confirm.
Select Confirm
Patient ID

## Slide 36: Confirming Patient
Return to selection screen with confirmed patient highlighted.
The press Next.

## Slide 37: Program Accession Number
Manually enter, or scan the tube barcode, to enter Accession Number.

## Slide 38: Confirm Patient Data
Select Confirm
If ADT Feed or ADT Query is implemented, the analyzer will extract patient identifying information.
Confirm if patient details are correct.

## Slide 39: Preparing Test
Barcode identifies cartridge type, test type, expiration date, and lot number
NOTE: Cartridge is inserted prior to addition of blood
When cartridge port is illuminated remove cartridge from sealed pouch.
Insert cartridge into slot (barcode on left side).

## Slide 40: Verify cartridge
Verify cartridge and sample type
Barcode information downloaded and cartridge tested.
Verify correct cartridge and sample type, then press Next.

## Slide 41: Enter Test Information
Optional text field
Optional – Enter information that will identify test.
Do not enter any PHI into this field.
Press Next.

## Slide 42: Load Sample
Arrow indicating fill line
Add blood to or above FILL line
Press Start Test
The sample will automatically:
  Be drawn in
  Mixed with reagents
  Initiate testing

## Slide 43: Sealed Cartridge
Sample reservoir emptied
Waste reservoir
Reagent mixing and testing occurs automatically within sealed cartridge.
Excess blood moved to sealed waste area in cartridge.

## Slide 44
Test completion

## Slide 45: Stopping a Sample Prematurely
Note: Only finalized results will be saved.
Patient and QC tests can be stopped prematurely if desired.
To stop early: press Stop from the results screen.
Choose Acknowledge to proceed with test termination or Cancel to continue testing.

## Slide 46: Finishing Test
Cartridge Port Light Flashing
Patient and QC tests stop automatically when all parameters finalize
Print if desired
Select Done
When prompted Remove Cartridge

## Slide 47
Results display
Display
Stored test
Icons

## Slide 48: Display
Date and time
Run time of test
Patient Identification
The results shown are from item 07-601-US which is indicated for use in cardiovascular surgery and cardiology procedures
Cartridge type

## Slide 49: Display: Results
Test Type
Parameter Units
Test results
The results shown are from item 07-601-US which is indicated for use in cardiovascular surgery and cardiology procedures
Print Results
Out of range warning
Reference range
Press to see tracings

## Slide 50: Display: Overlay Tracings
Color key for tracings. Touch to go to individual tracing
Return to results table
The results shown are from item 07-601-US which is indicated for use in cardiovascular surgery and cardiology procedures
Visual comparison of tracings
Press to cycle through tracing displays
Overlaid (all)
Offset (all)
Individual with numerical data

## Slide 51: Display: Offset Tracings
The results shown are from item 07-601-US which is indicated for use in cardiovascular surgery and cardiology procedures
Provides a better view when tracing shapes are similar

## Slide 52: Display: Kaolin Tracing
The results shown are from item 07-601-US which is indicated for use in cardiovascular surgery and cardiology procedures.

## Slide 53: Stored Test/QC Screens
Patient tests
QC tests
05/08/2014 10:39 AM Patient 1 CM User 1

## Slide 54: Printout on Thermal Printer
Paper width 76 mm (3.0 in), variable length

## Slide 55
Running Quality Control (QC) Reagents

## Slide 56: Electronic Analyzer QC
The TEG® 6s analyzer and reagents have been designed and manufactured following ISO 14971 guidelines.
The TEG® 6s has robust self-monitoring and quality control mechanisms.
QC self-checks validated the analyzer’s internal components are functioning correctly, and occur:
  Every time the analyzer is powered on (Power On Self-Test, or POST)
  Every time a cartridge is inserted
  Every time a test is run
Through the incorporation of these three levels of internal QC, all essential functions of the analyzer are checked and verified without the need for external QC materials
Any operational issues are automatically detected and an error or warning is displayed to notify you.
Provides confidence that every time the analyzer is used, it is performing in a consistent manner and producing reliable results.

## Slide 57: Reagent Quality Control (QC)
Reagent cartridges are manufactured under current Good Manufacturing Practices (GMP) with a validated manufacturing process with in-process controls to ensure lot-to-lot reproducibility and product performance.
Haemonetics recommends that two levels of QC testing be performed for each new shipment of assay cartridges.
Additional QC checks may be performed at predetermined intervals based on the laboratory’s quality control policies. (NOTE: Manufacturers IQCP aid is available upon request)

## Slide 58: Quality Control (QC)
*Normal and abnormal blood supplied by user from donors who meet the institution’s criteria.
Each type of cartridge can be quality controlled by conducting two levels of QC. These include running two levels of commercially available QC material or donor blood samples.
Global Hemostasis (07-601-US, 07-604-US & 07-605-US)
    Level 1 QC Reagent kit (07-664-US)
    Level 2 QC Regent kit (07-665-US)
PlateletMapping® (07-615-US & 07-614-US)
    Normal donor blood*
    Abnormal donor blood*/ AntiPlatelet Reagent (07-670-US)**
**A General Purpose Reagent available for user-specific application.

## Slide 59: Preparing Quality Control Material
Each QC Reagent kit contains:
12 vials of lyophilized QC reagent
12 vials of DI water
Remove from refrigerator and allow to come to room temperature (10 mins)
Vial of water + Vial of QC material from kit
Citrate – Global Hemostasis cartridge + pipette
Required materials

## Slide 60: Preparing Quality Control Material
QC reconstitution
reconstitute
shake
rest
shake
rest
Remove stopper with care, metal seal may have sharp edges.
Tap the water on counter to ensure not water in cap.
Uncap and slowly pour into vial. All water must go into vial. Reinsert stopper and shake vigorously, let stand for 5 minutes then shake vigorously again and let stand for another 5 minutes.
Step 1 Step 2 Step 3 Step 4 Step 5

## Slide 61: Running Cartridge Reagent QC (07-601-US, 07-604-US or 07-605-US)
Remove desired cartridge from sealed pouch
From the Home screen, press New QC
Do NOT select “New Test” or you will not receive an MA value when running QC
Add any desired ID info on the following “Test Information” screen

## Slide 62: New QC- QC Level Screen
Select the QC test that you wish to run and then touch Next.
Select to run Proficiency/Validation samples and then touch Next.
Note: The “Proficiency/Validation” checkbox is disabled if a QC test previously failed on the analyzer. To re-enable the checkbox, run the applicable QC test.

## Slide 63: Insert Cartridge
Insert cartridge into slot (barcode on left side)
Barcode identifies cartridge type, test type, expiration date and lot number
Press Acknowledge to confirm that QC is being run on patient cartridge

## Slide 64: Start Test
Add prepared QC reagent or normal blood material to cartridge, filling to or above line
Press Start Test
Arrow indicating fill line

## Slide 65: Finish QCTest
When prompted, remove cartridge and properly dispose
View or print test results
Press Done to return Home
Indicates Pass
Embedded Reference Ranges*
*Note: Embedded Reference Ranges only available for Citrated cartridges and show when Level 1/Level 2 is selected.

## Slide 66: Proficiency/Validation Test
Select when running Validation or Proficiency samples
Does not include reference ranges, analytical measurement ranges (AMRs), or pass/fail status.

## Slide 67
Troubleshooting and maintenance

## Slide 68: Status Icon Definitions
Remote user
Exclusive mode
Saving data
Critical error
Error
Warning
Data being collected
Test Timeout
QC Overdue

## Slide 69: Errors
Make note of the error code
Follow instructions in error and press Acknowledge

## Slide 70: Errors & Warnings During Testing
An error causes the test to stop and may require you to replace the cartridge or to turn the power off and back on in an attempt to clear the error.
A warning does not require you to take a specific action, and the test continues to run.

## Slide 71: Critical Alerts
System encounters an unrecoverable error
Note error code and turn analyzer off and restart
If alert does not clear, contact customer support

## Slide 72: Cleaning the Analyzer
Wipe down analyzer surface weekly or as needed
    Disposable gloves
    70 – 95% isopropyl alcohol plus 2” X 2” Gauze pads or germicidal disposable wipes

## Slide 73
Thank you