REAGENT QUALITY CONTROL

Principle

Good laboratory practice requires reagents be tested on the day of use to ensure that the products are reactive and specific.  Any deviation from the expected result warrants investigation prior to use.  Reagents are tested when first placed in service.  When opening a new reagent bottle, write the date opened and initial the bottle.  When new lot numbers of reagents are received, a note should be left either on the board or on the QC log so the technologist performing the QC the next day can QC the new lot number.  The exception to this rule is when you QC a lot number immediately and place it into service.  The new lot number is electronically added to the Rack, tested and results entered in the computer.

QC for RACKS 1, 2 and 3 are performed daily on first shift.  QC for the reagents on Rack 4-9 will be performed when needed for use in patient testing. When QC is completed on a reagent from Rack 4-9, place your initials on the Monthly Reagent Clipboard to show that QC has been completed for that reagent.  Rack 9 is used for BIO-RAD reagents.  Rack 10 is used for new shipments.

Clinical Significance

In order to ensure that the patient results are accurate you must be able to assure the reagents and equipment you are using will perform as expected.  This must be done prior to reporting any patient results.

Reagents

Anti-Human Globulins

Anti-Human Globulin Anti-IgG

BIO-RAD Anti-IgG Rabbit

Anti-Human Globulin Polyspecific Anti-IgG,-C3d

BIO-RAD AHG Anti-IgG + Anti-C3d blended Rabitt and Murine                             Monoclonal

Anti-Human Globulin Anti-C3b-C3d

Immucor/Gamma Gamma-clone Anti-C3b-C3d (Murine Monoclonal)

Ortho BioClone Anti-C3b-C3d (Murine Monoclonal)

Blood Grouping Reagents

A1, A2, B Reagent Red Blood Cells

BIO-RAD Biotestcell reagent Red Blood Cells A1, A2, B

Anti A, Anti B

BIO-RAD Seraclone Murine IgM Monoclonal

Anti A,B

Ortho BioClone (Murine Monoclonal Blend)

Anti-A1 Lectin (Dolichos biflorus) blood grouping lectin

Immucor/Gamma

Ortho

Anti D

BIO_RAD Seraclone Anti-D Blend Human IgG/IgM Blended Monoclonal                   Anti-D (Detects weak D))

Anti-C, Anti-c,

BIO-RAD Seraclone Human IgM Monoclonal

Anti-E, Anti-e

BIO-RAD Seraclone Human IgM Blended Monoclonal

Anti-Cw

Immucor/Gamma (By Modified Tube Test)

Anti-Fya

BIO-RAD Seraclone Human IgG monoclonal

Anti-Fyb

Immucor/Gamma (By Indirect Antiglobulin Test)

Anti-Jka, Anti-Jkb

BIO-RAD Seraclone Human IgM Monoclonal

Anti-K

BIO-RAD Seraclone Human IgM Monoclonal

Anti-Kpa

Immucor/Gamma (By Indirect Antiglobulin Test)

Anti-Lea, Anti-Leb

Immucor/Gamma Gamma-clone (Murine Monoclonal)

Anti-M, Anti-N

BIO-RAD Seraclone Murine IgM Monoclonal

Anti-P1, Anti-S

BIO-RAD Seraclone IgM Monoclonal

Anti-s

BIO-RAD Seraclone Murine IgG Monoclonal

Antibody Identification by Gel Reagents

Anti-IgG Cards (Rabbit)

Micro Typing Systems

MTS Diluent

Micro Typing Systems

Selectogen Gel Screen

Ortho

Resolve .8% Reagent Cells Gel Panel

Ortho

Antibody Identification by Tube Reagents

Bovine Albumin Immucor/Gamma 22% solution Serological Reagent

Ortho Antibody Enhancement Solution

BIO-RAD Biotestcell 1&2 (tube screen)BIO-RAD Biotestcell 1 11 Eleven Cell Antibody Identification Panel (tube panel)

ELU-KIT II Immucor/Gamma

Warm Autoantibody Removal Medium (W.A.R.M.) Immucor/Gamma

Control Cells

BIO-RAD Coombscell-E IgG Coated RBC	Immucor/Gamma Compliment Control Cells

Fetal Bleed Screening Test

Immucor/Gamma

Fetal Hemoglobin Stain

Sure-Tech

Canine Coombs Reagents

VMRD

corQC Test System

Immucor/Gamma

INTERNALQC

ARC rbc used for internal QC

Instrumentation/Equipment

Incubators – Gel and Tube

Centrifuges – Gel and Cell Washer / Immufuges

Impact Pipette / BioHit Pipette/ Finnpipette

Procedure: Inspect all reagents being used for evidence of contamination or deterioration (i.e. marked turbidity of the antiserum, hemolysis of the red cells).  Check the IgG card for signs of discoloration, bubbles, crystals, drying (a liquid layer should appear on top of the Gel in each microtube), or seals which appear damaged or opened.

- In the Patient Order Module, Click “File” from the menu and “Open”, “Reagent Rack” from the drop down boxes.  Click “Query”.

- Select the rack you are testing and click OK.  Check the lot numbers, expiration dates and open dates of all the reagents.

- If you have a new lot number of reagent to add to the Rack, add a row to the Reagent list by right clicking in the Reagent ID box.

- Right click again and click Find.  The Find Reagent Lot Number Box will pop up.  Click the query button, and select your reagent from the drop down box.  Click OK.

- Enter the open date of the reagent into the computer.  Write the date and your initials on the vial.

- To remove an expired reagent or incorrect reagent from your RACK, right click on the row to remove and select “Delete Row” from the drop down box.  Save.

- In the Reagent Order Module, Click “File” from the menu and “Open”, “Reagent Lot”.  Enter the lot number of incorrect reagent.  Click “Query”.

- Change status to “R” Recalled.  Save.

Reporting Results:

- To Enter QC results, Select “File” from the Menu and “New”, “Reagent Test Batch” from the drop down boxes.

- Enter the Rack number that you are testing or click Find in the Rack ID box and select Rack from the drop down box.  Insure the “Include all reagents” button is checked and Click OK.  The “Reagent Test Batch Window” opens.

- Click the “Delete” icon to remove reagents from your rack that you will not be testing.  This will not delete them from the Rack, but will delete them from your test profile.

- Check the “Result and Go to Result Entry Box” box.

- Enter reactions and interpretations into the computer.

- For reagents that require visual checks ensure that the lot number of the reagent being used is the lot number that is listed on your rack.  Verify lot number and expiration and report interpretation as pass.

- For reactions that require check cells, enter the check cell reaction strength into the comment line of the test. (such as 2+cc)

- Any deviations from the expected results should be brought to the attention of the team leader.  This discrepancy must be resolved prior to using the equipment or reagents for patient testing.

RACK 1 & 2:

- Label test tubes 1-14.

- To tubes 1 and 2 (QCCC) add 1 drop of polyspecific antihuman globulin.

- To tube 3 and 4 (QCAntiA) add 1 drop of Anti-A.

- To tubes 5 and 6 (QCAntiB) add 1 drop of Anti-B.

- To tubes 7 and 8 (QCAntiAB) add I drop of Anti-A,B.

- To tubes 9 and 10 (QCAnti-D) add 1 drop of Anti-D.

- To tubes 11 and 12 (QCA) add 1 drop of A1 cells.

- To tubes 13 and 14 (QCB) add 1 drop of B cells.

- Add 1 drop of Combscell-E IgG coated red blood cells to tube 1.

- Add 1 drop of CorQC reagent cells (group AB, D positive) to tubes 3, 5, 7 and 9.

- Add 1 drop of CorQC reagent antisera to tubes 11 and 13.

- Add 1 drop of Internal QC (ARC group O negative cells) to tubes 2, 4, 6, 8 and 10.

- Add 1 drop of Anti-B to tube 12.

- Add 1 drop of Anti-A to tube 14.

- Mix and centrifuge all tubes for 15 seconds at 3400 RPM.  Examine macroscopically for agglutination.  Grade and enter reactions and interpretations into the blood bank computer system.

- Discard all tubes except tube 10.

- Incubate tube 10 for 15 minutes at 37 degrees.  After incubation, wash 3 times with normal saline.  Add 2 drops of Anti-IgG, mix well and centrifuge for 15 seconds at 3400 RPM.  Examine macroscopically for agglutination.

- Grade and enter the reaction, confirming the validity of the negative test tube 10 with 1 drop of Combscell-E IgG coated RBCs.  Centrifuge and record results in the blood bank computer system.

- CorQC is a visual check only.

- Save all results and discard tubes.

Note:

- Record all negative results on the comment line of the corresponding

reagent (there is a formatted comment under results).

- Anti-D does not need a negative comment.

- Remember to record the check cell results for tube 10.

RACK 3:

- Label test tubes 1-8.

- To tubes 1 and 2, add 1 drop of Anti-A from the first vial of Anti-A in rack 3.

- To tubes 3 and 4, add 1 drop of Anti-A from the second vial of Anti-A in rack 3.

- To tubes 5 and 6, add 1 drop of Anti-B from the first vial of Anti-B in rack 3.

- To tubes 7 and 8, add 1 drop of Anti-B from the second vial of Anti-B in rack 3.

- Add 1 drop of internal QC (ARC group AB positive cells) to tubes 1, 3, 5 and 7.

- Add 1 drop of internal QC (ARC group O negative cells) to tubes 2, 4, 6 and 8.

- Mix and centrifuge for 15 seconds at 3400 RPM.  Examine macroscopically for agglutination.

- Grade and enter reactions and interpretations in the blood bank computer system.

- Discard tubes.

- (QCRBC) is a visual check of the RBC unit number being used as an internal control.

Note:

- Record all negative results on the comment line of the corresponding reagent ( there is a formatted comment under results).

For the Gel QC:  QCGSC &QCcard

- Label 2 wells of the MTS Gel Card I and II (QCGSC).

- Remove the foil seal from the 2 wells to be used.

- Using the Impact / Biohit pipette or the Finn pipette:  Add 50 microliters of 0.8% cell suspensions of the Selectogen Gel Screening Cells into the appropriately labeled wells.

- Add 25 microliters of corQC reagent antiserum into each of the wells.

- Place the card in the Gel 37degree Incubator for 15 minutes.  (Not to exceed 40 minutes)

- After incubation, place the card in the Gel Centrifuge and spin for 10 minutes at 900 RPM.

- Read the front and back of each microtube macroscopically and record the reactions. For the Negative Control it is acceptable to use the first patient who exhibits a negative antibody screen test as your negative result.

- (QCcard) is a visual check.

- Save.

RACKS 4-8:  Performed only when needed for patient testing

1.	Carefully inspect the antiserum you are using.  Note the manufacturer and type of antiserum (monoclonal, series 1, series 5, gamma-clone, etc).

2.	Choose the corresponding reagent and RACK number.  Follow the instructions on the reagent insert for the correct procedure. (Immediate spin, 37C, Room Temp, IAT, etc.)

3.  Enter reactions and interpretations for the positive and negative controls into the computer.

4.  Testing phases performed and reported should be equivalent if the correct reagent is chosen from the reagent rack.

RACK 4:

QC for Diluent:  (QCDIL)

- Prime the MTS diluent dispenser so there is no evidence of bubbles.

- Dispense 1ml of diluent into a properly labeled tube.

- Add 10ul of donor packed red blood cells.  Mix.

- Label 2 wells on a MTS Gel Card as Positive and Negative.

- Remove the foil seal from the 2 wells and pipette 50ul of the donor unit cell suspension into both of the wells.

- Add 25 microliter of corQC reagent antiserum to the positive well.

- Add 25 microliter of saline to the negative well.

- Place the card in the Gel 37degree Incubator for 15 minutes.  (Not to exceed 40 minutes)

- After incubation, place the card in the Gel Centrifuge and spin for 10 minutes at 900 RPM.

- Read the front and back of each microtube macroscopically and record the reactions.

- Type the unit number of packed red blood cells used into the comment field of QCDIL.  Save.

QC for Tube screen: (QCTSC  &  QCalbumin) /or QC OAES (Ortho Antibdy Enhancement Solution)

QCTSC:

- Label tubes #1-4.

- Add 1 drop of Biotestcell reagent cells 1 to tubes 1 and 2.

- Add 1 drop of  Biotestcell reagent cells 2 to tubes 3 and 4.

- Add 1 drop of corQC reagent antiserum to tubes 1 and 3 and 2 drops of saline to tubes 2 and 4.  Mix

- Centrifuge for 15 seconds at 3400 RPM.  Examine macroscopically for agglutination and record results.

- Add 2 drops of Albumin/OAES.

- Mix and incubate at 37 degrees for 15 minutes for albumin or 10 minutes for OAES.

- Centrifuge for 15 seconds at 3400 RPM.  Examine macroscopically for agglutination and record results.

- Place tubes in the cell-washer and wash 3 times with saline.  Add 2 drops of Anti-IgG antiserum to each tube.  Mix and centrifuge.  Examine macroscopically for agglutination.

- Enter the reactions and interpretations into the computer.

- QCalbumin/OAES is a Visual check.

- Record all negative results on the comment line of the corresponding reagent (there is a formatted comment under results).

- Save.

QC for Compliment Check Cells: (QCCOMPCC & QCcompl)

QCCOMPCC:

- Add 1 drop of Anti-C3b,C3d  anti-human globulin to a properly labeled tube.

- Add 1 drop of compliment control cells.  Mix and let sit for 5 minutes.

- Centrifuge for 15 seconds at 3400 RPM.  Examine macroscopically for agglutination.

- Enter the reaction and interpretation into the computer for (QCCOMPCC).

- QCcompl is a Visual check.

- Save.

Rack 5:

Anti-C, Immucor monoclonal Gamma-clone (QCANTI C)

Anti-c, Immucor series 1 (QCANTIcS1)

Anti-Cw, Immucor (QCANTICw)

Anti-Kpa, Immucor (QCANTIKpa)

Rack 6:

Anti-Fyb, Immucor (QCANTIFyb)

Anti-Jka, Immucor (QCANTIJka)

Rack 7:

Anti-Lea, Immucor Gammaclone (QCANTILea)

Anti-Leb, Immucor Gammaclone (QCANTILeb)

For subgroups of A:  QCAntiA1L & QCA2

- (QCAntiA1L)  Label two tubes as Positive and Negative.

- Add one drop of Anti-A1 Lectin to each tube.

- Add one drop of A1 red cells to Positive control tube.

- Add one drop of A2 red cells to Negative control tube.

- Mix and centrifuge for 15 seconds at 3400 rpm.

- Re-suspend the packed cells by gentle agitation and examine for agglutination.

- Record result.

- (QC-A2) is a visual check

- Save.

Rack 8:

Visual checks that require the lot number of reagent is checked to insure the correct lot number is being used for testing.

Elution Kit II, Immucor (QCelute) (This consists of the lot number on the kit.)

Canine Coombs Reagent, VMRD (QCCCR)

W.A.R.M., Immucor (QCWARM)

Fetal Hemoglobin Stain, Sure-Tech (QCFStain)

Rack 9:

Anti-C, BIO-RAD Seraclone Anti-C (BR Anti C)

Anti-c, BIO-RAD Seraclone Anti-c (BR Anti c)

Anti-E, BIO-RAD Seraclone Anti-E (BR Anti E)

Anti-e, BIO-RAD Seraclone Anti-e (BR Anti e)

Anti-K, BIO-RAD Seraclone Anti-K (BR Anti K)

Anti-Jka, BIO-RAD Seraclone Anti-Jka (BR Anti Jka)

Anti-Jkb, BIO-RAD Seraclone Anti-Jkb (BR Anti Jkb)

Anti-M, BIO-RAD Seraclone Anti-M (BR Anti M)

Anti-N, BIO-RAD Seraclone Anti-N (BR Anti N)

Anti-S, BIO-RAD Seraclone Anti-S (BR Anti S)

Anti-s, BIO-RAD Seraclone Anti-s (BR Anti s)

Anti-P1, BIO-RAD Seraclone Anti-P1 (BR Anti P!)

Anti-Fya, BIO-RAD Seraclone Anti-Fya (BRFya)

Procedural Notes/Problem-Solving Tips

- A discrepancy in the AB and Rh antiserum.

- May be due to a loss of reactivity.  To determine if this is the case, repeat testing using a different vial of the same lot number.  If this doesn’t resolve the discrepancy repeat testing using a different lot number of antiserum.

- May be due to under-centrifugation or over-centrifugation.  Repeat testing taking care to observe centrifugation times indicated by last centrifuge calibration. (This can be found on the side of each centrifuge)  If discrepancies are still noted, recalibrate the centrifuge.

- A discrepancy with A1 or B Cells

- May be due to loss of antigenicity.  Repeat testing using different A1 or B Cells of the same lot number. If a new lot number is not available, or is discrepancy is still seen, repeat testing using a different source of A1 or B Cells.  To enhance reactivity you may also incubate the A1 and B cells at room temperature for 15 to 30 minutes.

- May be due to under-centrifugation or over-centrifugation.  See steps to follow under #1 in resolving discrepancies in the AB and Rh antiserum.

- A discrepancy involving tubes that are incubated.

- May be due to improper incubation.  Repeat testing using a different 37C incubator.  If results are acceptable, check original incubator for fluctuations in temperature. (Hot/Cold spots).  Mark those wells in the incubator that should not be used for testing, or take the incubator out of service.

- A discrepancy with the Screening Cells.

- May be due to loss of antigenicity of the Screening Cell(s).  Repeat testing using a different vial of the Screening Cell.  There will be times when we have access to another lot number of Screening cells, repeat testing with these cells if using a new vial of the same lot number does not resolve the discrepancy.  If the discrepancy is still not resolved repeat testing using panel cells or cells known to possess antigen corresponding to IgG antibody specificity. (Anti-D and Anti-c).  If the discrepancy still exists see how to resolve discrepancies due to Anti AHG reagent.

- A discrepancy seen in tubes containing Anti-human serum (IgG).

- May be due to improper washing or neutralization of the serum.

- If automated washing was used – Check the saline source and lines to ensure proper flow.  Check washer on multiple cycles to ensure proper fill of all tubes.  Correct any problems encountered and repeat testing.

- If manual washing was used, repeat testing taking care to ensure all cell buttons are completely re-suspended after each wash and decanting of saline is complete after last wash.

- If your reactions are weaker than those normally seen, (or not reacting at all) check the Coombs serum by testing the original vial and a second vial of the same lot number as follow: Add one drop of Coombs control cells (CC’s) to two drops of Anti- AHG.   Centrifuge and read.  Compare the two sets of results.  If the reactions with the two vials differ – with the original vial used having a lesser strength of reactivity the probable cause is neutralization of the vial. Repeat testing using a new vial of anti–AHG.  If the reactions are essentially the same with both vials, the probable cause is loss of reactivity of the entire lot of anti-AHG.  Repeat testing using a different lot of Anti-AHG.

- A discrepancy seen with Coombs Control Cells.

- May be due to loss of reactivity of the cells.  Prepare your own IgG sensitized cells and repeat test using these cells for control phase.

- If any antiserum or cells are found to be inadequate to use, they should be removed from the racks and returned to the company.  Any reagent returned to the vendor or taken out of service for ineffectiveness should be recorded on the Reagent Deviations Log.

- When testing kits with multiple reagents the reagents in each kit will not be used in conjunction with another lot number of the same kit.  If one of the reagents in the kit runs out, or fails to perform as expected, the whole kit will be discarded if another kit with the same lot number is not available.  The new kit will be placed in use, and QC will be performed at the time of use of this new kit.

- All kits will be run in tandem with positive and negative controls from the current kit, when first received and prior to first use for patient testing.

References

Gamma QC Test Systems

Ortho Reagent Inserts

Immucor/Gamma Reagent Inserts

MMCI Blood Bank QC Manual / Mary McCallister

Author:  Jane Harrington MLT.  09-03-2010

Medical Director Approval: Elizabeth Bauer-Marsh, MD

| REVISION HISTORY (began tracking 2012) |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 0.1 | Removed old reagents and added BIO-RAD reagents. | Kathy Turpin | 05/29/2014 |
| 0.2 | Added negative controls to the QC procedures for Racks 1,2 and 4 | Kathy Turpin | 08/22/2014 |

Reviewed

| Reviewed by | Date | Coordinator/Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
| K. Maher | 1/16/12 |  | 1/16/12 |  | 1/16/12 |
|  |  |  | 8/23/14 |  | 8/29/14 |