Remel VRE Screening

BHI agar w/vancomycin

- PRINCIPLE

The recommended medium for agar screening is BHI agar.  This medium contains meat                           infusion, peptone, and dextrose which provide nitrogen, carbon, sulfur, vitamins, and a                          carbohydrate source.  Agar is a solidifying agent.  BHI agar is supplemented with 6 μg/ml of                  vancomycin to detect resistant strains of Enterococcus sp.

- CLINICAL SIGNIFICANCE

Vancomycin resistant Enterococcus can account for high morbidity and mortality since few

other agents are effective in treating the organism.  Because of this, the detection of this

organism from clinical specimens must be detected.  Surveillance of these patients for

colonization or infection is accomplished by using this agar.

- SPECIMEN

Any plated culture that contains raised, gray gamma hemolytic colonies that are gram positive                cocci, catalase negative, and PYR positive.

- REAGENT

BHI infusion agar with vancomycin

Tryptic soy broth (TSB) broth

- INSTRUMENTATION/EQUIPMENT

35° ambient incubator

1 μl disposable loop (blue loop)

- QUALITY CONTROL

All lot numbers of BHI agar w/ vancomycin have been tested using the following quality                        control organisms and have been found to be acceptable.  This quality control testing                              conforms with or exceeds CLSI standards.  Controls should be included each time a                                susceptibility test is performed or weekly if satisfactory performance can be documented                         according to CLSI standards.  If aberrant quality control results are noted, patient results should not be reported.

CONTROL				INCUBATION		RESULTS

Enterococcus faecalis		Ambient, 24 h @		No growth

ATCC 29212				35 ± 2° C

Enterococcus faecalis		Ambient, 24 h @		Growth

ATCC 51299				35 ± 2° C

- PROCEDURE:

- Obtain a pure, 18-24 hour old culture of enterococcus.

- Prepare a suspension of the culture in tryptic soy broth and adjust to the turbidity of a          McFarland 0.5 standard or equivalent visually or by utilizing a photometric device.

- Inoculate the BHI agar plate with the adjusted organism suspension using a 1 μl                   calibrated loop (blue loop).

- Spread the inoculum over a 10-15 mm area and allow to absorb into the agar surface.

- Invert the plate and incubate in ambient air at 35 ± 2° C for a full 24 hours prior to              interpretation.

- After incubation, examine plate for evidence of small colonies (>1 colony) or a film of        growth.

- REPORTING RESULTS

Growth - Resistant (good growth, light growth, or >1 individual colony), VRE

No Growth – Susceptible, not VRE

- PROCEDURAL NOTES/PROBLEM-SOLVING TIPS

- This product is a screening system for use in the detection of vancomycin-resistant              enterococcus.  Data is not available for other organisms.

- A blood agar plate may be used as a growth control to monitor organism viability and         purity.

- Susceptibility testing does not need to be performed on VRE as this is a screening.

- REFERENCES

- Leclercq, R., E. Derlot, J. Duval, and P. Courvalin. 1988 . N Eng. J. Med. 319:157-161.

- Clark, N.C.,R.C.Cooksey, B.C. Hill, J.M. Swenson, and F.C. Tenover. 1993.                       Antimicrobial Agents and Chemother. 37: 2311-2317.

- Wiley, B.M., B.N. Kreiswirth, A.E. Simor, G. Williams, S.R. Scriver, A. Phillips, and         D.E. Low. 1992. J. Clin, Microbiol. 30: 1621-1624.

- Clinical and Laboratory Standards Institute (CLSI). 2009.  Methods for Dilution                  Antimicrobial Susceptibility Tests for Bacteria that Grow Aerobically; Approved                Standard, 8th ed. M7-A8.  CLSI, Wayne, PA.

MMCI Laboratory is a CAP accredited facility, as of 7/1/11 the responsibility of new and/or substantially revised policies and procedures will be restricted the Laboratory Director whose name appears on the CLIA certificate, whose signature appears below. The biennial review will be completed by the Administrative Director.

| POLICY CREATION : | Date |
| --- | --- |
| Author: Teresa Nuese | Date: October 26, 2017 |
| Medical Director: Elizabeth A. Bauer-Marsh, MD | Date: October 31, 2017 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| October 31, 2017 | Elizabeth A. Bauer-Marsh, M.D. |  |
| SECTION MEDICAL DIRECTOR |  |  |
| October 31, 2017 | Lori Racsa, D.O. |  |

| REVISION HISTORY (began tracking 2011) |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Initial Release | M. Bishoff/ T. Nuese | 11/2/17 |

REVIEWED BY

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
| Marsha Bishoff | 11/2/17 |  | 10/26/17 11/3/17 |  | 11/3/17 |