CORD BLOOD TESTING

Principle

Cord bloods are sent down to the laboratory on all newborn babies.  Some will have an order for an ABO, Rh typing and direct coombs.  Others are stored just in case testing is needed at a later time. Typing can be done to assist in determining if an Rh negative mother needs Rh immune globulin or to assist in the diagnosis of Hemolytic Disease of the newborn (HDN).  In the case of HDN, a direct coombs is also helpful.

Specimen

- Umbilical cord blood, anticoagulated (EDTA) specimen, if available.

- All specimens are kept for 10 days stored at 2-8˚C.

- Testing should be performed within 72hrs of collection.

Reagents

| Reagent | Storage | Stability |
| --- | --- | --- |
| Isotonic Saline, NERL | 15-30°C | Unopened: marked expiration Opened: 1 month |
| Monoclonal Anti-A antisera | 2-8°C | Stable through marked expiration |
| Monoclonal Anti-B antisera | 2-8°C | Stable through marked expiration |
| Anti-D sera | 2-8°C | Stable through marked expiration |
| Anti-IgG,-C3d; Polyspecific, BIO-RAD | 2-8°C | Stable through marked expiration |
| Coombscells-E, BIO-RAD | 2-8°C | Stable through marked expiration |

Instrumentation/Equipment

- 10 x 75 mm Tubes

- Centrifuge

- Pipettes

Quality Control

To be performed once per day of testing.  See Blood Bank Quality Control Procedure for more information.

| Reagent to be Tested | Positive Reactivity | Negative Reactivity |
| --- | --- | --- |
| Anti-A | CorQC Reagent Cells | ARC group O negative cells |
| Anti-B | CorQC Reagent Cells | ARC group O negative cells |
| Anti-D | CorQC Reagent Cells | ARC group O negative cells |
| Anti-Human Globulin, Anti-IgG, -C3d; polyspecific | Coombscells-E | ARC group O negative cells |

Procedure

- Prepare cord blood suspension:

- Label a 10x75mm tube with the patient’s last name.

- Add a few drops of umbilical cord blood.

- Add saline to form a 3-4% suspension.

- Label tubes according to tests required.

- For ABORH testing:  label A, A, B, B, D, D, CT, CT

- For Direct Coombs (DAT) testing: label DC, DC

- Add 1 drop of cord blood suspension to each tube.

- Wash all tubes three times with saline, completely decanting last wash.

- Dispense one drop of antisera to each corresponding tube: Anti-A to A tubes, Anti-B to B tubes, Anti-D to D tubes, Saline to CT tubes, and 2 drops of Anti-Human Globulin, Anti-IgG, -C3d:polyspecific to DC tubes.

- Centrifuge each tube for 15-20 seconds at 3400 rpm.

- Read macroscopically for ABORh and macroscopically with the tube agglutination viewer for DAT.  Agglutination is considered a positive reaction.

- To negative DC tubes, add 1 drop of Coombscells-E and centrifuge 15-20 seconds at 3400 rpm.  Agglutination confirms a valid negative result.

- To non-reactive Anti-D, perform Du testing.  Refer to Testing for Du Antigen Procedure.

###### Interpretation

| Anti-A | Anti-B |  | Group Interpretation |  |  |
| --- | --- | --- | --- | --- | --- |
| NEG | NEG |  | O |  |  |
| POS | NEG |  | A |  |  |
| NEG | POS |  | B |  |  |
| POS | POS |  | AB |  |  |
| Anti-D |  | PT Control |  |  | Interpretation |
| POS |  | NEG |  |  | D POSITIVE |
| NEG |  | NEG |  |  | Test for Weak D antigen |
| POS |  | POS |  |  | Invalid result |

Reporting Results

Call the Nursery with all positive Direct Coombs results.

For Positive DAT:

- If the following conditions are found:  baby’s ABO type is incompatible with mother’s ABO type, AND mother has a negative antibody screen.

- Report the DAT positive.

- Comment “Positive DAT is probably due to maternal ABO antibody.  An elution may be ordered if more information is needed”, or appropriate comment as utilized by the LIS.

- DO NOT PERFORM ELUTION unless ordered by the physician. If elution is requested, perform the acid elution technique and test the eluate against the A1 cells, B cells as well as screening cells I & II. If screen is positive, perform antibody panel for antibody identification.

- Record mother’s type in inventory book on same line with baby’s.

- For the following:  baby and mother have same ABO type, or baby is Type O.

- Report the DAT positive.

- Comment “The antibody is NOT due to an ABO incompatibility.  It is probably due to an antibody in the mother’s serum.  An elution may be ordered if more information is needed or appropriate comment as utilized by the LIS.

For Positive DAT with a positive Du antigen*:

- If baby’s Du reaction is equal in strength to the DAT reaction and Patient Control Reaction:  Report baby as Rh NEGATIVE.

- If baby’s Du reaction is greater in strength than the DAT and Patient Control Reaction:  Report baby as Rh POSITIVE.

| Examples* |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| REACTION STRENGTHS: |  |  | REPORT AS: |  |  |
| Du | CT | DC | Pt Rh | RhIG |  |
| 1+ | 1+ | 1+ | Rh= | no |  |
| 2+ | 1+ | 1+ | Rh+ | yes |  |

*If any of these scenarios occur, repeat all testing before turning out – make sure results are reproducible.

Procedural Notes/Problem-Solving Tips

- In newborns only cell grouping is done, as the expected Anti-A and/or Anti-B do not usually appear until the infant is about 3-6 months of age.

- Cord blood samples must be washed adequately before testing because of the Wharton's jelly which can cause false-positive results.

- False-negative results of Rh typing may occur when baby's red cells are strongly sensitized with Rh antibody from the mother.  In such cases, the direct coombs test will be positive.  An elution will reveal anti-D: the baby should be considered Rh positive.

- Babies born to mothers that received ante-partum RhIG may have a positive direct coombs.  No evidence of hemolysis due to RhIG has been noted.

References

Roback, J.  Combs, M.  Grossman, B.  Hillyer, C, AABB Technical Manual, 17th Ed.  Bethesda: AABB, 2011.  P. 634.Bio-Rad Medical Diagnostics GmbH, Blood grouping reagent insert.

| POLICY CREATION : |  |
| --- | --- |
| Author: Mary McCallister | September 24, 2003 |
| Medical Director: Juan G Gonzalez, MD | October 13, 2008 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| January 1, 2017 | Elizabeth A. Bauer-Marsh, M.D. |  |

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Formatting changes, additional interpretation information, Document ID added to header. | S.Schaffer | 12/6/11 |
| 2 | Changed reagent storage temperature. Added Bio-Rad as reference. Changed AABB edition and pages. | Kathy Turpin | 07/02/13 |
| 3 | Under Reporting Results in LIS: changed how results are entered into the computer if patient is a Du negative or Du positive. Minor formatting. | Kathy Turpin | 03/11/14 |
| 3 | Changed the QC negative Reactivity reagent to ARC Group O negative cells. Under procedure number 7 removed microscopically and added tube agglutination reader. | Kathy Turpin | 12/22/14 |
| 4 | Removed all LIS specific references, allowed for LIS-specific comments regarding positive DAT on cord blood testing. | Vincent Strow | 2/10/16 |
| 5 | Removed mention of LUI elution and replaced with acid elution since LUI elution is no longer performed. Minor formatting. | Leah Miller | 1/11/18 |

REVIEWED BY

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
| K. Maher | 3/1/12 |  | 3/1/12 |  | 3/1/12 |
| K. Maher | 8/5/13 |  | 8/5/13 |  | 8/8/13 |
| D. Allen | 5/4/14 |  | 6/10/14 |  | 6/10/14 |
| D. Allen | 12/23/14 |  | 12/22/14 |  | 1/30/15 |
| V. Strow | 2/10/16 |  | 7/6/16 |  | 7/7/16 |
| Leah Miller | 1/11/18 |  | 1/11/18 |  | 1/22/18 |