Vidas® Lyme IgG/Lyme IgM

Qualitative Enzyme Immunoassay

Vidas ® 3

- PRINCIPLE

IgM II assay principle combines a 2 step enzyme immunoassay sandwich method with a final fluorescent detection (The Vidas®3 Lyme IgG II and Lyme ELFA) (see User’s Manual).  The Solid Phase Receptacle (SPR®) serves as the solid phase as well as the pipetting device for the assay.  Reagents for the assay are ready-to-use and pre-dispensed in the sealed reagent strips.  All the assay steps are performed automatically by the instrument. The reaction medium is cycled in and out of the SPR several times.  After a preliminary wash step and a sample dilution step, the antibodies to B.  burdorferi,  present in the specimen will bind to the B.  burgdorferi   specific recombinant proteins coating the interior of the SPR.  Unbound sample components are washed away.  Anti-human IgG/IgM antibodies conjugated with alkaline phosphatase, will attach to the immunocomplex bound to the SPR wall.  A final wash step removes unbound conjugate.  During the final step, the substrate (4-Methyl-umvelliferyl phosphate) is cycled in and out of the SPR.  The conjugate enzyme catalyzes the hydrolysis of this substrate into a fluorescent product (4-Methyl-umbelliferone) the fluorescence of which is measured at 450nm.  The intensity of the fluorescence is proportional to the quantity of anti-B. burgdorferi  IgG/IgM antibody present in the sample. At the end of the Vidas® Lyme IgG/IgM assay, results area automatically calculated by the instrument.  A test value is generated and a report is printed.

- CLINICAL SIGNIFICANCE

The assay is a qualitative enzyme immunoassay used in the presumptive detection of           human antibodies to Borrelia  Burgdorferi  in plasma and or serum.  Lyme disease is a tick born disease caused by the bacterium, Borrelia burgdorferi sensu stricto in the US.

The disease is seasonal (from Spring to Autumn).  It occurs in three stages, evolving from early to late disease when left untreated. Early antibiotic treatment can resolve clinical symptoms and prevent progression of disease to later stages.

- SPECIMEN

2 ml of serum collected in a serum separator tube.

It is recommended not to use samples that are hemolyzed, lipemic or icteric and, if possible,         to collect a new sample.

Testing will be performed on Tuesdays and Fridays only. (Testing maybe referred out for        confirmation, if positive, therefore serum is the only acceptable specimen.)

Specimen preparation:

All specimens should be centrifuged according to manufacturer’s recommendations.  All   specimens that have been frozen must be spun after thawing, removing excess fibrin or particulate matter, if necessary.   Frozen specimens should be mixed using a vortex type mixer.

Specimen stability:

If specimens cannot be tested on the day of collection, store them at 2-8° in stoppered tubes   for up to 6 days.  If longer storage is required, freeze the serum at -25 ±6° C for up to 6    months.

Specimens have been validated through two freeze/thaw cycles.

- REAGENTS

Strips and SPR’s specific for both IgG and IgM

Ready to use standards and controls

For a more detailed description of the contents of the strips and SPR’s please see the package insert.

Some reagents in the strips can be irritating to the eyes.  In case of contact with eyes, rinse immediately with plenty of water and seek medical advice.  For further information, refer to the SDS available on request.

- INSTRUMENTATION/EQUIPMENT

Pipette with disposable tips to dispense 100 µl.

Powderless, disposable gloves

VIDAS® 3 instrument

- CALIBRATION

Calibration, using the standard provided in the kit, must be performed each time a new lot of reagents are opened, after the master lot data has been entered.  Calibration should then be   performed every 28 days.  This operation provides instrument-specific calibration data and compensates for possible minor variations in assay signal throughout the shelf-life of the kit.

The standard, identified by S1, must be tested in duplicate, (See User’s Manual). The standard value must be within the set RFV “Relative Fluorescence Value” range.  If this is not the case, recalibrate.

The expected values for the standard are also stored using the MLE data.  If the standard values deviate from the expected values they will be flagged on the result sheet and the result should not be reported.

- QUALITY CONTROL

One positive and one negative pre-diluted control are included in each Vidas® Lyme IgM II and Lyme IgG II kits.

These controls must be tested immediately after opening a new kit to ensure that reagent performance has not been altered.  Each calibration must also be checked using these controls.  The instrument will only be able to check the control values if they are identified by C1 and C2.

Results cannot be validated if the control values deviate from the expected values. It is the           responsibility of the user to perform Quality Control in accordance with any local applicable        regulations.

*NOTE: The negative and positive controls supplied with the kit are in a buffer solution.             The positive control contains monoclonal antibodies.  The controls are intended to monitor          for substantial reagent failure, but, will not ensure precision at the assay cutoff.  The user               should use serum based controls, as necessary.

- PROCEDURE

1.        Remove the required reagents form the refrigerator and let sit at RT for 30 minutes.

2.     Use one “LYM” or “LYG” strip and one “LYM” or “LYG” SPR for each serum                      sample, control or standard to be tested.

3.     Make sure the storage pouch has been carefully resealed after the required SPRs                       have been removed.

4.        The test is identified by the “LYM” or “LYG” code on the instrument.  If a standard            is to be tested, enter “S1” for the sample ID.  The standard must be identified by                 “S1”, and tested in duplicate.  If the positive control is to be tested, it should be                     identified by “C1”.  If a negative control is to be tested, it should be identified by                “C2”.

5.     Mix the standard, controls, and samples using a vortex-type mixer, (for serum that                  has been separated from the pellet.  Test volume for all of the above is 100µl.

6.      Insert the VIDAS® “LYM” or “LYG” SPRs and the “LYM” or “LYG” strips into                 the analyzer. *Must load the strips with specimen prior to loading onto the                     analyzer when doing manual pipetting.  Check to make sure the color labels with             the assay code on the SPRs and the reagent strips match.

7.      Initiate the assay as directed in the User’s Manual or Quick guide (located on board              the analyzer or in the drawer below).  All the assay steps are performed                                 automatically by the instrument.

8.       Reclose the control and standard vials and return them to 2-8° after pipetting.

9.    The assay will be completed within approximately 27 minutes.  After the assay is                    completed, remove the SPRs and strips from the instrument.

10.     Dispose of the used SPRs and strips into an appropriate recipient.

- REPORTING RESULTS

VIDAS ® Lyme IgG II

| Index | Result | Interpretation |
| --- | --- | --- |
| Index <0.20 | Negative | No detectable anti-B.burdorferi IgG antibodies detected. |
| Index ≥0.20 | Postive | Presumptive diagnosis of anti-B. burdorferi IgG antibodies. The specimen should be tested further by an IgG Western Blot for supplementary information |

VIDAS® Lyme IgM II

| Index | Result | Interpretation |
| --- | --- | --- |
| Index <0.12 | Negative | No detectable anti-B burgdorferi IgM antibodies detected. |
| 0.12 ≤Index <0.32 | Equivocal | Imprecise detection of anti-B.burdorferi IgM antibodies. The specimen should be tested further by an IgM Western Blot for supplementary information. |
| Index ≥ 0.32 | Positive | Presumptive diagnosis of anti-B. burdorferi IgM antibodies. The specimen should be tested further by an IgM Western Blot for supplementary information |

All specimens with the presumptive results of equivocal and or positive will be referred               out for immunoblot (Western Blot) testing per CDC recommendation.  Results are                        considered positive only if the EIA and immunoblot are both positive.  Negative EIA                    samples need not be tested further.   In general, these results should be interpreted taking               into consideration the clinical presentation, the case history, and the results of any other                biological tests performed.

Enter the results into the 	LIS in the “Hematolgy Manual Result Entry” screen. Modify the results, hit the << arrow to put all patients to the left, hit the result button, scan the barcode, hit the tab key, check 2 patient identifiers prior to saving the results.

Remember to document the Lot # of the kit, expiration date and QC ok on the line below the results.

Note:  Lyme disease is a IDPH reportable disease.  All suspected positives and                       confirmed cases of this disease must be reported to the IDPH within 7 days by calling 309-679-6655, M-F between 8:00AM-4:30PM

- PROCEDURAL NOTES/PROBLEM-SOLVING TIPS

- Antibody detection methods do not provide definitive results for establishing or             ruling out a diagnosis of Lyme disease.

- The recombinant proteins used in the VIDAS® Lyme IgM II and the Lyme IgG II          assay contain antigen constituents from B. afzelii and B. garinii in addition to B.            burgdorferi. The presence of additional antigens may interfere with the                           interpretation of the 2nd tier results from Western Blots that contain antigens from          B. burgdorferi only.

- A negative result in the Lyme IgG II or the Lyme IgM II assay does not rule out              the possibility of B. burgdorferi infection in a patient.  Patients in early stages of            infection or who have undergone antibiotic therapy may not produce measurable           antibodies.  Patients with clinical history and/or symptoms suggestive of Lyme               disease, but with negatie test results, should be reported as “No detectable                      antibodies to B. burgdorferi.  A second specimen should be collected in 4-6                    weeks.

Note: It is estimated that 50% of subjects, in the primary stage of disease, antibody levels in blood remain below the detectable threshold.

- Positive results in the VIDAS® Lyme IgG II and or Lyme IgM  II assay must be             interpreted with caution.  Cross-reactivity may be observed with certain diseases.           Refer to the package insert for a complete list.

- The effects of Babesiosis, Erhlichiosis, and Rocky Mountain spotted fever                      pathologies on the Vidas® Lyme IgM II and or Lyme IgG II  performance is not             known.

- It is recommended not to use samples that are hemolyzed, lipemic or icteric and, if          possible, to collect a new sample.

- Clinical symptoms, epidemiological information and other laboratory test results            must all be considered in addition to Vidas® Lyme IgM II/Lyme IgG II assay                 results.

- Interference may be encountered with certain sera containing antibodies directed           against reagent components.  For this reason, assay results should be interpreted             taking into consideration the patient’s history,  and the results of any other tests               performed.

- Testing should be done only when exposure history, epidemiology, and clinical              symptoms suggest Lyme disease.

- Interpret test results of specimens from immunosuppressed patients with caution.

- This test is intended for qualitative determination only.  The numeric value of the           final result above the cutoff is not indicative of the amount of Borrelia burgdorferi          antibody present.  The values obtained from different Lyme disease assays cannot          be used interchangeably.

- Dispose of used or unused reagent as well as any other contaminated disposable             materials following procedures for infections or potentially infectious products.

- REFERENCES

For a complete list of literary references please see the package insert.

MMCI Laboratory is a CAP accredited facility, as of 7/1/11 the responsibility of new and/or substantially revised policies and procedures will be restricted the Laboratory Director whose name appears on the CLIA certificate, whose signature appears below. The biennial review will be completed by the Administrative Director.

Policy Created by: ______Kim Paige___________________________	 Date: 12-17-17___________

Medical Director Approval: __________________________	 Date: ________________________

Change of Medical Director: _________________________	 Date: _________________________

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 0 | Initial Release | Kim Paige |  |
| 1 | Added manual result entry procedure | Kim Paige | 1-9-18 |

Reviewed by

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |