| METHODIST CAMPUS |  |  | Page 1 of 7 | Section: UPM HEMO | Policy #: 24 |
| --- | --- | --- | --- | --- | --- |
|  | HEMATOLOGY |  | Approved by: see signature block at end of document |  | Date: 3/9/18 Review by: 3/9/20 |
|  | LABORATORY |  | Policy Created: 4/22/13 Date Revised: 3/9/18 Supersedes: 8/31/10, 12/7/10, 11/8/11, 12/10/13, 7/5/16, 9/5/16, 9/12/17, 3/9/18 |  |  |
|  |  |  | Primary Responsible Parties: Kim Paige Secondary Responsible Parties: June Bembenek |  |  |
|  |  |  | CAP Standard: NA |  |  |
| SUBJECT: |  | CSF CELL COUNT & DIFFERENTIAL - MANUAL AND AUTOMATED |  |  |  |

# PRINCIPLE

A CSF Cell Count & Differential consists of a WBC, RBC and Differential.  The report also includes color, clarity, & volume.  The counts are performed routinely by the hemacytometer methodology but can be run on the Sysmex XN-3000 if the specimen is hazy, cloudy or bloody.  A total nucleated count and red blood cell count are commonly run on the XN analyzer. A cytocentrifuged smear is stained with Wright's stain and the differential performed manually.

## CLINICAL SIGNIFICANCE

CSF examination can aid the physician in the diagnosis of infections, hemorrhage, neurologic disease, malignancy, and tumor detection.  It is also used to monitor treatment.

## POLICY SCOPE

## The scope of this policy applies to all Laboratory staff that prepares or performs testing on laboratory specimens at UnityPoint Methodist.

## SPECIMEN

Tube #3 of cerebral spinal fluid collected by lumbar puncture, unless physician requests cell counts on alternate tube.

Counts should be performed as soon as possible.

CSF stored 2-8 ºC for 7 days.

## Reagents

- Wright Stain

- Albumin

- Nerl saline

- Instrumentation/Equipment

- Hemacytometer & Cover Glass

- Capillary tubes

- Slides

- Pipettes

- Cytospin Cups & Pads

- Cytocentrifuge

- Microscope

- Warmed Box (opt.)

- Slide Stainer

- XN analyzer

- Cell counter

- Sysmex SP-10

- Wright’s stain

- Saline

- Glass test tubes

- Haematokrit 200 centrifuge

- CALIBRATION:

The automated Sysmex XN-3000 is calibrated twice yearly to ensure accuracy of the system.  An ongoing verification of calibration is done by hematology technical staff by review and documentation fo commercial control.  If a potential problem is noted, re-calibration will be performed by following the procedure located in the Sysmex “XN-3000 Instructions for Use” manual (Chapter 12).

- QUALITY CONTROL :

When running body fluid on the XN analyzer:  Quality control is performed in order to monitor analyzer performance. XN CHECK BF control levels 1 and 2 will be run by technical staff in the Manual BF mode as needed before patient samples are processed. Perform once every 24 hours of patient testing.  See UPPM HEMO procedure 20 on how to run BF controls on the XN.

- For performing manual body fluid counts:

- Level 2 Quantimetric Microscopic Urinalysis Controls, once per 8 hours of patient testing.

- Count WBCs and RBCs in 4 corner squares on both sides of hemacytometer.

- Calculate and record results on Body Fluid Control Form.

- Verify counts are within acceptable range, as established, prior to patient testing.

- Correlate number of cells seen on slide to cell count.

  - The proportion of RBCs and WBC as well as the # of cells seen are QC checks.

  - If there is any discrepancy, repeat both the cell count and the preparation of the smear.

3.   If performing a dilution, background checks must be performed on the Nerl saline at the time of the counts at least once per day of testing or with a change in the lot number or the saline. Plate the Nerl saline on the hemocytometer and observe for contamination. Document on the Body Fluid Cell Count control form.

4.  DCL Cell Pack:  This is used for making dilutions when specimens are run on the XN analyzer.  A portion of this must be run in the BF mode, prior to analyzing the specimen you are diluting it with.  Observe for contamination.

## Procedure:

SPECIMEN PREPARATION:  All CSF specimens must go to microbiology first for proper routing and paper work.  Microbiology will place the volume and description for each tube submitted on the CSF worksheet.  If this has not been done, check with microbiology before beginning processing to insure proper information on worksheet Clear, colorless spinal fluids must be performed by the hemacytometer counting method.  If the CSF is hazy, cloudy or bloody it can be run on the Sysmex XN-3000 analyzer provided the counts are within the linearity of the instrument. If specimen counts exceed the linearity of the analyzer, appropriate dilutions must be made.  Follow the XN operating procedure when running body fluids on the analyzer. When red blood cell counts are below instrument linearity, they should be counted manually by hemocytometer counting method and those results should be reported instead of the automated count. a  When cell counts are above linearity, the specimen must be diluted and repeated in order to obtain results within the reportable range of the analyzer.  Multiply the instrument results by the dilution factor to obtain the reportable result.

Also, remember to test a portion of your diluting fluid in the BF mode on the XN when indicated, and check the results to make sure there isn’t any contamination.

When counting using the hemacytometer method:

- Cell Counts

- Mix specimen well.

- Observe the clarity of the specimen to aid you in determining dilution to use.

    - If clear, charge both sides of the hemacytometer with undiluted fluid using the microhematocrit tubes or 10ul pipette.

    - If cloudy.   Run on the XN hematology analyzer under the body fluid mode. You may also perform a manual dilution using Nerl saline.

- Charge both sides of the hemacytometer.

- Perform both a WBC and RBC:

    - Count the number of cells, differentiating RBC and WBC, on both sides of the counting chamber (18 sqmm). Document the cell counts on the Body Fluid Worksheet (under number 5, Cell Count).

    - Counts should agree to within 15%; if not recharge the hemacytometer chamber and recount.

    - If count is extremely high, count only the cells in the 4 corner squares on both sides (8 sqmm) or the central square millimeter on both sides (2 sqmm). If specimen was run on the XN analyzer, skip this step.

- Calculate the RBC and WBC using the following formula:

Count =     total # cells X depth (10) X dilution

total #sqmm counted

- Differential

- Prepare 2 cytospin smears. (Definitely prepare two cytospin smears if cytology is ordered)

- Write name and container on slide with pencil If there is room,write the specimen type as well.

- Place filter on slide with frosted side facing the filter.  Be sure filter is square with bottom of slide (opposite end of frosted area).  On back of slide, take felt tip pen and mark around the opening in the filter.

- Place slide and filter in holder with frosted side facing out and at the top.

- Place the cup piece next to filter paper and snap in place (slide-filter-cup).

- Place cup-assembly in cytocentrifuge.

- Be sure to balance with another cup-assembly.

- Place 6 drops of well mixed "clear" fluid or 2-3 drops of "cloudy" fluid into cup.

- Add 1-2 drops of albumin.

- Place cap on cup.

- Close lid on centrifuge and press "MODE" to spin.

  - The speed should be set at 7 and the time at 10.

  - Centrifuge will buzz at completion of spin cycle.

- Remove the cup-assembly from centrifuge.

- Disassemble, throwing cup , holder, and filter in biohazard container.

- Dry slide(s) thoroughly.  To enhance drying, place slide on warm box.

- Stain two slides with Wright's stain by placing in Hematek stainer or SP-100 Sysmex  slide stainer.

- Leave both stained smears for cytology, if ordered and notify the clinical pathologist that a bodyfluid count was performed on the specimen.

- Scan slide for malignant cells under 10X power.

- Perform a leukocyte differential cell count using 50X and/or 100X oil.

  - Include monocytes & macrophages/histiocytes under mononuclears.

  - Solid tumor malignant cells and ependymal (lining) cells should not be included in the differential but footnoted as being present.

  - Do not verify prior to Path review.

  - Correlate # cells seen on slide to cell count.  The proportion of RBCs and WBCs as well as the # of cells seen are QC checks.  If there is any discrepancy, repeat both the cell count and the preparation of the cytospin.

- If cytology is ordered, place two Wright stained smears in small cardboard folder along with patient results for cytology in the path review plastic tray.

- Place remaining specimen in refrigerator unless needed for serology and/or chemistry.

## Calculations

Count = total # cells X depth (10) X dilution

total #sqmm counted

Differential Percentage when 100 cells are not counted:

Diff % =   # Cell Type   Total # of Cells Counted X 100

Reporting Results

1.    Report out in LIS.

| If no WBC and/or RBC seen on first plating | Replate |
| --- | --- |
| If no cells are seen on 2nd plating and RBCs/WBCs are present on the cytospin | Report count as 0-1 WBC/mm3 and/or 01-RBC/mm3 |
| If no cells are seen on either (count & cytospin) | Report as 0 WBC/mm3 And 0 RBCmm3 |
| If abnormal cells (tumor cells, immature cells or cells with intracellular bacteria) are seen | Take to pathology for review prior to verification and to allow for appropriate correlation for other ancillary tests. (ex. Cytology) |
| If specimen was ran on the Sysmex XN analyzer and the RBC count is below linear limit. | You must confirm by plating on the Hemocytometer before turning out the results. |

2.  Manual Hct may be ordered on some specimens.  For those crits that are less     2.0% turn as <2.0

Reference Ranges

| Cell Counts |  |
| --- | --- |
| WBC/mm3 | 0-10 |
| RBC/mm3 | 0 |

| Differential Count |  |  |
| --- | --- | --- |
|  | Adults | Neonates |
| Lymphocytes | 40-80% | 5-35% |
| Monocytes | 15-45% | 50-90% |
| Neutrophils | 0-6% | 0-8% |

Procedural Notes/Problem-Solving Tips

- If physician wants a count on more than one tube, order another CSF-Cells test. (do not add onto same accession)

- Footnote any body fluid that has a clot, clumped cells, or other debris present and indicate that counts may be inaccurate.

- Document cell counts on Body Fluid Worksheet.

- Phone results for all CSF fluids footnoting name of the person phoned to and time phoned.

- CSF WBC >10/cu mm phone as critical value.

- Order Path Review only if requested.

## References

1.	"CLINICAL DIAGNOSIS", Davidsohn & Henry, W.B. Saunders, 18th Edition.

2.	"FUNDAMENTALS OF URINE & BODY FLUID ANALYSIS", Brunzel, W.B. Saunders, 1994.

3.	"DIAGNOSTIC HEMATOLOGY", Rodak, W.B. Saunders, 1995

| POLICY CREATION : |  |
| --- | --- |
| Author: |  |
| Medical Director: Horea Baila, MD | August 31, 2010 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| February 11, 2017 | Elizabeth A. Bauer-Marsh, M.D. |  |
| SECTION MEDICAL DIRECTOR |  |  |
| May 3, 2012 | Wei Liu, M.D., PhD |  |
| July 17, 2014 | Julia Adams, M.D. |  |

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Removed Unopette technique, replaced with methylene blue dilution. Minor formatting changes. Added RBC’s to reporting results section. | T. Mikolajczyk | 8/31/10 |
| 2 | Added reporting of xanthrochromia evaluation | T. Mikolajczyk | 12/7/10 |
| 3 | Deleted step 2 from the procedural notes due to new LIS. Added 1:10 methylene Blue diluent preparation | N. Rutledge | 11/8/11 |
| 4 | Added document cell counts on body fluid worksheet under cell counts 4a. Clarified sentence under procedural notes/problem-solving tips.Also added to section II differentials notify the clinical pathologist that a bodyfluid count was performed on the specimen. Changed slide requirement for cytology. System name and Logo updated. | Kathy Turpin | 12/10/13 |
| 5 | Removed New Methylene Blue diluent for dilution | Kim Paige | 7/5/16 |
| 6 | Changed wording under Sect. 1 2-B and 4-C. Added the wording for background checks if needed. | Kim Paige | 9/5/16 |
| 7 | Added in the QC and procedural notes for running the CSF on the XN analyzer. Added to the reagents and equipment, added in calibration information and other general wording changes. | Kim Paige | 9/12/17 |
| 8 | Added clarification of what to do with RBC counts below analyzer linearity | Kim Paige | 3/9/18 |

Reviewed by

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 8/31/10 |  | 8/31/10 |
|  |  |  | 11/8/11 |  | 11/8/11 |
|  |  |  | 1/17/14 |  | 1/17/14 |
| R. Fitzgerald | 5/8/14 |  | 7/2/14 |  | 7/17/14 |
| K. Paige | 7/5/16 |  | 7/6/16 |  | 7/10/16 |
| K. Paige | 9/5/16 |  | 9/7/16 |  | 9/8/16 |
| K. Paige | 9/12/17 |  | 9/13/17 9/13/17 |  | 9/19/17 |
| K. Paige | 3/9/18 |  | 3/9/18 |  | 3/9/18 |