| METHODIST |  |  | Page 1 of 17 | Section: UPM CHEMC | Policy #: 01.012 |
| --- | --- | --- | --- | --- | --- |
|  | CHEMISTRY, CENTAUR |  | Approved by: see signature block at end of document |  | Date: 3/20/18 Review by: 3/21/18 |
|  | LABORATORY |  | Policy Created: 7/22/02 Supersedes: 1/10/12, 5/15/12, 6/9/14, 4/1/16, 6/9/16, 2/24/17, 3/5/17, 3/20/18 |  |  |
|  |  |  | Primary Responsible Parties: Ray Gross Secondary Responsible Parties: Amy Gibbs |  |  |
|  |  |  | CAP Standard: NA |  |  |
| SUBJECT: |  | PTH INTACT PARATHYROID HORMONE |  |  |  |

LIS ORDER CODE = PTHWCA

I.     Principle

The ADVIA Centaur Intact PTH assay is a two-site sandwich immunoassay using direct chemiluminescent technology, which uses constant amounts of two anti-human PTH antibodies. The first antibody in the Lite Reagent is a monoclonal mouse anti-human PTH (N-terminal) antibody labeled with acridinium ester. The second antibody is a biotinylated l  monoclonal mouse anti-human PTH (C-terminal) antibody that is bound to streptavidin-coated paramagnetic latex particles in the Solid PhaseThere is a direct relationship between the amount of PTH present in the patient sample and the amount of relative light units (RLUs) detected by the system.

II.   Clinical Significance

Parathyroid hormone (PTH), produced by the parathyroid gland is the major circulating factor regulating extracellular calcium concentration.  Parathyroid hormone is secreted by the chief cells of the parathyroid glands as a polypeptide that contains 84 amino acids.  PTH has a molecular weight of 9.4 kDa and is excreted through the kidneys with a half-life of approximately 4 minutes.   Abnormally low ionized calcium concentrations trigger the secretion of PTH.  The PTH molecules bind to type 1 parathyroid hormone receptors in target tissues and initiate a sequence of reactions that result in an increase in extracellular calcium concentrations.  PTH stimulates osteoclastic bone resorption resulting in the release of calcium from bone.  PTH stimulates transcellular calcium reabsorption from the renal tubules and stimulates the kidney to produce 1,25-dihydroxyvitamin D which acts on the intestines to increase calcium reabsorption.  In most clinical conditions, rising levels of extracellular calcium will suppress PTH secretion through a negative feedback mechanism.

Parathyroid hormone increases the rate of bone metabolism.  Depending on the age of the patient, the bones involved, and the concentrations of the hormone in circulation over time, the effect on the bone can be either catabolic or anabolic.  Consistently high concentrations of PTH generally have a catabolic effect and intermittent slightly elevated concentrations have an anabolic effect.

Quantification of circulating intact PTH assists in the differential diagnosis of hypercalcemia and hypocalcemia.  In conjunction with the measurement of ionized calcium, intact PTH evaluations can be used to distinguish between patients with hyperparathyroidism, hypoparathyroidism, or hypercalcemia of malignancy.  The diagnosis of primary hyperparathyroidism, a common cause of hypercalcemia, is confirmed by elevated ionized calcium concentrations and elevated parathyroid hormones.

Intact PTH levels are also used to assess and manage other metabolic bone disorders, including osteoporosis and renal osteodystrophy. Additionally, intra-operative PTH measurements can be used, in conjunction with improved pre-operative localization methods (ultrasound and Sesta-MIBIscan), to control the success of parathyroidectomy for both primary and renal hyperparathyroidism.  The measurement of PTH using two-site immunoassays provides a more accurate assessment of parathyroid tissue secratory status, especially in patients with renal impairment.

The National Academy of Clinical Biochemistry7 recommends the use of intra-operative

parathyroid hormone testing for:

• Patients undergoing initial surgery for primary hyperparathyroidism

• Patients undergoing re-operative hyperparathyroidism

• During pre-operative localization in patients with primary hyperparathyroidism

PTH values can vary depending on the testing procedure used. When monitoring patients over

time, PTH values obtained with different methods should not be used interchangeably.

#### III.   Specimen Collection and Handling

EDTA plasma (lavender top) for intra-operative PTH (only) , Serum (red or gold top) or green top are the recommended specimen types for this assay. Centrifuge green, red or gold top tubes within 1hr after collection and lavender top tubes immediately after collection. For Intra-operative PTH always collect lavender top tubes and perform testing immediately after centifugation. Do not add on any testing once the operation and analysis of those specimens from pre-op to post-op is complete.

The type of specimen used (serum or EDTA plasma) may influence Intact PTH measurements.  During routine monitoring of iPTH levels, to avoid bias in results, use the same specimen type throughout the monitoring period.

Patient Sample Stability is shown in the table: Blood separated from cells

| Temperature | EDTA Plasma | Serum | Lithium or Sodium Heparin |
| --- | --- | --- | --- |
| Room temp | 25 hrs | 8 hrs | 9 hrs |
| 2 – 8 C | 14 days hrs | 8 hrs | 72 hours |
| -20°C | Not recommended | 1 month | Not recommended |

Thawed frozen specimens must be clarified by centrifugation prior to testing.

For Reference Lab and other outside clients aliquots of refrigerated or frozen samples must meet the patient sample stability table shown above, otherwise the sample will be rejected and the test request canceled.  Intact PTH has been demonstrated to be labile and is susceptible to fragmentation.

In case the Centaur is out of service , the Immulite 1000 rapid PTH is the back up method to use for intraoperative PTH samples.   If any of the intraoperative samples were performed on the Centaur before taken out of service, repeat those samples on the back up method, Immulite.   All of the patient’s intraoperative samples must be performed on the same analyzer.

When used intraoperatively to verify a decline in PTH production after resection, it is recommended that two or more samples be collected at least 5–10 minutes apart.

The sample volume pipetted for the PTH-INT assay is 200 ul, excluding the unusable volume in the sample container or the additional volume required when performing duplicates or other tests on the same sample.  Ensure samples are free of fibrin, particulate matter, and/or bubbles prior to placing on system.

IV.  Reagents

- ADVIA Centaur Intact PTH primary reagent pack

- Lite Reagent – Unopened:Stable until expiration date at 2-8°C.  Onboard, 28 days   .

10   ml acridinium ester-labeled mouse monoclonal  antihuman PTH  antibody (~ 0.6   μg/mL) in  buffered saline with mouse gamma globulin , bovine serum  , bovine serum albumin, and preservatives.

- Solid Phase- Unopened Stable until expiration date at 2-8°C.  Onboard, 28 days

20.0 ml biotinylated mouse monoclonal  anti-human PTH antibody bound to streptavidin- coated paramagnetic particles(~0.4 mg/mL) in  buffered saline  with  bovine gamma globulin, bovine serum albumin, and preservatives.

- Multi-Diluent 13  - Unopened Stable until expiration date at 2-8°C or 28 consecutive days after accessing the ancillary reagent pack.

10  mL/reagent pack.  Buffer with surfactant and sodium azide (<0.1%)  .

- PTH Calibrator – Unopened; Stable  until expiration date at 2-8 C

Reconstituted : 8 hours at 18-25 C.  Reconstituted : 60 days at </= 20C.

After reconstitution, low or high levels of intact PTH synthetic peptide in 		buffered saline with human EDTA plasma (10%), surfactants and          preservatives.

- Wash 1500 mL/pack and 2500 mL/pack, Phosphate-buffered saline with sodium azide (<0.1%) and surfactant.   Unopened: Stable until the expiration date on the product at 2-25 C.  On-system: 1 month

Storing and Stability

Store unopened reagents upright at 2–8°C.

Protect unopened reagent packs from all heat and light sources. Reagent packs loaded on the

system are protected from light.

All reagents are stable at 2–8°C until the expiration date on the product.

A.   Precautions

- This device contains material of animal origin and should be handled as a potential carrier and transmitter of disease.

- Do not use kit components beyond expiration date.

- Contains human source material.   Each donation of human blood or blood component was tested by FDA-approved methods for the presence of antibodies to human immunodeficiency virus typy 1 (HIV-!) and type 2 (HIV-2) as well as for hepatitis B surface antigen (HBsAg) and antibody to hepatitis C virus (HCV).   The test results were negative (not repeatedly reactive).  No test offers complete assurance that these or other infectious agents are absent; this material should be handled using good laboratory practices and universal  precautions.

- Contain sodium azide as a preservative.

- Contains 2 –methyl-2H-isothiazol-3-one.  May produce an allergic reaction. ADVIA Centaur PTH Calibrator

- Safety data sheets (MSDS/SDS) available on www.siemens.com/diagnostics.

B.   Loading Reagents

- Ensure that the system has sufficient primary and ancillary reagent packs.

- CAUTION: Mix all primary reagent packs by hand before loading them onto the system. Visually inspect the bottom of the reagent pack to ensure that all particles are dispersed and re-suspended.

- Load the Ready Pack reagent packs in the primary reagent area using the arrows as a placement guide. The system automatically mixes the primary reagent packs to maintain homogenous suspension of the reagents.

C.   Onboard stability

- The ADVIA Centaur intact PTH assay has an onboard stability of 28 days.

- Discard the primary reagent packs at the end of the onboard stability interval

- Do not use reagents beyond the expiration date.

# V.    Instrumentation/Equipment

The ADVIA CENTAUR system is an automated, immunoassay analyzer that offers optimal productivity and efficiency. No-pause reloading of reagents, samples, and supplies means that the system is always ready to process samples. All assays use direct chemiluminescent technology. Chemiluminescence is a chemical reaction that emits energy in the form of light. When used in combination with immunoassay technology, the light produced by the reaction indicates the amount of analyte in the sample. Direct chemiluminescent reactions directly measure the light energy without the use of added steps or amplifying molecules. The ADVIA Centaur assays use acridinium ester as the chemiluminescent label, since it does not require the addition of a catalyst or substrate.

When the sample start key is pressed, the barcode labels on the sample cups are read, sample is aspirated, reagent is dispensed, and the assay process begins.  Particles are magnetically separated in the cuvette incubation ring. The addition of hydroxyl groups to complete the flash reaction is accomplished by the addition of Reagent 1 & 2; Acid and Base. The chemiluminescent reaction occurs in the luminometer.  The photomultiplier tube measures the chemical light reaction that takes place.

There is one (1) main system operation key on the ACS:CENTAUR, the “Sample Start button”.  Pressing this key performs the following actions:

- Homes the subsystems.

- The system starts specimen sampling.

- If the start button is pressed while the instrument is running, it stops sampling additional specimens, however it continues to process the specimens in the incubation ring.

## Additional Equipment and Supplies

Reagent Water

Sample cups / tubes

Cuvettes

Sample tips

Reagent 1  (0.5% H2O2; 0.1N HNO3)

Reagent 2  (less than 0.25N. Sodium Hydroxide and surfactant)

ACS:CENTAUR Cleaning Solution

ACS:CENTAUR primary and ancillary reagents.

VI.  Calibration

The ADVIA Centaur system uses a Master Curve and a two-point operator initiated calibration to calibrate assays. The Master Curve is established as part of the manufacturing process for each assay lot number.

The ADVIA Centaur intact PTH assay requires a Master Curve calibration when using a new lot number of Lite Reagent and Solid Phase. For each new lot number, use the barcode reader to enter the Master Curve values on the system. The Master Curve Card contains the Master Curve values. Refer to the system operating instructions for more information.

A two-point calibration must be performed at regular, assay specific intervals. Replicates for two calibrators of known value are processed. If the calibrators meet defined validity criteria, the system is adjusted. Refer to the Centaur Operating Procedures for calibration procedure.

The low and high calibrators provided in this kit are matched to the ReadyPack primary reagent pack.  Do not mix calibrator lots with different lots of reagent packs.

A.   Calibration Frequency

Two-point calibration of the intact PTH assay is required:

- every 14 days

- when changing lot numbers of primary reagent packs

- when replacing system components

- when QC results are unacceptable

B.   Calibration Material

The intact PTH assay is calibrated with PTH Calibrators, Low and High.

- Reconstitute each vial with 1.0 ml reagent grade water using a volumetric pipet.

- Let the calibrator stand for 30 minutes at room temperature to allow the lyophilized material to dissolve.

- Gently swirl and invert the vials until homogenous.  Avoid bubbles.

- Each lot of calibrators contains a Calibrator Assigned Value card. Calibrator values may be entered by using the barcode wand or the keyboard.

- Affix the Low and High Calibrator barcode labels to the appropriate calibrator sample cups so that the system recognizes the sample as a calibrator.  Calibrator barcode labels are lot-number-specific.  Do not use barcode labels from one of calibrators with any other lot of calibrators.

- Refer to the Centaur Operating Procedures for scheduling and running a calibration.

C.   Stability / Storage

- Lyophilized calibrator is stored at -20C until expiration date on the vial

- Reconstituted calibrator is stable at room temp. or onboard for 8 hrs.

For longer storage, store at _- </=20 C for up to 60 days.   Freeze and thaw the calibrators one time only.

D.   Precautions

- Sodium Azide can react with copper and lead plumbing to form explosive metal azides. On disposal, flush reagents with a large volume of water to prevent the buildup of azides.

- Potential biohazard, human and/or other biological source material. Handle as if potentially infectious, according to established good laboratory practices.

- Do not return any calibrators back into the vials after calibration because evaporation can occur, which may affect performance.

- Dispose of any calibrator remaining in the sample cups after 4 hours.

- Do not refill calibrator sample cups when the contents are depleted. If required, dispense fresh calibrators into new cups.

#### VII.   Quality Control

A.   QC product:  BioRad Liquichek Specialty Immunoassay Control, Levels 1, 2, & 3.  BioRad ECS Division, Anaheim, CA/

B.   Preparation/handling

- Stored frozen.  This QC product is stable until the expiration date on the label when stored unopened at –20 to -70C.

- Allow the frozen control to stand at room temperature until it is completely thawed. Before sampling, gently swirl the vial several times to ensure homogeneity. Promptly replace the stopper and refrigerate at 2 to 8C

- Once thawed and opened, all analytes are stable for thirty days when stored tightly capped at 2 to 8C, except for PTH-INT which is stable for 23 days.

- Do not use past the expiration date.

- Discard the vial if there is evidence of microbial contamination or excessive turbidity in the product.

- Frequency

All three levels of quality control are run on each day on first shift.

- Acceptability Criteria

Acceptability of QC is determined by the lab internal QC policy.  Corrective action for out-of control QC is outlined in the QC policy and actions taken must be documented in the LIS system.  No patient results may be released until QC results are acceptable.

VIII.   Procedure

Performed daily. Frozen patient aliquots will be thawed by Centaur operator. This assay requires 50 μL of sample for a single determination. This volume does not include

the unusable volume in the sample container or the additional volume required when

performing duplicates or other tests on the same sample. For detailed information about

determining the minimum required volume, refer to the system operating instructions.

Before placing samples on the system, ensure that samples have the following characteristics:

• Samples are free of fibrin or other particulate matter and free of bubbles.

- Prepare the sample container for each sample, ensuring that a barcode label is affixed. Avoid bubbles.

- Samples can be loaded on the Aptio in the appropriate lane or if front loaded, use the appropriately coded sample racks for the type of sample tube to be used:

- Position 1 – aliquot tube (blue screw cap)

- Position 2 – primary sample tube

- Position 3 – sample cup (Siemens)

- Load each sample tube into a rack, ensuring that the barcode is visible through the slot in the rack.

- Place the rack(s) in the entry queue.

- Press ’START’ only if the system is not currently ‘In Process’. The analyzer will read the barcode label and run the appropriate tests.

- For the PTH assay, the system automatically performs the following steps:

- dispenses 50   ul of sample into a cuvette.

- dispenses 100  ul of Lite Reagent

- dispenses 200 ul of Solid Phase and incubates 8.25  min. at 37C.

- separates, aspirates, and washes the cuvettes with reagent water.

- dispenses 300 ul each of Acid Reagent and Base Reagent to initiate the chemiluminescent reaction.

- reports results.

### IX.   Reporting Results

A.    Calculations

The system automatically reports serum PTH-INT results in pg/ml. For detailed information about how the system calculates results, refer to the system operating instructions.

B.    Limitations of Procedure

- Assay range (AMR): 	 6.3 – 2000  pg/ml

(manufacturer’s stated 2.5-1900.0 pg/mL)

Results less than 6.3 pg/ml, report as “< 6.3 pg/ml”

- Automatic dilution:	The system is programmed to perform an automatic

repeat with a x5 dilution for patient samples reading

>2000  pg/ml. If a repeat test is required, the system

automatically schedules and performs the repeat test.  No operator intervention is required, as the sample racks are held in the in process queue until results are completed.

3.    Manual dilution: 	Not normally needed.  A smaller dilution, if needed, may be made by using Multi-Dil   13

# C.   Reference Interval

18- 80   pg/mL

D.   Reporting

- Analytical measurement range (AMR):  6.3 – 2000  pg/ml

- Clinical reportable range (CRR):  6.3 –  10,000    pg/ml (using x5 autodilute)

Report results less than 6.3 as “< 6.3 pg/ml”.

X.   Procedural Notes/Problem-Solving Tips

- It is extremely important to ensure that patient samples have been handled and stored correctly.  Incorrect handling of samples will result in a loss of intact PTH.

- Expedite processing and analysis of intraoperative PTH samples. The calcium component should be entered as ‘NA’, and results must be phoned ASAP. There will usually be a preoperative specimen as well as a postoperative specimen.  . Add-ons are not acceptable for intra-operative specimens. Once the operation is over and analysis is complete with the preoperative and postoperative specimens, any new orders will need to be drawn using serum specimens.

- Interpretation of intact PTH values should always take into account serum calcium results and the interrelationship between these two elements in various disorders involving PTH and calcium.  It is recommended that the intact PTH results should always be interpreted with caution and with consideration of the overall clinical manifestations even when used in conjunction with calcium values.

- Some overlap of intact PTH values does exist for patients with various parathyroid disorders.  Measurement of intact PTH is useful in differentiating between hypercalcemia due to hyperparathyroidism and hypercalcemia of malignancy.  However, the assay is not intended as, and should not be relied upon as, a diagnostic indicator of malignancy.

4.The ADVIA Centaur PTH assay will detect non-intact PTH, such as PTH fragment (7-84). PTH

fragment (7-84) may cause falsely elevated PTH results in patients with abnormal renal

function because these patients may have various concentrations of PTH fragment (7-84)

in their blood. In patients with abnormal renal function, interpret the PTH result with

caution, and do not make patient management decisions on the PTH result alone. A study

of characterized PTH fragments is provided in Lopez et al “Selected Reaction Monitoring-

Mass Spectrometric Immunoassay Responsive to Parathyroid Hormone and Related

Variants.”

- Patient samples with high intact PTH levels can cause a paradoxical decrease in RLUs (high dose hook effect). In this assay, patient samples with PTH levels as high as 100,000 pg/ml do not demonstrate a high dose hook effect.

- Lipemic (<3000 mg/dL of triglycerides), icteric (<40 mg/dL of bilirubin), and hemolyzed  (<200 mg/dl of hemoglobin) demonstrate <10% change in results.

- Patient samples may contain  heterophilic antibodies that could react in immunoassays to give falsely elevated or depressed results.  This assay is designed to minimize the interference from heterophilic antibodies.  t .

- Performance Characteristics

The ADVIA centaur PTH assay measures intact parathyroid hormone concentrations from  6.3-2000 pg/ml(0.668 – 212 pmol/L).  The low end of the measuring interval is constrained by the low end of the measuring interval of the predicate assay.

Report results below the measuring interval as <6.3 pg/mL (<0.668 pmol/L).  When samples exceed the measuring interval, refer to Dilutions.

Specificity

Interference was determined according to CLSI Document EP07-A12.

The cross-reactivity of the ADVIA Centaur PTH assay was determined by spiking samples with the PTH fragments and compounds listed below at the indicated levels.  The following results were obtained:

| Cross-Reactant | Amount Added (pg/mL) | Cross-Reactivy (%) |
| --- | --- | --- |
| PTH (1-34) fragment | 12000 | <0.1 |
| PTH (39-68) fragment | 100000 | <0.1 |
| PTH (39-84) fragment | 100000 | <0.1 |
| PTH (44-68) fragment | 100000 | <0.1 |
| PTH (53-84) fragment | 100000 | <0.1 |
| PTH (7-84) fragment | 300 | 37.4 |
| Beta-Cross Laps | 10000 | <0.1 |
| Calcitonin | 100000 | <0.1 |
| Osteocalcin | 50000 | <0.1 |
| PTH RP (1-34) | 100000 | <0.1 |

Detection Capability The Limit of Blank (LoB), the Limit of Detection (LoD), and the Limit of Quantitation (LoQ) were determined as described in CLSI protocol EP17-A2 using the ADVIA Centaur XP system.  The assay is designed to have an LoB of <6.0 pg/mL, and LoD of </=6.0 pg/mL, and an LoQ of </= 6 pg/mL.

Representative detection capability data are shown below.  Assay results obtained at individual laboratories may vary from the data presented.

The LoB is defined as the highest measurement result that is likely to be observe for a blank sample.   The ADVIA Centaur PTH assay has an LoB of 1.5 pg/mL (0.160 pmol/L)

The LoD is defined as the lowest concentration of intact PTH that can be detected with 95% probability.  The ADVIA Centaur PTH assay has an LoD of 3.2 pg/mL (0.339 pmol/L)

The LoQ is defined as the lowest concentration of intact PTH that can be detected at a total CV of 20%.  The ADVIA Centaur PTH assay has an LoQ of 4.6 pg/mL (0.488 pmol/L).

Precision

The ADVIA Centaur PTH assay is designed to meet the following repeatability and within-lab precision requirements:

Intact PTH Level						Design Requirements

| (pg/mL) | (pmol/L) | Repeatability (within –run) %CV | Within-Lab (total precision)_ %CV |
| --- | --- | --- | --- |
| 10.0-20.0 | 1.06-2.12 | </=8.0 | </=10. |
| >20.0-700 | >2.12-74.2 | </=6.0 | </=8.0 |
| >700 | >74.2 | </=8.0 | </=10.0 |

Precision was evaluated according to the CLSI protocol EP05-A3.19 Six EDTA plasma samples

were assayed 2 times per day in replicates of 2, for 20 days (n = 80 replicates per sample) using

the ADVIA Centaur PTH assay. The following results were obtained:

Mean			Repeatability				Within-Lab

(Within-Run)				(Total Precision)

| Sample | pg/mL | Pmol/L | Pg/mL | Pmol/L | %CV | SD pg/mL | SD pmol/L | %CV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample 1 | 16.9 | 1.79 | 0.9 | 0.09 | 5.16 | 1.1 | 0.12 | 6.78 |
| Sample 2 | 50.3 | 5.34 | 1.4 | 0.15 | 2.78 | 2.3 | 0.24 | 4.57 |
| Sample 3 | 146.4 | 15.52 | 1.9 | 0.20 | 1.30 | 2.9 | 0.30 | 1.95 |
| Sample 4 | 409.2 | 43.37 | 6.8 | 0.72 | 1.67 | 16.9 | 1.79 | 4.13 |
| Sample 5 | 596.3 | 63.20 | 7.3 | 0.78 | 1.23 | 9.7 | 1.02 | 1.62 |
| Sample 6 | 1080.4 | 114.52 | 11.8 | 1.25 | 1.09 | 25.1 | 2.66 | 2.32 |

ADVIA Centaur PTH Quality Controls

| Sample | pg/mL |  |  | Pmol/L |  | Pg/mL |  | Pmol/L |  | %CV |  | SD pg/mL |  | SD pmol/L |  | %CV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Control 1 |  | 41.6 | 4.41 |  | 0.4 |  | 0.04 |  | 0.91 |  | 1.4 |  | 0.15 |  | 3.36 |  |
| Control 2 |  | 235.9 | 25.00 |  | 1.9 |  | 0.20 |  | 0.81 |  | 4.4 |  | 0.46 |  | 1.85 |  |
| Control 3 |  | 874.0 | 92.64 |  | 8.0 |  | 0.85 |  | 0.92 |  | 18.4 |  | 1.96 |  | 2.11 |  |

Actual results obtained at individual laboratories may vary from the data provided.

Accuracy/Method Comparison

The ADVIA Centaur PTH assay is designed to have a correlation coefficient ≥ 0.95 when

compared to a commercial intact PTH assay. Typical data obtained showed 0.99.

For 349 EDTA plasma samples within the range of the commercial intact PTH assay

(6.3–1900 pg/mL (0.67–201 pmol/L)), the relationship between the ADVIA Centaur PTH

assay (y) and the commercial intact PTH assay (x) is described using Passing-Bablok regression:

ADVIA Centaur PTH (y) = 1.02 (x) - 2.18 pg/mL (0.231 pmol/L), r = 0.99

The correlation of the assay may vary depending on the study design, comparable method, and

sample population. Results obtained at individual laboratories may vary from the data

presented.

Specimen Collection Tube ComparisonThe ADVIA Centaur PTH assay was evaluated using different specimen matrices and tube collection types. The assay is designed to have a correlation of coefficient ≥ 0.95, and a slope of test tube type vs. reference of 1.0 ±0.1.

A total of 56 matched dipotassium EDTA plasma, serum, serum separator tube, lithium

heparin, and sodium heparin samples that span the range of the assay were examined. The following results were obtained using the ADVIA Centaur XP system:

| Comparison | Regression Equation | r |
| --- | --- | --- |
| Serum vs. dipotassium EDTA plasma | Serum = 0.99 EDTA-1.85 | 0.996 |
| Serum separator tube vs. dipotassium EDTA plasma | Serum separator tube= 1.03 EDTA +0.20 | 0.996 |
| Lithium heparin vs. dipotassium EDTA plasma | Lithium heparin=0.99 EDTA +1.95 | 0.998 |
| Sodium heparin vs. dipotassium EDTA plasma | Sodium heparin = 1.00 EDTA+1.03 | 0.997 |

This study was performed using Becton Dickinson tubes.   Siemens recommends that laboratories evaluate performance when using other manufacturers’ tubes.

Interferences

Potential interference in this assay from the interferents listed below is designed to be ≤ 10%.  Interfering substances at the levels indicated in the table below were tested as described in CLSI Document EP07-A217 using the ADVIA Centaur PTH assay.

| Interferent | Highest Concentration Tested | Observed Interference (%) |
| --- | --- | --- |
| Aliskiren | 200 ug/mL | 5.3 |
| Bilirubin (conjugated) | 60 mg/dL | 5.9 |
| Bilirubin (unconjugated) | 60 mg/dL | -2.2 |
| Biotin | 1000 ng/mL | 6.9 |
| Caffeine | 308 umol/L | 2.3 |
| Calcitrol | 360 pg/mL | 2.5 |
| Cholesterol | 500 mg/dL | -6.5 |
| Enalaprilat | 0.86 umil/L | -2.6 |
| Epoetin alfa | 15 mU/L | -1.2 |
| Fosrenol | 20 ng/mL | 3.4 |
| Furosemide | 181 umol/L | -2.6 |
| Hemoglobin | 500 mg/dL | -4.5 |
| IgG | 6 g/dL | 1.3 |
| Total protein (high) | 12 g/dL | -4.2 |
| Total protein (low) | 6g/dL | -7.5 |
| Triglycerides | 3275 mg/dL | -0.4 |

Six human plasma samples containing human anti-mouse antibodies (HAMA) were tested at 2 PTH concentrations. Samples at a PTH concentration of 60 pg/mL (6.36 pmol/L)

demonstrated a difference of -11.5% to 3.2% from the expected values. Samples at a PTH concentration of 340 pg/mL (36.04 pmol/L) demonstrated a difference of -3.8% to 2.4% from the expected values. Refer to Limitations.

Assay results obtained at individual laboratories may vary from the data presented.

Linearity

The ADVIA Centaur PTH assay is linear from 6.3–2000 pg/mL (0.668–212 pmol/L).

Linearity was evaluated according to the CLSI Document EP06-A.20 An EDTA plasma sample containing a high concentration of intact PTH was mixed with an assay-specific diluent (ADVIA Centaur Multi-Diluent 13, REF 10492364). The resulting sample mixtures were tested with the ADVIA Centaur PTH assay

Dilution Recovery

This assay is designed to have a mean dilution recovery of 90%–110%.

A low sample pool (~22 pg/mL PTH) was spiked with a commercially available synthetic

PTH peptide to reach levels of approximately 3000 pg/mL (318 pmol/L), 6000 pg/mL

(636 pmol/L), and 9000 pg/mL (954 pmol/L). Each sample was diluted 1:5 with ADVIA Centaur Multi-Diluent 13 onboard the ADVIA Centaur XP system. The following results were obtained:

| Sample | Observed Mean (pg/mL) | Expected (pg/mL) | Observed Mean (pmol/L) | Expected (pmol/L) | Recovery (%) |
| --- | --- | --- | --- | --- | --- |
| 1 | 3068.8 | 3020.9 | 325.3 | 320.2 | 101.6 |
| 2 | 6270.8 | 6019.8 | 664.7 | 638.1 | 104.2 |
| 3 | 9253.2 | 9018.7 | 980.8 | 956.0 | 102.6 |
| Mean |  |  |  |  | 102.8 |

Each value in this column is the mean of 3 onboard dilution results.

Spiking Recovery

This assay is designed to have a mean spiking recovery of 90%–110%.

Varying amounts of intact PTH were added to 4 human EDTA plasma samples with endogenous intact PTH levels of 26.9–74.4 pg/mL (2.9–7.9 pmol/L). The recoveries ranged from 96.9%–101.8% with a mean of 99.7%.

| Sample | Observed (pg/mL) | Expected (pg/mL) | Observed (pmol/L) | Expected (pmol/L) | Recovery (%) |
| --- | --- | --- | --- | --- | --- |
| 1 | 72.2 | 73.1 | 7.65 | 7.75 | 98.8 |
|  | 727.4 | 730.9 | 77.10 | 77.48 | 99.5 |
| Mean |  |  |  |  | 99.2 |
| 2 | 119.4 | 119.7 | 12.66 | 12.69 | 99.7 |
|  | 754.9 | 779.3 | 80.02 | 82.61 | 96.9 |
| Mean |  |  |  |  | 98.3 |
| 3 | 80.1 | 78.7 | 8.49 | 8.34 | 101.8 |
|  | 735.7 | 738.0 | 77.98 | 78.23 | 99.7 |
| Mean |  |  |  |  | 100.7 |
| 4 | 116.6 | 115.1 | 12.36 | 12.20 | 101.3 |
|  | 769.2 | 772.1 | 81.54 | 81.84 | 99.6 |
| Mean |  |  |  |  | 100.5 |
| Mean |  |  |  |  | 99.7 |

High-Dose Hook Effect

Patient samples with high intact PTH levels can cause a paradoxical decrease in the RLUs (high-dose hook effect). In this assay, patient samples with intact PTH levels as high as 100,000 pg/mL (10,600 pmol/L) are reported as > 2000 pg/mL (> 212 pmol/L).

Standardization

The ADVIA Centaur PTH assay standardization is maintained with internal standards using purified human PTH (1-84). Assigned values for calibrators are traceable to this

standardization.

- . References

1. Mundy GR, Guise TA. Hormonal control of calcium homeostasis. Clin Chem.

1999;45:1347–1352.

2. Marx SJ. Hyperparathyroid and hypoparathyroid disorders. N Engl J Med.

2000;343:1863–1875.

3. Friedman PA. Mechanisms of renal calcium transport. Exp Nephrol. 2000;8(6):343–350.

4. Akizawa T, Fukagawa M. Modulation of parathyroid cell function by calcium ion in health and uremia. Am J Med Sci. 1999;317(6):358–362.

5. Itani O, Tsang RC. Bone disease. In: Kaplan LA, Pesce AJ, eds. Clinical Chemistry: Theory, Analysis, and Correlation. 5th ed. St. Louis, MO: Mosby; 2009:614–650.

6. Risteli J, Winter WE, Kleerekoper M, Risteli L. Disorders of bone and mineral metabolism. In: Burtis CA, Bruns DE, eds. Tietz Fundamentals of Clinical Chemistry and Molecular Diagnostics. 7th ed. St. Louis, MO: Elsevier; 2015: 741–768.

7. Nichols JH ed. Evidence-based practice for point-of-care testing. In: National Academy of Clinical Biochemistry Presents Laboratory Medicine Practice Guidelines. AACC Press,

2006;105–119.

8. Centers for Disease Control. Update: Universal precautions for prevention of transmission of human immunodeficiency virus, hepatitis B virus and other bloodborne pathogens in healthcare settings. MMWR. 1988:37:377–382, 387,388.

9. Clinical and Laboratory Standards Institute. Protection of Laboratory Workers From

Occupationally Acquired Infections; Approved Guideline—Fourth Edition. CLSI Document

M29-A4. Wayne, PA: Clinical and Laboratory Standards Institute; 2014.

10. Federal Occupational Safety and Health Administration. Bloodborne Pathogens Standard, 29 CFR 1910.1030.

11. Clinical and Laboratory Standards Institute. Procedures for the Collection of Diagnostic Blood Specimens by Venipuncture; Approved Standard—Sixth Edition. CLSI Document GP41-A6. Wayne, PA: Clinical and Laboratory Standards Institute; 2007.

12. Clinical and Laboratory Standards Institute. Tubes and Additives for Venous and Capillary Blood Specimen Collection; Approved Standard—Sixth Edition. CLSI document GP39-A6. Wayne, PA: Clinical and Laboratory Standards Institute; 2010.

13. Clinical and Laboratory Standards Institute. Procedures for the Handling and Processing of Blood Specimens for Common Laboratory Tests; Approved Guideline—Fourth Edition. CLSI Document GP44-A4. Wayne, PA: Clinical and Laboratory Standards Institute; 2010.

14. Lopez MF, Rezai T, et al. Selected reaction monitoring-mass spectrometric immunoassay responsive to parathyroid hormone and related variants. Clin Chem. 2010;56:281–290.

15. Boscato LM, Stuart MC. Heterophilic antibodies: a problem for all immunoassays. Clin

Chem. 1988;34:27–33.

16. Clinical and Laboratory Standards Institute. Defining, Establishing, and Verifying Reference Intervals in the Clinical Laboratory. CLSI Document EP28-A3c. Wayne, PA: Clinical and Laboratory Standards Institute; 2010.

17. Clinical and Laboratory Standards Institute. Interference Testing in Clinical Chemistry;

Approved Guideline—Second Edition. CLSI Document EP07-A2. Wayne, PA: Clinical and

Laboratory Standards Institute; 2005.

18. Clinical and Laboratory Standards Institute. Evaluation of Detection Capability for Clinical Laboratory Measurement Procedures; Approved Guideline—Second Edition. CLSI

Document EP17-A2. Wayne, PA: Clinical and Laboratory Standards Institute; 2012.

19. Clinical and Laboratory Standards Institute. Evaluation of Precision Performance of

Quantitative Measurement Methods; Approved Guideline—Third Edition. CLSI Document

EP05-A3.Wayne, PA: Clinical and Laboratory Standards Institute; 2014.

20. Clinical and Laboratory Standards Institute. Evaluation of the Linearity of Quantitative

Measurement Procedures: A Statistical Approach; Approved Guideline. CLSI Document

EP06-A. Wayne, PA: Clinical and Laboratory Standards Institute; 2003.

21.Intact PTH (iPTH) Instructions for Use, Siemens Centaur and Centaur XP, Document ID 10492444_EN, Rev. B, 5/2011.  Siemens Healthcare Diagnostics.

22.  Intact PTH Calibrator Instructions for Use, Document 10492392 Rev.A, 08/2010. Siemens Healthcare Diagnostics.

23.  Intact PTH(iPTH) Instructions for use:  Siemens Centaur and Centaur XP, Document 10699545_EN Rev D, 2014-08.

| POLICY CREATION : |  |
| --- | --- |
| Author: Brenda Mahoney, CLS/ Judy Wetzel, CLS | July 22, 2002 |
| Medical Director: Donald L. Frederick, PhD | July 22, 2002 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| March 4, 2017 | Elizabeth A. Bauer-Marsh, M.D. |  |
| SECTION MEDICAL DIRECTOR |  |  |
| September 20, 2012 | Robert Benirschke, PhD |  |
| April 21, 2015 | Devendra Trivedi, MD |  |
| February 29, 2016 | Lori Racsa, DO |  |

| REVISION HISTORY (began tracking 2011) |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Added revision history chart, revised for new LIS system, changes in QC | R. Pavlacic | 9/9/11 |
| 2 | Revised due to new biotin free assay method. | T. King | 12/7/2011 |
| 3 | Title/header formatting changes, added manufacturer sens. and assay range, smaller dilution protocol addition. | R Pavlacic | 2/1/12 |
| 4 | Updated Centaur to Centaur XP, removed downtime notification associated with two centaurs | D Roth/ S Burton | 6/03/14 |
| 5 | Updated quality control prep and handling | D Roth/ S Burton | 3/30/16 |
| 6 | Updated LIS code due to Sunquest install. | M. Greer | 6/1/16 |
| 7 | Updated from XP to XPT, Added sample can be loaded on Aptio lane | A.Gibbs | 2/13/17 |
| 8 | Added back up method and all of the patient’s intraoperative samples must be performed on the same analyzer. | A.Gibbs | 05/26/17 |
| 9 | Updated SOP with PTH reagent which replaced iPTH reagent per Siemens | A.Gibbs | 11/27/17 |
| 10 | Added that add-on is not acceptable for interoperative PTH testing. | Dr. Racsa | 03/16/18 |

REVIEWED BY

| Lead | Date | Coordinator/Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 1/10/12 |  | 1/10/12 |
|  |  |  | 5/8/12 |  | 5/15/12 |
|  |  |  | 6/3/14 |  | 6/9/14 |
| Donna Roth | 3/30/16 |  | 3/30/16 |  | 4/1/16 |
| M. Greer | 6/1/16 |  | 6/6/16 |  | 6/9/16 |
| R. Gross | 2/16/17 |  | 2/13/17 |  | 2/24/17 3/5/17 |
| R. Gross | 5/30/17 |  | 5/26/17 |  | 6/7/17 |
| R. Gross | 3/19/18 |  | 3/16/18 |  | 3/20/18 |