| METHODIST |  |  | Page 1 of 3 | Section: UPM MICRO | Policy #: 07.009 |
| --- | --- | --- | --- | --- | --- |
|  | MICROBIOLOGY |  | Approved by: see signature block at end of document |  | Date: 3/15/18 Review by: 3/15/20 |
|  | LABORATORY |  | Policy Created: Supersedes: 10/3/17, 3/13/18 |  |  |
|  |  |  | Primary Responsible Parties: Marsha Bishoff Secondary Responsible Parties: Audrey Davis |  |  |
|  |  |  | CAP Standard: COM 04250 |  |  |
| SUBJECT: |  | INSTRUMENT COMPARISON STUDIES IN MICROBIOLOGY |  |  |  |

- POLICY STATEMENT:

Instrument Comparison Studies:  Quantitative test which can be analyzed on more than one Instrument or method will be checked for result correlation twice a year.

- PURPOSE:

To standardize performance of studies which verify instrument agreement on patient results for analytes performed on multiple instruments and/or methods, declare performance frequency, and define acceptability criteria.

- POLICY SCOPE

The scope of this policy applies to all Laboratory staff that prepares or performs testing on laboratory specimens at UnityPoint Methodist.

- GENERAL INFORMATION:

A minimum of five samples should be selected for each analyte that is currently performed on more than one instrument. Specimens selected should be patient samples/isolates. QC is acceptable when patient samples are not usable due to QNS, especially with BD Affirm. Comparison studies should be performed twice a year, January and July.

- PROCEDURE

- C. difficile

- Test a total of five patient stool specimens for C. difficile (three positive specimens and two negative specimens) on the Biofire FilmArray and the Cepheid GeneXpert.

- Acceptability is 100% congruent results with GeneXpert results compared to Biofire.   Some negative specimens on the GeneXpert may be positive on the Biofire due to toxin A being detected on Biofire but not GeneXpert, this would still be considered an acceptable study.

- MicroScan

  - For the Microscan comparison the following panels should be tested on the Walkaway 96 and the Walkaway 40:  Neg Combo 62, Pos Combo 44, MICroSTREP +1, RPID2, and RNID3.  Use a mix of three Gram-negative organisms and three Gram-positive organisms (can be anything that we have the day we run the comparison studies – E. coli, S. aureus, Enterococcus, etc.)  In addition to the panels being compared on both Walkaway instruments, the rapid ID panels need to be compared to the corresponding combo panels. Use the same three Gram-negative organisms and Gram-positive organisms for the panel comparisons.

    - Susceptibility: Acceptability would be 3% major errors and 7% minor errors for susceptibility interpretations. Overall agreement of 90% is acceptable.

    - Identification: Acceptability would be 100% agreement with the organisms’ genus and 83% genus and species.

- BD Affirm

  - Due to the small amount of patient specimen, three positive controls and two negative controls within the same lot (both controls and reagents) will be run on all instruments.

      - Acceptability will be 100% agreement between the instruments.

- MAINTENANCE AND STORAGE

  - All policies and procedures are reviewed every two years by Laboratory Administration and or the Medical Director of the Laboratory or designee.

  - The Laboratory Administration and Medical Director review policies and procedures when there are changes in practice standards, or requirements.

  - All policies and procedures are reviewed every two years by staff or at the time new or revised ones are put in effect.

  - All policies are retained 8 years after being discontinued or revised.

  - All procedures are retained 2 years after being discontinued or revised.

UnityPoint Health Methodist Laboratory is a CAP accredited facility, as of 7/1/11 the responsibility of new and/or substantially revised policies and procedures will be restricted the Laboratory Director whose name appears on the CLIA certificate, whose signature appears below. The biennial review will be completed by the Administrative Director.

| POLICY CREATION : |  |
| --- | --- |
| Author: Teresa Nuese | September 7, 2017 |
| Medical Director: Dr. Elizabeth Bauer Marsh | October 2, 2017 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| October 2, 2017 | Elizabeth A. Bauer-Marsh, M.D. |  |
| SECTION MEDICAL DIRECTOR |  |  |
| October 2, 2017 | Lori Racsa, DO |  |

| REVISION HISTORY (began tracking 2011) |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Initial Release | T. Nuese | 10/3/17 |
| 2 | Updated requirements for comparison for MicroScan WA. Added BD Affirm comparisons. Added acceptability criteria for all. | A. Davis | 3/9/18 |

REVIEWED BY

| Lead | Date | Coordinator/Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
| Marsha Bishoff | 9/25/17 |  | 9/7/17 9/25/17 |  | 10/3/17 |
|  |  |  | 3/13/18 |  | 3/15/18 |