| METHODIST |  |  | Page 1 of 16 | Section: UPM CHEMG | Policy #: 03.032 |
| --- | --- | --- | --- | --- | --- |
|  | CHEMISTRY, GENERAL |  | Approved by: see signature block at end of document |  | Date: 10/15/18 Review by: 10/15/20 |
|  | LABORATORY |  | Policy Created: 10/15/18 Supersedes: |  |  |
|  |  |  | Primary Responsible Parties: Ray Gross Secondary Responsible Parties: Amy Gibbs |  |  |
|  |  |  | CAP Standard: NA |  |  |
| SUBJECT: |  | LACTIC ACID |  |  |  |

Amperometric electrode

ABL90 Flex

LA

- PRINCIPLE:  The ABL90 FLEX PLUS analyzer is an in vitro diagnostic, portable, automated analyzer that quantitatively measures lactate in heparinized whole blood, capillary, venous and arterial whole blood.

Amperometry is the measuring principle employed in the sensor in the ABL90 FLEX PLUS analyzer.  The magnitude of an electrical current that flows through an electrode

chain is proportional to the concentration of the substance that is oxidized or reduced at an electrode in the chain. The amperometric measuring principle is applied in the cLac sensor. A constant polarization voltage is applied between the anode and the reference electrode. The current runs through the anode and cathode chain.  This chain is measured by an ammeter.

Dissolved lactate molecules, in solution, are transported across the outer layer of a multilayer membrane system. The lactate oxidase, immobilized between the outer and inner layers, converts lactate according to these reactions:

Lactate: Lactate + O2 ® Pyruvate + H2O2

The oxygen for this reaction is supplied by the membrane system as well as by the

oxidation of H2O2 at the platinum anode.

The H2O2 produced by the enzyme reaction is transported across the inner membrane

to the platinum anode.

- Fluid transport system

Patient samples and solutions necessary for calibration, QC measurements and other procedures are transported through the fluid transport system of the analyzer.  The sample is aspirated from the inlet, transported through the Sensor Cassette and into the waste pouch of the Solution Pack.

After a patient sample analysis the system is rinsed. The CAL 1 solution from the Solution Pack is used.

- Measurement process

The measurement process is similar for all types of measurement, patient sample analysis, built-in QC measurements, ampoule-based QC measurements, calibration-verification measurements and calibration measurements.

  - The sample (patient sample, QC solution or calibration solution) is aspirated or drawn into the sensor measurement chamber.

  - Measurements are done as soon as the sample is in the chambers. Liquid sensors control the process and can detect sample inhomogeneity and air bubbles in the sample. If any problems are found or the sample volume is too low, the measurement is aborted and the problem reported in a message attached to the result.

  - A rinse is done.

  - A status calibration is done for all parameters.

- Rinse process

A rinse is done after a measurement is completed.

- CLINICAL SIGNIFICANCE:  Elevated levels of lactate are mainly found in conditions of hypoxia such as shock, hypovolemia, and left ventricular failure.  Elevated Lactic acid is found in conditions associated with diseases such as diabetes mellitus, neoplasia, liver disease and in conditions associated with drugs or toxins such as ethanol, methanol or salicylates.

- POLICY SCOPE:  The scope of this policy applies to all Laboratory staff that prepares or performs testing on laboratory specimens at UnityPoint Methodist.

- SPECIMEN:  Fresh whole blood from a full draw tube containing sodium or lithium heparin anticoagulant (green top, non-gel tube) is acceptable.   Whole blood should be analyzed immediately, within 30 minutes from collection time for optimum results.   Samples may also be placed on ice and run within 45 minutes of the time of collection.  Maintain anaerobic conditions, do not uncap the tube or centrifuge.  Remix the tube by gentle inversion ten times prior to testing.   The sample volume for analysis is 65 ul.

- Unacceptable Conditions:

- Clotted specimens

- Insufficient quantity

- Wrong anticoagulant used for collection

- Air contamination

- Gross hemolysis

- REAGENTS/Materials:

- Solution Pack SP90, REF 944-157

- Stable until the expiration date on the label when stored at 2-25C.

- Pack must be at room temperature prior to use.

- On board stability 30 days or when the number of activities is zero.

- Sensor Cassette SC90, REF 946-008 (600 tests)

- Stable until the expiration date on the label when stored at 2-8 C.

- Once installed, the cassette is stable for 30 days.

- Sensor Cassette must come to room temperature prior to replacement. 	       Minimally 4 hours but not to exceed 48 hours.

- Inlet Gasket Holder

- Inlet Connector Gasket

- Syringe Clot Catcher, 38883 200/PK

- Flush Device, REF 905-918]

- QUALICHECK Opener/Adapter, REF 925-214

- Qualicheck 5+ external QC Level 1, REF 944-017 (red color code)

a. All three levels are stable until the expiration date on the label when

stored at 2-25 C.  If stored at 25 -32 C, stable for 15 days. Use the            solution immediately after the ampoule is opened.

- Qualicheck 5+ external QC Level 2, REF 944-018 (yellow color code)

- Qualicheck 5+ external QC Level 3, REF 944-019 (blue color code)

- Qualicheck Calibration Verification Kit, REF-VK-R5

  - Stable until the expiration date on the label when stored at 2-25

  - Ampoules must be at a constant temperature between 18-32 C for 5         hours prior to use. Use the solution immediately after the ampoule is          opened.  Ampoules are light sensitive.

- Thermal printer paper, REF 984-070

- INSTRUMENTATION/EQUIPMENT :  ABL90 FLEX

- CALIBRATION:  The analyzer automatically performs a two-point calibration based on the internal analyzer schedule.  A 1-point calibration is performed automatically with each patient sample or at 2-hour intervals if a patient sample has not been run within that time frame.

- To find the status of calibrations:

1. Tap MENU>Analyzer status

2. Tap the Calibration button

- To request an unscheduled calibration from the Analyzer status screen.

1. Tap MENU>Analyzer status

2. Tap the Calibration button

3. Select Calibration as the Calibration Type.

4. Tap the Calibration button.

- To request an unscheduled calibration from the menu:

1. Make sure that the analyzer is Ready.

2. Tap Menu>Start programs>Calibration programs>Calibration.

- Calibration results:

To find a calibration result

1. Tap Menu>Data logs>Calibration log.

2. Select the calibration.

- Understanding calibration results:

1. Bold black –a result from the current calibration

2. Dark grey—a result from a previous calibration.  The result is still valid.

3. Red and Bold red—an error occurred.   A message attached to the 		                result describes the error.

- To see messages on a calibration result:

1. Tap Menu>Data logs>Calibration log.

2. Select the calibration.

3. Tap the Result button.

4. Tap the Messages button.

- To troubleshoot messages on results:

1. Select the message.

2. Tap the Troubleshoot button.

3. Follow instructions on the screen.

- To see trends in calibration results:

1. Tap the Trend button.

2. Select the parameters.

3. Tap the View Trend button.

- Calibration Verification:

Performed every 6 months using ABL90 FLEX Qualicheck Calibration Verification Kit to meet CAP requirements.

- To set up a calibration-verification mode

It is necessary to set up a calibration-verification mode before calibration-	verification measurements can be done.

1. Tap Menu > Utilities > Setup > Analysis setup > Syringe modes.

2. Tap a button with no text in the Primary modes or Secondary modes 			    field.

3. Select the Button is enabled: check button.

4. Tap the button until the button you selected in step 2 is given the name 			    Cal. Verification.

5. Tap the button for the mode you want to be the default mode.

A small black mark in the top right-hand corner of the button shows that 		 the mode has been selected as the default.

Note: The default mode is the mode the analyzer will use if no other measurement mode is selected.

6. Tap the Close button.

- QUALITY CONTROL:  Automatic quality control management (AQM) is the name given to quality control procedures that the analyzer is programmed to do automatically.  Automatic test sequences are done with each measurement to ensure that all parts of the analyzer operate within specifications.  A System Check is performed once a day at startup or when new consumables are installed.  An analysis check is performed with every patient sample or a minimum every hour if no samples are run according to prescheduled calibration and QC activities.

- Internal – Three QC solutions in the Solution Pack are scheduled by default to be done every 8 hours according to a preset internal schedule.  Each parameter is measured with at least 3 levels of quality control material during a 24 hour period.  The assigned values and acceptable ranges for each parameter and level are entered automatically into the ABL 90 FLEX each time a new solution pack is installed.  The ABL 90 automatically assesses all automatic QC results and flags any result that is outside the acceptance ranges. If a result falls outside the acceptance range for a parameter, that particular parameter will be inactivated until a subsequent QC analysis is successful for the out of range parameter.  Measurements are scheduled by default to be done in connection with: replacement of the Solution Pack, replacement of the Sensor Cassette or at Startup.

The analyzer automatically takes action to correct a problem.   If the action fails, a message is shown and the analyzer goes into these modes giving instructions about what to do: Operator Action Needed, Troubleshooting needed or Intervention Required.

To find the latest QC result:

1.	Tap Menu>Analyzer status>Quality Control

2.	In the Built-in QC field, select the measurement

3.	Tap the Result button.

- External—The ABL90 FLEX Qualicheck Controls Level 1, 2 and 3 are run with each new lot number of Solution Pack and Sensor Cassette or each shipment of a given lot number.   Run controls in patient testing, not ampule QC and record on the QC logsheet.   Results must be within the acceptable limits defined  on the QC logsheet.

- Manual Liquid Quality Control

1. 	Check that there are no calibration errors prior to running QC.

2.       The QC is at room temperature (13-32 C) for at least 5 hours.

3.	Hold the ampoule between the thumb and first finger when you shake it.

4.	Shake the ampoule vigorously for 15 seconds before it is opened.

5.	Use the Qualicheck Opener/Adapter to hold the ampoule during the QC 	                       measurement.

6.	Measure the QC material immediately.

7.	Tap Menu>Analyzer status>Quality Control.

8.	Select the correct lot of QC solution in the Ampoule-based QC field. QC solutions are identified by a Solution name and lot number.

9.	Tap the Start Ampoule QC button.

10.	Turn the Qualicheck Opener/Adapter with the ampoule so the Radiometer 	           logo faces upwards.

11.	Put the Qualicheck Opener/Adapter with the ampoule over the inlet 	                      gasket.

12.	Push the Qualicheck Opener/Adapter with the ampoule into the analyzer 	                       as far it will go and hold it there.

13.	Hold the adapter until the analyzer tells you to remove it.  The analyzer 	                      closes the inlet.

- Entering a new QC Lot into the analyzer software:

1.	Menu

2.	Utilities

3.	Setup

4.	QC setup

5.	QC solutions: verify that the desired QC slot is highlighted.   Do not 	                      override/delete previous QC materials.  Using the manufacturer’s package 	           insert form the new lot of QC, identify the correct barcode for the ABL90 	                       FLEX.

6.	Scan the barcode.

- PROCEDURE: Prior to a sample measurement, check the Analyzer status for Ready state.   Check the color of the analyzer status traffic light located on the main screen.  Once the inlet is opened, you only have a short time to complete the actions necessary.

- If sampling from a syringe:

- Log on by scanning employee badge.

- Mix the sample by inverting the tube 10 times to ensure the sample 	is 	homogeneous.

- Pull the whole blood sample from the patient’s tube using the blood 		transfer device and a syringe.

- Tap the Syringe button.   The analyzer opens the inlet.

- Follow the instructions on the screen.

- Place and hold the lip of the syringe against the collar of the Inlet 			Gasket.

- Push the syringe into the analyzer as far as it will go and hold it 			there.  Make sure that the probe extends into the sample and stays 		there during sample aspiration.   Do not bend the probe.

- Hold the syringe in the pushed in position until the analyzer tells 			you to remove it.

- When the analyzer tells you to, remove the syringe.  The analyzer 		     closes the inlet.

- Enter the necessary data in the Patient identification screen by 			scanning the patient’s barcode.  The relevant patient identification 			input fields will be filled in automatically.

- If the Patient result screen opens before you have entered the 			necessary data, tap the ID button.

- The Patient result screen will be displayed automatically when the 			measurement has been completed.

- Check the results and the parameter status.  If there is no marking 			next to the parameter, the parameter was measured without any 			problem.

- If sampling from the vacutainer:

- Log on by scanning employee badge.

- Mix the sample by inverting the tube 10 times to ensure the sample 		is homogeneous.

- Uncap the test tube.

- Hold the test tube and tap the Syringe button.   The analyzer opens 		the inlet.

- Follow the instructions on the screen

- Select the measurement mode.

- Place and hold the lip of the test tube against the collar of the Inlet 			Gasket.

- Push the test tube into the analyzer as far as it will go and hold it 			there.  Make sure that the probe extends into the sample and stays 		there during 	sample aspiration.   Do not bend the probe.

- Hold the test tube in the pushed in position until the analyzer tells 			you to remove it.

- When the analyzer tells you to, remove the test tube.  The analyzer 		 closes the inlet.

- Enter the necessary data in the Patient identification screen by 			scanning the patient’s barcode.  The relevant patient identification 			input fields will be filled in automatically.

- If the Patient result screen opens before you have entered the 			necessary data, tap the ID button.

- The Patient result screen will be displayed automatically when the 			measurement has been completed.

- Check the results and the parameter status.  If there is no marking 			next to the parameter, the parameter was measured without any 			problem.

- To edit data in the Patient identification screen:

- Find the patient result.

- Tap the ID button

- Edit the necessary data.

- To find a patient result:

- In the data log

- Tap Menu>Data logs>Patient results log.

- Select the measurement.

- Tap the Result button.

- Latest result

- Tap Menu> Latest Result

- To find a result under My Results:

- Tap Menu>My Results

- Select the measurement.

- Tap the Result button.

- Proficiency test modes

Proficiency test is a process that lets you verify the performance of your analyzer with

accredited test solutions.

- To set up a proficiency test mode

It is necessary to set up a proficiency test mode before proficiency tests can be done.

1. Tap Menu > Utilities > Setup > Analysis setup > Syringe modes.

2. Tap a button with no text in the Primary modes or Secondary modes field.

3. Select the Button is enabled: check button.

4. Tap the button until the button you selected in step 2 is given the name Prof.

Test.

5. Tap the button for the mode you want to be the default mode.

Note: Make sure a small black mark is shown in the top right-hand corner of the

button, as the mode is then set to be default.

6. Tap the Close button.

7. Tap Menu > Utilities > Setup > Analysis setup > Syringe modes.

8. Tap a button with no text in the Primary modes or Secondary modes field.

9. Select the Button is enabled: check button.

10. Tap the button until the button you selected in step 2 is given the name Prof.

Test.

11. Tap the button for the mode you want to be the default mode.

Note: Make sure a small black mark is shown in the top right-hand corner

button, as the mode is then set to be default.

12. Tap the Close button.

- REPORTING RESULTS

Analytical Measurement Range (AMR) : 0.0 – 29.0 mmol/L

Clinical Reportable Range (CRR):  0.0 – 29.0 mmol/L.

Reference Range:  0.4 – 1.9 mmol/L

For Sepsis Protocol:   Critical Results:  >/= 2.0 mmol/L

Report results by verifying the numeric value in the LIS.

Results that print >29.0 mmol/L, report as >29 mmol/L

- PROCEDURAL NOTES/PROBLEM-SOLVING TIPS

Troubleshooting is necessary when the analyzer goes into an Operator Action Needed, Troubleshooting needed or Intervention Required mode. It may also be necessary to troubleshoot messages in the Analyzer status screen.

- About guided troubleshooting

In the troubleshooting modes, Troubleshooting needed and Operator Action

Needed modes, text and video instructions guide you through each troubleshooting procedure and show you what to do to get out of the troubleshooting mode.

After each troubleshooting procedure, the analyzer makes checks to find out if the issue has been resolved. If not, a new troubleshooting procedure is shown on the screen. If the guided troubleshooting procedures do not resolve the issue, the analyzer will go into the Intervention Required.

- To get out of Operator Action Needed mode

1. Follow the text and video instructions on the screen.

- To get out of Troubleshooting needed mode

1. Follow the text and video instructions on the screen.

- To get out of Intervention Required mode

1. Do the first action shown in the Suggested actions frame.

2. Tap the Test again button.

3.  If the analyzer does not go out of Intervention Required mode, do the 		     next action.

4.  Tap the Test again button.

5.  If the analyzer does not go out of Intervention Required mode, do steps 		     3 and 4 again.

6.  If none of the actions cause the analyzer to go out of Intervention 			     Required mode, contact your local Radiometer representative.

- Troubleshooting modes – causes

| Troubleshooting mode | Possible causes |
| --- | --- |
| Operator Action Needed | • A consumable must be replaced |
| Troubleshooting needed | • Fluid transport errors were found |
| Intervention Required | • If the troubleshooting procedures in the Troubleshooting needed mode did not resolve the issue • All other possible errors |

- To find and troubleshoot messages in the Analyzer status screen

Prerequisite(s)

• The traffic light in the Analyzer status button is yellow or red

1. Tap Menu > Analyzer status.

2. Tap the button adjacent to a yellow or red traffic light.

3. Choose an option and follow the steps for it.

| Option | Steps |
| --- | --- |
| To troubleshoot a Recommended action | Follow the instructions on the screen. |
| To troubleshoot Quality control messages | To troubleshoot errors in the Built-in QC and Ampoule-based QC fields: a) Select the quality control measurement marked by a red ? clock symbol. b) Tap the Result button. c) Tap the Messages button. d) Select the message. e) Tap the Troubleshoot button. f) Follow the instructions on the screen. To troubleshoot messages in the QC Messages field: a) Select the message. b) Tap the Troubleshoot button. c) Follow the instructions on the screen. |
| To troubleshoot Calibrations messages | To troubleshoot calibrations marked by a red ? or clock symbol. a) Select the marked calibration. b) Tap the Result button. c) Tap the Messages button. d) Select the message. e) Tap the Troubleshoot button. f) Follow the instructions on the screen. To troubleshoot messages in the Message field: a) Select the message. b) Tap the Troubleshoot button. c) Follow the instructions on the screen. |
| To troubleshoot Consumables or System messages | a) Select the message. b) Tap the Troubleshoot button. c) Follow the instructions on the screen. |

- To flush the fluid transport system:

1.  Draw tap water into the Flush Device up to the 2.5 mL mark.

2.  Pull the plunger of the Flush Device up to the 5 mL mark to draw air into it.

3. Tap the Press to start video guidance button.

4.  Pull off the inlet cover.

5.  Tap the Action completed button.

The analyzer opens the inlet.

6.  Wait until the Solution Pack is ejected.

7.  Remove the Solution Pack.

8.  Tap the Action completed button.

9.  Pull out the Inlet Gasket Holder.

10. Tap the Action completed button.

The analyzer closes the inlet.

11. Put a tissue or a cloth under the inlet.

12. Tap the Action completed button.

13. Connect the tip of the Flush Device to the waste connector in the Solution 		      Pack compartment.

14. Tap the Action completed button.

15. Hold the Flush Device as shown.

16. Inject a very small quantity of air to fill approximately 1 cm of the tube.

17. Hold the Flush Device as shown.

18. Inject a very small quantity of water to fill approximately 1 cm of the tube.

19. Do steps 15 to 18 again repeatedly to clean the fluid transport system.

20. Tap the Action completed button.

21. Inject water until an unbroken stream of water comes out of the Inlet Probe.

Note: The fluid path is flushed, when this is possible.

Note: If it is not possible, do steps 15 to 18 and step 21 again.

22. Tap the Action completed button.

23. Disconnect the Flush Device.

24. Remove the tissue or the cloth.

25. Tap the Action completed button.

The analyzer opens the inlet.

26. Put the new Inlet Gasket holder over the slide and insert it. Make sure that the Inlet Probe is in the center of the gasket and that the Inlet Gasket Holder clicks in place.

27. Tap the Action completed button.

28. Put your thumbs on the white part of the Solution Pack and push the Solution Pack into its compartment until it clicks in place. The analyzer closes the inlet.

29. Put on the inlet cover.

30. Tap the Action completed button.

- Operator actions requested in analyzer messages

- To request a tubing refill

1. Tap Menu > Start programs > Auxiliary programs > Tubing refill.

- To request a liquid sensor adjustment

Note: This procedure measures and adjusts the settings of the liquid sensors.

1. Tap Menu > Start programs > Auxiliary programs > Liquid sensor adjust.

- To request a pump calibration

Note: This procedure makes sure that the pumps in the analyzer operate 	correctly.

1. Tap Menu > Start programs > Auxiliary programs > Pump calibration.

- To request a rinse

Note: This procedure starts a rinse process. A rinse is also done after all 	measurement activities.

1. Tap Menu > Start programs > Auxiliary programs > Rinse.

- To troubleshoot Analyzer messages

This procedure can be used to find out what operator actions are necessary to troubleshoot messages.

1. Note the message number (on the left of the message).

2. Find the message and operator actions in the Analyzer messages table.

Note: The messages in the table are sorted by number.

Note: If more operator actions are available, start with the first action listed and

see if this resolves the issue. If not, continue with the next action listed.

- Shutting down, moving and restarting the analyzer:

Shutdown is a safe procedure for you to close down the analyzer. There are 2 procedures a Temporary shutdown and a Long term shutdown.

Note: Do not use the power switch to shut down the analyzer.

- When to do a temporary shutdown

Usually, the analyzer is kept switched on so that it is ready to use at any time.

However, in some situations, it is necessary to do a temporary shutdown:

• When an analyzer without a battery must be moved to a new location

• When an analyzer with a low-charge level battery must be moved to a new 		  location

• When the analyzer tells you to do a shutdown (for example, during a 	  	  	  troubleshooting procedure)

• After a non-USB keyboard or mouse is connected to an analyzer that is 	    	              switched on.

Note: If the analyzer is shut down for more than 2 hours, the Sensor Cassette 	must be replaced.

- To do a temporary shutdown

1. Tap Menu > Utilities > Temporary shutdown.

2. Tap the Confirm shutdown button.

3. Wait until the Windows program tells you that it is shutting down.

4. When Windows program has shut down, push the analyzer power switch to 	the Off position (O).

- When to do a long-term shutdown

It is usually only necessary to do a long-term shutdown when the analyzer is stored.

Refer to operator manual for long-term shutdown.

- To move an analyzer that has a charged battery

Note: The charge level of the battery must be high enough to be able to move 	the analyzer and connect it to the mains power supply before the charge level 	drops below 11 %.

1. Disconnect the power cable and peripheral devices.

2. Lift the analyzer by its handle, keep it vertical and move it to its new location.

3. Connect the power cable and peripheral devices to the analyzer.

4. Connect the analyzer to the mains power supply before the analyzer  charge-	                level falls below 10 %.

- To move an analyzer that does not have a battery

1. Do a temporary shutdown.

2. Disconnect the power cable and peripheral devices.

3. Lift the analyzer by its handle, keep it vertical and move it to its new location.

4. Connect the power cable and peripheral devices.

5. Switch on the mains power supply.

6. Push the power switch to the On position (I).

7. If the analyzer does not restart, press the standby button on the back of the

analyzer.

- To restart the analyzer after a temporary shutdown

1. Push the power switch to the On position (I).

2. If the analyzer does not restart, press the standby button on the back of the

analyzer

Note: The analyzer is ready for use when it is Ready.

- To restart the analyzer after a long-term shutdown

Required item(s)

A Solution Pack

A Sensor Cassette

1. Use the power cable to connect the analyzer to the mains power supply.

2. Push the power switch to the On position (I) and wait until the Operator-	intervention required screen is shown.

3. If the analyzer does not restart, press the standby button on the back of the

analyzer.

4. Install a Solution Pack.

5. Install a Sensor Cassette.

6. Tap the Test again button.

The analyzer is ready for use when it is Ready.

- REFERENCES

1.  ABL90 Flex Plus Instruction for use manual from software 3.4, Version 2018051.

2.  ABL90FLEX Troubleshooting Guide, Radiometer Medical ApS, Copenhagen, Denmark, 2012.

UnityPoint Methodist Laboratory is a CAP accredited facility.  As of July 1, 2011,  the responsibility of new and/or substantially revised policies and procedures will be restricted the Laboratory Director whose name appears on the CLIA certificate, and the signature appears below. The biennial review will be completed by the appointed Section Medical Director.

| POLICY CREATION : |  |
| --- | --- |
| Author: Amy Gibbs | October 15, 2018 |
| Medical Director: Elizabeth A. Bauer-Marsh, MD | October 12, 2018 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| October 12, 2018 | Elizabeth A. Bauer-Marsh, M.D. |  |
| ADMINISTRATIVE DIRECTOR |  |  |
| October 15, 2018 | Lori Rasca, DO |  |

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | New Release | Ray Gross & Amy Gibbs | 10/15/18 |

REVIEWED BY

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
| Ray Gross | 10/15/18 |  | 10/15/18 |  | 10/15/18 |