| METHODIST |  |  | Page 1 of 17 | Section: UPM CHEMC | Policy #: 01.025 |
| --- | --- | --- | --- | --- | --- |
|  | CHEMISTRY, CENTAUR |  | Approved by: see signature block at end of document |  | Date: 07/25/18 Review by: 07/25/20 |
|  | LABORATORY |  | Policy Created: 7/25/18 Supersedes: |  |  |
|  |  |  | Primary Responsible Parties: Ray Gross Secondary Responsible Parties: Amy Gibbs |  |  |
|  |  |  | CAP Standard: NA |  |  |
| SUBJECT: |  | DHEA-SO4 |  |  |  |

DHEAS

ADVIA Centaur XPT

By Chemiluminescence

LIS ORDER CODE = DHEASI

- PRINCIPLE

The ADVIA Centaur DHEAS assay is a quantitative competitive immunoassay that uses direct chemiluminescent technology. The ADVIA Centaur DHEAS assay consists of streptavidin magnetic latex particles as the solid phase, biotin-labeled anti-DHEAS antibody as the capture conjugate, and an acridinium ester-labeled DHEAS conjugate. The anti-DHEAS antibody and the acridinium ester-labeled DHEAS complex is captured by the solid phase, followed by signal development.

The ADVIA Centaur® DHEAS assay is an in vitro diagnostic immunoassay for the quantitative determination of dehydroepiandrosterone sulfate (DHEAS) in human serum and plasma using the ADVIA Centaur XPT system.

- CLINICAL SIGNIFICANCE

Measurement of circulating levels of DHEAS, an adrenal steroid, is important in investigations of abnormal hair growth (hirsutism) and balding (alopecia) in women.1,2,3 It is also of value in the assessment of adrenarche and delayed puberty.2,4 DHEAS in circulation originates almost entirely from the adrenals, though in men some may also derive from the testes.5,6 Although DHEAS is weakly androgenic, it can metabolize to more potent androgens like androstenedione and testosterone, and thus be an indirect cause of hirsutism or virilization.1,7

Plasma levels of DHEAS increase steadily from about the seventh year of life, then gradually decline after the third decade.5,6,7 The upper limit of normal for young adults is approximately 300 μg/dL for women, and 500 μg/dL for men.8

DHEAS is secreted into the bloodstream at a rate only somewhat greater than DHEA, but has a much slower turnover and a half-life of nearly a full day. It maintains plasma levels almost a thousandfold higher than DHEA.2,7,9 Unlike cortisol, DHEAS does not exhibit significant diurnal variation.2,10 Also, unlike testosterone, it does not circulate bound to sex hormone-binding globulin.7 Its abundance, together with its within-day and day-to-day stability makes DHEAS an excellent indicator of adrenal androgen output.1,2,8,11

DHEAS is often assayed in conjugation with free testosterone as an initial screen for hyperandrogenismin hirsutism.12,13 High DHEAS levels are often encountered in polycystic ovary syndrome (PCOS).11 Levels of DHEAS greater than 700–800 μg/dL in women are suggestive of a hormone-secreting adrenal tumor.9

- POLICY SCOPE

The scope of this policy applies to all Laboratory staff that prepares or performs testing on laboratory specimens at UnityPoint Methodist.

- SPECIMEN

- Collect:  Serum is the recommended sample type for this assay.

The following recommendations for handling and storing blood samples are furnished by the Clinical and Laboratory Standards Institute (CLSI, formerly NCCLS):

- Collect all blood samples observing universal precautions for 	venipuncture.

- Keep tubes stoppered and upright at all times.

- Allow serum samples to clot adequately before centrifugation.

- Stability: Test samples as soon as possible after collecting. Do not use serum samples stored at room temperature for longer than 4 hours.

- If the assay is not completed within 4 hours, tightly cap and refrigerate serum or plasma specimens at 2° to 8°C for no longer than 6 days.

- If longer storage is necessary, freeze samples at or below -20°C up to 1 month in  no frost free freezers.

- Freeze samples only once, and mix thoroughly after thawing.

- Before placing samples on the system, ensure that samples have the following characteristics:

  - Samples are free of fibrin or other particulate matter.

  - Samples are free of bubbles.

- REAGENTS

Store the reagents upright at 2–8°C.

All reagents are liquid and ready to use.

Protect reagent packs from all heat and light sources. Reagent packs loaded on 	the system are protected from light. Store unused reagent packs at 2–8°C 	away from heat and light sources.

All reagents are stable at 2–8°C until the expiration date on the packaging

Mix all primary reagent packs by hand before loading them onto the system. Visually inspect the bottom of the reagent pack to ensure that all particles are dispersed and resuspended. For detailed information about preparing the reagents for use, see the system operator’s guide.

Note

- Discard the primary reagent packs at the end of the on-system stability interval.

- Do not use reagents beyond the expiration date.

| Reagent | Description | Storage | Reagent Stability |
| --- | --- | --- | --- |
| ADVIA Centaur DHEAS ReadyPack primary reagent pack; Lite Reagent | 5.7 mL/reagent pack DHEAS(~280 pg/mL) labeled with acridinium ester in phosphate buffer, surfactant, and preservatives. | 2 – 8 C | Unopened:Stable until the expiration date on the carton On-System: 28 days |
| ADVIA Centaur DHEAS ReadyPack primary reagent pack; Solid Phase Reagent. | 5.7 mL/reagent pack streptavidin coupled to latex paramagnetic particles (~400 ug/mL) in phosphate buffered saline with bovine serum albumin, mouse serum, surfactant and preservatives. | 2 – 8 C | Unopened: Stable until the expiration date on the carton. On-System: 28 days |
| ADVIA Centuar DHEAS ReadyPack primary reagent pack; Ancillary Well Reagent | 5.7 mL/reagent pack biotinylated monoclonal mouse anti-DHEAS antibody (~4 ug/mL) I phosphate buffered saline with bovine serum albumin, mouse serum, surfactant and preservatives. | 2 – 8 C | Unopened: Stable until the expiration date on the carton. On-System: 28 days |
| ADVIA Centaur DHEAS calibrators | 2.0 mL/vial human serum with detergents and preservatives. | 2 – 8 C | Unopened: Stable until the expiration date on the vial. On-System:4 hours |
| ADVIA Centaur Wash 1a | 1500 mL/pack phosphate-buffered saline with sodium azide (<0.1%) and surfactant. | 2 –25 C | Unopened: Stable until the expiration date on the pack. On-System: 1 month |
| ADVIA Centaur Wash 1a | 2500 mL/pack phosphate-buffered saline with sodium azide (<0.1%) and surfactant | 2– 25 C | Unopened: Stable until the expiration date on the pack. On-System: 1 month |
| ADVIA Centaur ReadyPack ancillary reagent pack; Multi-Diluent 1b | 25.0 mL/reagent pack equine serum with sodium azide (0.1%) and preservatives. | 2 -8 C | Unopened: Stable until the expiration date on the pack. On-System: 28 consecutive days after accessing the ancillary reagent pack. |
| ADVIA Centaur Multi-Duluent 1b | 50.0 mL/vial equine serum with sodium azide (0.1%) and preservatives. | 2 -8 C | Unopened: Until the expiration date on the vial. |

- Materials Provided

| Reference Number | Contents | Number of Tests |
| --- | --- | --- |
| 06489701 | 1 ReadyPack primary reagent pack containing ADVIA Centaur DHEAS Lite Reagent, Solid Phase and Ancillary Well Reagent. ADVIA Centaur and ADVIA Centaur CP DHEAS Master Curve Card. 1 vial ADVIA Centaur DHEAS low calibrator. 1 vial ADVIA Centaur DHEAS high calibrator ADVIA Centaur and ADVIA Centaur CP DHEAS Calibrator Assigned Value Cards. | 50 |

- Material Required but not Provided

| Item | Description |  |
| --- | --- | --- |
| REF 01137199 (112351) | ADVIA Centaur Wash 1 | 2 x 1500 mL/pack |
| REF 03373025 | ADVIA Centaur Wash 1 | 2 x 2500 mL/pack |

- Optional Materials

| Item | Description |  |
| --- | --- | --- |
| Ref 07907174 (110313) | ADVIA Centaur Multi-Diluent 1 | 6 ReadyPack ancillary reagent packs containing 25 mL/pack |
| Ref 07293184 (110312) | ADVIA Centaur Multi-Diluent 1 | 2 ReadyPack ancillary reagent packs containing 25 mL/pack |
| Ref 09348792 (672177) | ADVIA Centaur Multi-Diluent 1 | 50 mL/vial |
| Ref 06489752 | ADVIA Centaur DHEAS Master Curve Material | 5 x 1 mL |

Safety data sheets (MSDS/SDS) available on www.siemens.com/healthcare.

CAUTION: Contains human source material. While each human serum or plasma 	donor unit used in the manufacture of this product was tested by FDA-approved 	methods and found nonreactive for hepatitis B surface antigen (HBsAg), antibody to hepatitis C (HCV), and antibody to HIV-1/2, all products manufactured using human source material should be handled as potentially infectious. Because no test method can offer complete assurance that hepatitis B or C viruses, HIV, or other infectious agents are absent, these products should 	be handled according to established good laboratory practices. 14–16

This device contains material of animal origin and should be handled as a potential carrier and transmitter of disease.

Contains  sodium azide as a preservative.  Sodium azide can react with copper and lead plumbing to form explosive metal azides. On disposal, flush reagents with a large volume of water to prevent the buildup of azides, if disposal into a drain is in compliance with federal, state, and local requirements.

Dispose of hazardous or biologically contaminated materials according to the practices of your institution. Discard all materials in a safe and acceptable manner and in compliance with prevailing regulatory requirements.

For in vitro diagnostic use.

- INSTRUMENTATION/EQUIPMENT

The ADVIA CENTAUR XPT system is an automated, immunoassay analyzer that offers optimal productivity and efficiency. No-pause reloading of reagents, samples, and supplies means that the system is always ready to process samples. All assays use direct chemiluminescent technology. Chemiluminescence is a chemical reaction that emits energy in the form of light. When used in combination with immunoassay technology, the light produced by the reaction indicates the amount of analyte in the sample. Direct chemiluminescent reactions directly measure the light energy without the use of added steps or amplifying molecules. The ADVIA Centaur assays use acridinium ester as the chemiluminescent label, since it does not require the addition of a catalyst or substrate.

When the sample start key is pressed, the barcode labels on the sample cups are read, sample is aspirated, reagent is dispensed, and the assay process begins.  Particles are magnetically separated in the cuvette incubation ring. The addition of hydroxyl groups to complete the flash reaction is accomplished by the addition of Reagent 1 & 2; Acid and Base. The chemiluminescent reaction occurs in the luminometer. The photomultiplier tube measures the chemical light reaction that takes place.

## Additional Equipment and Supplies

- Reagent Water

- Sample cups / tubes

- Cuvettes

- Sample tips

.

- CALIBRATION

The ADVIA Centaur XPT DHEAS assay requires a two point calibration when changing lot numbers of primary reagent packs, when replacing system components or when quality control results are repeatedly out of range.

The onboard stability of DHEAS reagent is 28 days.

The calibration interval for DHEAS is 14 days.

The ADVIA Centaur DHEAS assay requires a Master Curve calibration when using a new lot number of Lite Reagent and Solid Phase. For each new lot number of Lite Reagent and Solid Phase, use the barcode reader or keyboard to enter the Master Curve values on the system. The Master Curve card contains the Master Curve values. For detailed information about entering calibration values, refer to the system operating instructions.

- Calibrator Description

Refer to the Calibrator Assigned Value Card for the assigned value of each calibrator. The assigned value is traceable to the standardization of the DHEAS assay. For additional information, refer to the reagent instructions for use.

- Preparing the Calibrators

- Let the calibrators stand for 5 minutes at room temperature (20–30°C) to warm to room temperature.

- Gently mix the calibrators before using.

- Using the Calibrator Assigned Value Card

Each lot of calibrators contains a Calibrator Assigned Value card to facilitate entering the calibrator values on the system. Enter the values using the barcode wand or the keyboard. For more information about entering calibrator values, refer to the system operating instructions or to the online help system.

- Scheduling the Calibrators

Refer to Calibration Interval in the appropriate product instructions for the frequency of

calibration.

- Schedule the calibrators to the worklist:

  - Schedule DHEAS Calibrator for the DHEAS test.

  - Continue to schedule quality control for the DHEAS test.

- Label the sample cups, one with a Low Calibrator barcode label and another with a High Calibrator barcode label. Because calibrator barcode labels are lot number specific, use barcode labels that correspond to the lot number of the calibrator used.

- Gently mix the Low and High Calibrators.

- Dispense at least 300 μL of Low and High Calibrator into the labeled sample cups. The sample cups are marked at 1.0 mL intervals to assist in filling. The required calibrator volume depends on how many assays are being 	calibrated using these calibrators and the number of calibrator replicates.

- Load the calibrator sample cups on the system.

The Low Calibrator cup position must precede the High Calibrator cup position.Ensure that the DHEAS reagents are loaded on the system.

- Continue loading quality control samples.

- Start the system, if required.

- Limitations

  - Do not return any calibrators back into the vials after calibration because evaporation can occur, which may affect performance.

  - Dispose of any calibrator remaining in the sample cups after 8 hours.

  - Do not refill calibrator sample cups when the contents are depleted. If required, dispense fresh calibrators.

- QUALITY CONTROL

Follow government regulations or accreditation requirements for quality control frequency. For detailed information about entering quality control values, refer to the system operating instructions or to the online help system. To monitor system performance and chart trends, as a minimum requirement, assay 2 levels of quality control material on each day that samples are analyzed. Assay quality control samples when performing a 2-point calibration. Treat all quality control samples the same as patient samples.

Siemens recommends the use of commercially available quality control materials with at least 2 levels (low and high). A satisfactory level of performance is achieved when the analyte values obtained are within the Acceptable Control Range for the system or within your range, as determined by an appropriate internal laboratory quality control scheme.

If the quality control results do not fall within the Expected Values or within the laboratory’s established values, do not report results. Take the following actions:

  - Verify that the materials are not expired.

  - Verify that required maintenance was performed.

  - Verify that the assay was performed according to the instructions for use.

  - If necessary, contact your local technical support provider or distributor for assistance.

- PROCEDURE

Ensure that the system has sufficient primary and ancillary reagents packs.

Load the ReadyPack reagent packs in the primary reagent area using the arrow as a placement guide.   The system automatically mixes the primary reagent packs to maintain homogeneous suspension of the reagents.

If automatic dilution of a sample is required, load ADVIA Centaur Multi-Diluent 1 in the ancillary reagent area.

- Prepare the sample container for each sample, ensuring that a barcode 		label is affixed. Sample volume requirement is 25 µl plus dead space. 			Before placing samples on the system ensure that the samples are free of 		fibrin or other particulate matter.   Samples should be free of bubbles.

- Samples can be loaded on the Aptio in the appropriate lane or front 			loaded, use the appropriately coded sample racks for the type of sample 		tube to be used:

- Position 1 – aliquot tube (blue screw cap)

- Position 2 – primary sample tube

- Position 3 – sample cup (Siemens)

- Load each sample tube into a rack, ensuring that the barcode is visible through the slot in the rack.

- Place the rack(s) in the entry queue.

- Press ’START’ only if the system is not currently ‘In Process’. The analyzer will scan the barcode label and run the appropriate tests via the LIS interface

- For the DHEAS assay, the system automatically performs the following steps:

  - Dispenses 25 μL of sample into a cuvette.

  - Dispenses 100 μL of Ancillary Well Reagent and incubates for 2.75 minutes at 37°C.

  - Dispenses 100 μL of Lite Reagent and incubates for 2.75 minutes at 37°C.

  - Dispenses 100 μL of Solid Phase and incubates for 2.75 minutes at  37°C.

  - Separates, aspirates, and washes the cuvettes with Wash 1.

  - Dispenses 300 μL each of Acid Reagent and Base Reagent to initiate the chemiluminescent reaction.

  - Reports results according to the selected option, as described in the system operating instructions.

An inverse relationship exists between the amount of DHEAS present in the patient sample and the amount of relative light units (RLUs) detected by the system.

If sample is placed on the Aptio, the robotic line will send the specimen to the Centaur XPT.  Once the Centaur is done sampling the specimen, the sample is stored in the Refrigerated Storage Module.

- REPORTING RESULTS:

- Effect of Ambient Temperature on Results

CAUTION

The operating range for the ADVIA Centaur DHEAS assay is 20–25°C (68–77°F).

Do not report results for the ADVIA Centaur DHEAS assay if your laboratory temperature is below 20°C (68°F) or above 25°C (77°F). The ADVIA Centaur DHEAS assay is susceptible to ambient temperature change, which has the potential to affect recoveries of patient samples and controls. The recoveries of patient samples and control materials may change by ± 3.7% for each Celsius degree change in an ambient temperature range of 20–25°C (68–77°F).

Quality control results for this assay will reflect any temperature effects on assay results. To ensure optimal assay performance, each laboratory should determine how often quality control material is run based on the ambient temperature conditions for that laboratory

The system reports DHEAS results in ug/dL.    For umol/L: the conversion formula is 1 μg/dL x 0.02714 = μmol/L.

- Limitations of Procedure

Assay range (AMR):  3 – 1500 ug/dL

Clinical Reportable range (CRR):  3 – 3000 ug/dL

- Dilutions

The sample volume required to perform onboard dilution differs from the sample volume required to perform a single determination. Refer to the following information for the sample volume required to perform onboard dilutions:

| Serum Dilution | Sample Volume (uL) |
| --- | --- |
| 1:2 | 100 |

- The following information pertains to dilutions:

- Samples with DHEAS levels greater than 1500 ug/dL must be diluted and retested to obtain accurate results.

- Samples can be automatically diluted by the system or prepared manually.

- For automatic dilutions, ensure that ADVIA Centaur Mulit- Diluent 1 is loaded and set the system parameters as follows:

Dilution point : ≤ 1500 ug/dL

Dilution factor: 2

For detailed information about automatic dilutions, refer to the  system operating instructions.

- Manually dilute serum samples, when sample results exceed the linearity of the assay using automatic dilution, or when laboratory protocol requires manual dilution.

- Use Multi-Diluent 1  to manually dilute patient samples, and then load the diluted sample on the sample tray, replacing the undiluted sample.

- Ensure that results are mathematically corrected for dilution. If a predilution factor is entered when scheduling the test, the system automatically calculates the result.

- Interpretation of Results

Results of this assay should always be interpreted in conjunction with patient's medical history, clinical presentation and other findings.

- Expected Values:  The expected results for the ADVIA Centaur DHEAS assay were obtained using serum samples from 334 apparently healthy individuals. Based on a central 95% interval, the following reference ranges were established:

|  | Median (ug/dL) | Range Central 95% (ug/dL) |
| --- | --- | --- |
| Females | 149.4 | 25.9 – 460.2 |
| Males | 203.0 | 34.5 – 568.9 |

Report result less than3.0 as ” <3.0 ug/dL”

Report results greater than 3000 as “>3000 ug/dL”

X.      Procedural Notes/Problem –Solving Tips

- Limitations

This assay has not been validated for newborn testing.

Always use measurements of DHEAS in conjunction with other diagnostic procedures, including information from the patient’s clinical evaluation.

Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays.18  Routine exposure  to animals or to animal serum products can be prone to this interference and anomalous values may be observed. Additional information may be required for diagnosis.

- Specimens that contain biotin at a concentration of 12.5 ng/mL demonstrate a less   than or equal to 10% change in results. Biotin concentrations greater than this may lead to falsely elevated results for patient samples.

- Results from patients taking biotin supplements or receiving high-dose biotin therapy should be interpreted with caution due to possible interference with this test.

- Specificity

The ADVIA Centaur DHEAS assay shows high specificity for DHEAS. The compounds listed in the table below were spiked into a steroid-free diluent and the following results were obtained. The percent cross-reactivity is the apparent result compared to the amount added.

Percent cross-reactivity is calculated as:

% cross-reactivity = (concentration of spiked sample x 100) / amount of 				compound spiked

| Compound | Amount Added (ug/dL) | % Cross - reactivity |
| --- | --- | --- |
| DHEA | 4000 | 0.04 |
| Aldosterone | 5000 | ND |
| Androstenedione | 1000 | ND |
| Androsterone | 2000 | ND |
| Androsterone-glucuronide | 5000 | 0.01 |
| Cortisol | 10000 | ND |
| 5-dihydrotestosterone | 5000 | 0.04 |
| Estradiol | 5000 | 0.05 |
| Estriol | 5000 | 0.01 |
| Estrone | 5000 | 0.01 |
| Testosterone | 2000 | ND |
| 19-hyroxyandrostendione | 5000 | 0.04 |
| Progesterone | 5000 | ND |
| b-estradiol-3-sulfate-17-glucuronide | 5000 | 0.01 |

Interference testing was determined according to CLSI Document EP7-A2

ND= Not detectable

- Sensitivity

The ADVIA Centaur DHEAS assay measures DHEAS concentrations up to 1500 ug/dL with a minimum detectable concentration of 3 ug/dL  Minimum detectable concentration is defined as the concentration of DHEAS that corresponds to the RLUs that are two standard deviations less than the mean RLUs of 20 replicate determinations of the DHEAS zero standard.

- Precision

Precision was evaluated according to CLSI protocol EP5-A2.21 Three samples were assayed 2 times, in 2 runs, on 1 ADVIA Centaur system, (n = 80 for each sample), over a period of 20 days. The following results were obtained:

| Sample | Mean (ug/dL) | Mean (umol/L) | Within-run% CV | Total % CV |
| --- | --- | --- | --- | --- |
| Sample 1 | 56.5 | 1.5 | 7.9 | 9.0 |
| Sample 2 | 222.9 | 6.0 | 3.9 | 5.7 |
| Sample 4 | 639.5 | 17.4 | 4.8 | 5.9 |

- Accuracy / Method Comparison

The relationship between the ADVIA Centaur DHEAS assay and the IMMULITE® 2000 DHEAS was established by testing 196 serum samples on both systems. These relationships, as determined by linear regression, are described by the equation:

ADVIA Centaur DHEAS = 1.02 (IMMULITE 2000 DHEAS) + 14.2 μg/dL; r = 0.98

- Matched Sample Analysis – Plasma Versus Serum

Matched heparinized plasma, EDTA plasma, and serum samples were collected from 50 patients in the range of 25.7 to 550 μg/dL. The relationship between DHEAS recoveries for the three sample types is described by Linear regression analysis.

ADVIA Centaur DHEAS (Lithium Heparin) = 0.95 (ADVIA Centaur DHEAS 	serum) + 8.53 μg/dL

50 Sample Pairs; r = 0.989

ADVIA Centaur DHEAS (EDTA) = 1.01 (ADVIA Centaur DHEAS serum) - 2.87 	μg/dL

50 Sample Pairs; r = 0.986

- Interference

| Serum or plasma specimens that are…….. | Demonstrate </= 10% change in result up to…………. |
| --- | --- |
| Icteric | 5 mg/dL conjugate bilirubin |
| icteric | 5 mg/dL unconjugated bilirubin |
| lipemic | 1000 mg/dL triglycerides |
| Hemolyzed | 500 mg/dL hemoglobin |
| Hyperproteinemic | 5 g/dL albumin |
| Biotin-spike | 12.5 ng/mL of biotin |

- Dilution Recovery

Six serum samples in the range of 912 to 1239 μg/dL (24.8 to 33.6 μmol/L) of DHEAS were diluted 1:2 and 1:10 with Multi-Diluent 1, and assayed for recovery. The recoveries ranged from 99 to 115% with a mean of 107%.

| Sample | Dilution | Observed (ug/dL) | Expected (ug/dL) | Recovery % |
| --- | --- | --- | --- | --- |
| 1 | ------ | 1132 |  |  |
|  | 1:2 | 578 | 566 | 102 |
|  | 1:10 | 112 | 113 | 99 |
|  | Mean |  |  | 101 |
| 2 | ---------- | 1125 |  |  |
|  | 1:2 | 579 | 563 | 103 |
|  | 1:10 | 125 | 113 | 111 |
|  | Mean |  |  | 107 |
| 3 | ------- | 1205 |  |  |
|  | 1:2 | 654 | 603 | 108 |
|  | 1:10 | 130 | 121 | 107 |
|  | Mean |  |  | 108 |
| 4 | -------- | 912 |  |  |
|  | 1:2 | 494 | 456 | 108 |
|  | 1:10 | 105 | 91 | 115 |
|  | Mean |  |  | 112 |
| 5 | ------ | 1080 |  |  |
|  | 1:2 | 562 | 540 | 104 |
|  | 1:10 | 121 | 108 | 112 |
|  | Mean |  |  | 108 |
| 6 | ------- | 1239 |  |  |
|  | 1:2 | 621 | 620 | 100 |
|  | 1:10 | 140 | 124 | 113 |
|  | Mean |  |  | 107 |
| Mean |  |  |  | 107 |

- Spiking Recovery

Three patient samples were spiked with varying amounts of DHEAS and assayed. The recoveries ranged from 89 to 109% with a mean of 97%.

| Sample | Amount Added (ug/dL) | Observed (ug/dL) | Expected (ug/dL) | Recovery % |
| --- | --- | --- | --- | --- |
| 1 | ------------------- |  |  |  |
|  | 0 | 12 |  |  |
|  | 25 | 40 | 37 | 109 |
|  | 53 | 66 | 65 | 102 |
|  | 106 | 120 | 118 | 102 |
|  | 213 | 227 | 225 | 101 |
|  | 445 | 457 | 457 | 100 |
|  | 888 | 948 | 900 | 105 |
|  | Mean |  |  | 103 |
| 2 | ----------------- |  |  |  |
|  | 0 | 39 |  |  |
|  | 25 | 61 | 64 | 95 |
|  | 53 | 90 | 92 | 97 |
|  | 106 | 137 | 145 | 95 |
|  | 213 | 242 | 253 | 96 |
|  | 445 | 464 | 485 | 96 |
|  | 888 | 896 | 928 | 97 |
|  | Mean |  |  | 96 |
| 3 | ----------- |  |  |  |
|  | 0 | 183 |  |  |
|  | 25 | 196 | 207 | 95 |
|  | 53 | 209 | 236 | 89 |
|  | 106 | 263 | 289 | 91 |
|  | 213 | 366 | 396 | 93 |
|  | 445 | 558 | 628 | 89 |
|  | 888 | 961 | 1071 | 90 |
|  | Mean |  |  | 91 |
| Mean |  |  |  | 97 |

- Standardization

The ADVIA Centaur DHEAS assay is traceable to an internal standard manufactured using purified DHEAS. Assigned values of calibrators and ranges of controls are traceable to this standardization.

- Technical Assistance

For customer support, please contact your local technical support provider or distributor. www.siemens.com/diagnostics

- REFERENCES

- Braunstein GD. Coping with hirsutism – management strategies. Contemp Ob/Gyn.1981;17(5):189–93.

- Haning RV. Using DHEAS to monitor androgen disorders. Contemp Ob/Gyn.1981;18(9):117–32.

- Kasick JM, Bergfeld WF, Steck WD, Gupta MK. Adrenal androgenic female-pattern alopecia: sex hormones and the balding woman. Clev Clin Quarterly. 1983;50(2):111–22.

- Cohen HN, Wallace AM, Beastall GH, Fogelman I, Thomson JA. Clinical value of adrenal androgen measurement in the diagnosis of delayed puberty. Lancet. 1981;1:689–92.

- de Peretti E, Forest MG. Pattern of plasma dehydroepiandrosterone sulfate levels in humans from birth to adulthood: evidence for testicular production. J Clin Endocrinol Metab. 1978;47:572–77.

- Zappulla F, Ventura D, Capelli M, et al. Gonadal and adrenal secretion of

dehydroepiandrosterone sulfate in prepubertal and pubertal subjects. J Endocrinol Invest. 1981;4:197–202.

- Vermeulen A. Androgen secretion by adrenals and gonads. In: Mahesh V, Greenblatt RB, editors. Hirsutism and Virilism. Boston: John Wright–PSG; 1983:17–34.

- Lobo RA, Paul WL, Goebelsmann U. Dehydroepiandrosterone sulfate as an indicator of adrenal androgen function. Obstet Gynecol. 1981;57:69–73.

- Sciarra F. Diagnosis of virilizing syndromes: endocrinological parameters. In: Molinatti G, et al, eds. Androgenization in Women. New York: Raven Press; 1983:85–113.

- Tourniaire J, Pugeat M. Strategic approach of hyperandrogenism in women. Horm Res. 1983;18:125–34.

- Buvat J, Dewailly D, Marcolin G, Buvat-Herbaut M, Racadot A, Fossati P. Investigative strategy of hyperandrogenism in women. Horm Res. 1983;18:106–16.

- Haning RV Jr, Carlson IH, Shapiro SS, Nolten WE. Testosterone free index correlates best with dehydroepiandrosterone sulfate. Fertil Steril. 1981;36:757–65.

- Lobo RA, Paul WL, Goebelsmann U. Serum levels of DHEAS in gynecologic endocrinopathy and infertility. Obstet Gynecol. 1981;57:607–12.

- Centers for Disease Control. Update: Universal precautions for prevention of transmission of human immunodeficiency virus, hepatitis B virus and other bloodborne pathogens in healthcare settings. MMWR. 1988;37:377–82, 387–8.

- Clinical and Laboratory Standards Institute (formerly NCCLS). Protection of Laboratory Workers From Occupationally Acquired Infections; Approved Guideline - Third Edition. Wayne, PA: Clinical and Laboratory Standards Institute; 2005. NCCLS Document M29-A3.

- Federal Occupational Safety and Health Administration, Bloodborne Pathogens Standard, 29 CFR 1910.1030.

- Clinical and Laboratory Standards Institute (formerly NCCLS). Procedures for the Handling and Processing of Blood Specimens; Approved Guideline - Third Edition. Wayne, PA: Clinical and Laboratory Standards Institute; 2004. NCCLS Document H18-A3.

- Boscato LM, Stuart MC. Heterophilic antibodies: a problem for all immunoassays.Clin Chem. 1988;34:27–33.

- Clinical and Laboratory Standards Institute (formerly NCCLS). How to Define and Determine Reference Intervals in the Clinical Laboratory; Approved Guideline - Second Edition. Wayne, PA: Clinical and Laboratory Standards Institute; 2000. NCCLS Document C28-A2.

- Clinical and Laboratory Standards Institute (formerly NCCLS). Interference Testing in Clinical Chemistry; Approved Guideline - Second Edition. Wayne, PA: Clinical and Laboratory Standards Institute; 2005. NCCLS Document EP7-A2.

- Clinical and Laboratory Standards Institute (formerly NCCLS). Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline – Second Edition. Wayne, PA: Clinical and Laboratory Standards Institute; 2004. NCCLS Document EP5-A2.

- MAINTENANCE AND STORAGE:

  - All policies and procedures are reviewed every two years (safety standards are reviewed yearly) by Laboratory Administration and/or the Medical Director of the Laboratory or designee.

  - The Laboratory Administration and Medical Director review policies and procedures when there are changes in practice standards, or requirements.

  - All policies and procedures are reviewed every two years by staff or at the time new or revised ones are put in effect.

  - All policies are retained 8 years after being discontinued or revised.

  - All procedures are retained 5 years after being discontinued or revised.

UnityPoint Methodist Laboratory is a CAP accredited facility.  As of 7/1/11,  the responsibility of new and/or substantially revised policies and procedures will be restricted the Laboratory Director whose name appears on the CLIA certificate, whose signature appears below. The biennial review will be completed by the Administrative Director.

| POLICY CREATION : |  |
| --- | --- |
| Author: Amy Gibbs | July 17, 2018 |
| Medical Director: Elizabeth A. Bauer-Marsh, MD | October 23, 2018 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| July 25, 2018 | Elizabeth A. Bauer-Marsh, M.D. |  |
| SECTION MEDICAL DIRECTOR |  |  |
| October 23, 2018 | Lori Rasca, DO |  |

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 0 | Initial Release | A. Gibbs | 07/17/18 |

| REVIEWED BY |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Lead | Date | Coordinator/Manager | Date | Medical Director | Date |
| Ray Gross | 10/22/18 |  | 7/17/18 |  | 10/23/18 |