| METHODIST |  |  | Page 1 of 6 | Section: UPM BBPRO | Policy #: 02.006 |
| --- | --- | --- | --- | --- | --- |
|  | BLOOD BANK PROCEDURES |  | Approved by: see signature block at end of document |  | Date: 9/24/19 Review by: 9/24/21 |
|  | LABORATORY |  | Policy Created: 9/25/02 Supersedes 12/29/11, 8/8/13, 7/11/16, 12/21/16, 3/11/19, 5/21/19, 7/8/19, 9/24/19 |  |  |
|  |  |  | Primary Responsible Parties: Secondary Responsible Parties: June Bembenek |  |  |
|  |  |  | CAP Standard: NA |  |  |
| SUBJECT: |  | ANTIBODY IDENTIFICATION, TUBE METHOD |  |  |  |

Principle

The serum of any patient or donor in which an antibody is detected must be examined by means of a cell panel for identification of the antibody.  Knowledge of the identity of the antibody or antibodies in a patient's serum is necessary when selecting compatible donor blood for transfusion or as an aid for the physician in the management of obstetrical patients.  The serum must be tested against the cells in the manner in which the antibody was originally detected.

Scope

This policy applies to all Blood Bank technologists.

Specimen

Patient Preparation: No special preparation of the patient is required prior to specimen collection.  Blood should be collected by approved techniques.

Note for Antibody Screens and/or Compatibility Testing: If patient has been transfused with blood or blood component(s) containing red cells, or has been pregnant within the preceding three months, or if transfusion history is unknown or uncertain, sample must be drawn within three days of transfusion.  Otherwise, sample can be drawn within 7 days of transfusion.

Requirements:  K2 EDTA pink or lavender top tube is preferred.

Minimum Volume:

- Adult:  5.0 mL whole blood with Blood Bank ID band label/number

- Pediatric: 2 - K2 EDTA microtainers (each with 300-500 uL) or cord blood specimen

Specimen Stability:

If stored at room temperature 15-30°C, stable for testing 24 hours.

If stored 2-8°C, stable for 7 days.

Storage:  2-8°C for a minimum of 7 days after transfusion, or 10 days post crossmatch.

Rejection Criteria: Hemolysis.  **In rare occasions where sample cannot be redrawn, hemolyzed specimen may be used for testing as long as the testing personnel can interpret the reactions.

BEFORE ACCEPTING A TYPE AND SCREEN SPECIMEN, the technologist should visually verify that enough plasma would be left in the tube after performing the T&S to crossmatch approximately 12 units of blood.  This will avoid another sample having to be drawn and repeating testing during an emergency.

Reagents

| ###### Reagent | Storage | Stability |
| --- | --- | --- |
| Isotonic Saline | 15-30°C | Unopened: to Exp. date Open: 1 month. |
| Biotestcell Reagent Red Blood Cells, Bio-Rad | 2-8°C | Exp. date |
| Gamma Peg, Immucor | 1-10°C | Exp. date |
| Anti-Human Globulin, Anti-IgG, Rabbit, Bio-Rad | 2-8°C | Exp. date |
| Coombscell-E, Bio-Rad | 2-8°C | Exp. date |

Instrumentation/Equipment

- Test Tubes 10 x 75 mm

- Centrifuge

- Pipettes

##### Procedure

Phase 1: Immediate spin

- Label tubes to be tested with each of the panel red cells numbers (1A, 2A,…11A) and  patient control (PC).

- Add 2 drops of the patient plasma to each tube.

- Add 1 drop of corresponding reagent red blood panel cell suspension to the    panel cell labeled tubes, 1-11.

- To the patient control, add 1 drop of a 2-4% suspension of patient

cells washed in saline and mix well.

- Gently dislodge the red blood cell button and observe for macroscopic agglutination or hemolysis.  Negative reactions may be examined with an agglutination viewer, however, microscopic examination is not recommended.

- Record results

(Often the immediate centrifugation test shows expression of anti-M, -N, -P and cold reactive antibodies)

Phase 2: Incubation

- Add 2 drops of  PEG and mix tubes

- Incubate tubes at 36º - 38º C for 15 - 30 minutes.

DO NOT CENTRIFUGE. At the end of incubation, observe for hemolysis before proceeding to the wash step.

Phase 3: Indirect Antiglobulin Test (IAT)

- Wash the red blood cells three times with isotonic saline.  Decant the supernatant saline completely.

- Add 2 drops of IgG Coombs to the cell sediments and mix well.

- Centrifuge for 15-20 seconds at 3400 rpm.

- Gently dislodge the red blood cell button and observe for macroscopic agglutination.  Negative reactions may be examined with an agglutination viewer, however, microscopic examination is not recommended.

- Record results on the panel sheet.

- To control and all negative antiglobulin tests, add Coombscell-E to all negative antiglobulin tubes and centrifuge.  Positive results indicate a valid test.  Negative results indicate an invalid test and testing must be repeated.

(This test shows IgG antibodies such as anti-Duffy, anti-Kidd, all Rh-antibodies.)

Interpretation

| 4+ | Agglutination, one solid aggregate |
| --- | --- |
| 3+ | Agglutination, several large aggregates, clear background |
| 2+ | Agglutination, medium-sized aggregates, clear background |
| 1+ | Agglutination, small aggregates, turbid background |
| w+ + | Agglutination, tiny or microscopic aggregates |
| (-) (=) | No agglutination or test results negative |

Positive:  Agglutination of any panel red cell at any phase, or hemolysis at the Room Temperature or 37ºC phases, indicating the presence of an unexpected antibody.

Negative:  Absence of agglutination and hemolysis through the test procedure indicating that no unexpected antibody was present.

Guideline to Identify Antibody:

- Review reactions.  Cross off antigens present on the nonreactive red cells at each phase of testing.

- To rule out an antibody - a negative reaction on at least one cell for the antigen being tested.

- In order to confirm an antibody - a positive reaction on a minimum of two cells for the antigen being tested.

- In order to rule out other clinically significant antibodies, select additional cells from current in-date panels or out-dated panels that are less than four months old.

- Perform positive and negative QC on each additional outdated cell used for testing. Use the antisera of the antigen you are trying to confirm or exclude. You cannot use the cell to rule in or out if the QC does not give the expected reaction.

- Compare the pattern/strengths of agglutination with the antigen profiles not crossed-off.

- If one antigen remains:  the antibody is preliminarily identified.

- If more than one antigen remains, further identification steps need to be performed.

- Check for dosage on the Rh, Duffy, Kidd, and MNS groups.  The antibody may be reacting more strongly with cells that are homozygous for the     antigenic determinant than with the heterozygous cells.

- Variation in reaction strength may be seen in P1, SdA, LeA, LeB, and I. Sometimes diminished strength can be related to disease or to the length of time the suspension has been stored.

- If positive and negative reactions do not fit any pattern, it may be suggestive of multiple antibodies or antibodies of unspecified antigens.

- Test the patient red blood cells with antisera corresponding to suspected antibody.  The patient should be negative for an antigen for a corresponding antibody they exhibit.

- If patient control is positive, an auto-antibody or allo-antibody coated transfused cells could be the cause of reactivity. Perform DAT. See UPMBBPRO 02.016, Direct Antiglobulin Testing.  (If negative, it is alloantibody)

### Reporting Results

When an antibody is discovered, the following steps should be followed:

- Add on an Antibody ID in the LIS

- Enter the results of the Antibody ID.

- If only a type and screen is ordered on an inpatient or Pre-admit patient:  2 packed cells must be set up (Coombs crossmatches) and antigen testing must be done on the units. In LIS order 2 packed red blood cells to the order. Refer to Compatibility Testing and Antigen Typing for Other Blood Group Systems procedures.

- Perform a DAT on any positive Patient Control reactions.

- Place a comment in the computer for any significant information pertaining to the antibody Identification.

- An antibody identification is only necessary once a month, unless the reaction of the screening cells has changed or the antigen negative units are incompatible.

- Record all tests appropriately.

- Fill out the panel sheet completely, including patient's name, ABORH, date tested, your initials, existing antibodies, interpretation, panel results, screening cell results, direct Coombs, elution, and antigen typing results.

- Leave the panel sheet in the review folder for the Lead Tech to review.

If Anti-D is found:

- If anti D is identified on an Rh positive patient, inquire whether or not the patient has received WinRho.  If yes, see UPM BBPOL 01.018, PATIENTS RECEIVING ANTI-D FOR TREATMENT OF ITP; RELEASING ANTI-D TO PHARMACY FOR TREATMENT OF ITP, AND TRANSFUSION REQUIREMENTS for how to transfuse patients who have received WinRho.

- A pathologist and patient’s physician should discuss if patient should receive Rh pos or Rh neg blood.

- If anti-D is found due to Rh immune globulin report the antibody screen   interpretation as “Anti D due to Rh Immune Globulin given on date”.  See UPM POL 01.017, ANTI D DUE TO RHIG.

Procedural Notes/Problem-Solving Tips

- Complete grouping and antigen phenotyping of the sample may be of assistance in confirming the identity of the antibody.  Difficult samples may be sent to the Reference Laboratory at Red Cross for identification or verification of the antibody.

- Rouleaux can be identified by microscopic examination of the tubes.  The cells give the classic "stack of coins" appearance.  See Saline Replacement Procedure, UPM BBPRO 02.032-Resolving ABO Discrepancies, for special technique to be used with rouleaux.

- The patient's diagnosis and medication the patient is receiving may supply useful information.

- A cold antibody of no specificity is present if the control and all the cells are

agglutinated.  This type of antibody may exhibit wide thermal amplitude and react with Coombs when tested with alternate technique.

References

1.	AABB Technical Manual, 18th Edition, 2014, pages 400-419.

2.	Manufacturers' Directions – Bio-rad and Immucor

3.	Bio-Rad Medical Diagnostics GmbH Anti-Human Globulin (anti-IgG) insert.

4.  Immucor Gamma PeG reagent product insert.

| POLICY CREATION : |  |
| --- | --- |
| Author: Gayle Shaikh, MT Supervisor | March 5, 1978 |
| Medical Director: John Taraska, MD | March 5, 1978 |

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Minor formatting changes, added document ID to header. Updated pretransfusion sample acceptability to 30 days. | T. Mikolajczyk | 12/6/10 |
| 2 | Updated for new LIS processes. | S. Schaffer | 5/5/11 |
| 3 | Changed appropriate reagents to BIO-RAD reagents. Changed storage temperatures of reagents. Added BIO-RAD as a reference. Changed Edition number to the AABB Reference. | K. Turpin | 7/17/13 |
| 4 | Changed pretransfusion sample acceptability to 7 days. Removed LIS specific references. | V. Strow | 2/10/16 |
| 5 | Added all antibody identification will be performed on site | K. Turpin | 12/21/16 |
| 6 | Wording clarifications | J. Bembenek | 3/7/19 |
| 7 | Removed revision history #5 from Procedural Notes/Problem-Solving Tips; Updated procedure with product insert phases. | J. Bembenek | 5/6/19 |
| 8 | Added albumin step to the procedure; Updated Technical Manual to 18th edition. | - Cannon | 7/3/19 |
| 9 | Changed potentiator from 22% albumin to Gamma Peg; Minor wording changes. | - Cannon | 9/18/2019 |

Reviewed by

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
| Kathy Maher | 12/30/11 |  | 12/29/11 |  | 12/29/11 |
| Kathy Maher | 8/6/13 |  | 7/17/13 |  | 8/8/13 |
| Vincent Strow | 7/6/16 |  | 7/6/16 |  | 7/11/16 |
|  |  |  | 12/21/16 |  | 12/21/16 |
|  |  |  | 3/7/19 |  | 3/11/19 |
|  |  |  | 5/6/19 |  | 5/21/19 |
| April Cannon | 7/3/19 |  | 7/3/19 |  | 7/8/19 |
| April Cannon | 9/24/19 |  | 9/24/19 |  | 9/24/19 |