| METHODIST |  |  | Page 1 of 7 | Section: UPM HEMO | Policy #: 27 |
| --- | --- | --- | --- | --- | --- |
|  | HEMATOLOGY |  | Approved by: see signature block at end of document |  | Date: 8/23/19 Review by: 8/23/21 |
|  | LABORATORY |  | Date Created: 8/31/10 Date Revised: 8/31/10, 8/8/12, 2/4/14, 9/8/16, 9/19/17, 8/23/19 |  |  |
|  |  |  | Primary Responsible Parties: Kim Paige Secondary Responsible Parties: June Bembenek |  |  |
|  |  |  | CAP Standard: NA |  |  |
| SUBJECT: |  | SYNOVIAL FLUID CELL COUNT AND DIFFERENTIAL MANUAL AND AUTOMATED METHOD |  |  |  |

Principle

Synovial fluid is produced by dialysis of plasma across the synovial membrane and by the secretion of a hyaluronate protein complex by the synovial membrane.  The Profile consists of Appearance, Volume, TNC and RBC cell count, Differential and Crystal Identification.   A total nucleated count and red blood cell count are commonly run on the XN analyzer but can be performed by the hemacytometer method if necessary.  The differential is performed on a Wright’s stained smear. Crystals are identified by a wet mount using polarized microscopy.

Clinical Significance

Synovial fluid aspirations provide useful information for diagnosis and classification of joint disorders.  Normal WBC is less than 200 cells/mm3.  A low WBC with predominately mononuclear cells (<25% polys) suggests a non-inflammatory joint fluid such as traumatic arthritis, osteoarthritis, or systemic lupus.  A high WBC count suggests an inflammatory disease, and a high WBC (>100,000/mm3) with a high proportion of polys strongly suggest infection.

Crystal identification is useful in the differentiation of gout and pseudogout.  Monosodium urate crystals are characteristic of gouty arthritis.  Calcium pyrophosphate crystals are indicative of pseudogout   Cholesterol crystals are associated with chronic inflammatory conditions.

## POLICY SCOPE

## The scope of this policy applies to all Laboratory staff that prepares or performs testing on laboratory specimens at UnityPoint Methodist.

Specimen

| Specimen collection | Minimum of 1 ml of specimen is collected by arthrocentesis and placed in a sterile potassium EDTA lavender anticoagulated collection tube. |
| --- | --- |
| Stability | 18-25°C for 4 hours |
|  | 2-8°C for 24 hours |
| Storage/Retention | 2-8°C for 7 days |

Instrumentation/Equipment/ Reagents

Saline solution

Wright's stain

Capillary tubes

Hemacytometer Counting Chamber & Cover Glass

Pipettes

Slides

Cytocentrifuge Cups and pads

Cytocentrifuge

Microscope

Glass Test Tubes

Hema-Tek 2000 stainer

XP-50 stainer

Sysmex XN-3100 analyzer (see operating procedure for all reagents)

Capillary tubes

Slide warmer, opt.

Haematokrit 200 centrifuge

Calibration

The automated Sysmex XN-3100 is calibrated twice yearly to ensure accuracy of the system.  An ongoing verification of calibration is done by hematology technical staff by review and documentation of commercial control.  If a potential problem is noted, re-calibration will be performed by following the procedure located in the Sysmex “XN-3000 Instructions for Use” manual (Chapter 12).

Quality Control

When running body fluid on the XN analyzer:

XN CHECK BF control levels 1 and 2 will be run by technical staff in the Manual BF mode as needed before patient samples are processed. Performed once every 24 hours of patient testing.  See UPPM HEMO procedure 20 on how to run BF controls on the XN.

For performing manual body fluid counts:

Level 2 Quantrimetrix Control:  Perform once per 8 hours of patient testing.

Plate the Level 2 Control directly onto the hemacytometer and count both WBC’s and RBC’s on the 4 corner squares on both sides of the hemacytometer.  Calculate and record results on the Hemocytometer & Cell Count Control Document. UPMHEMO 24.01

With each patient testing:  Correlate the number of cells seen on the slide to the cell count.  The proportion of RBC’s and WBC’s as well as the number of cells seen are QC checks.  If there is any discrepancy, repeat both the cell count and the preparation of the smear.

Nerl Saline: When using for dilution perform background check microscopically with slide and coverslip once per day of patient testing, observe for contamination.  Document on Body Fluid Cell Count Control Document, UPMHEMO 24.01

DCL Cell Pack:  This is used for making dilutions when specimens are run on the XN analyzer.  A portion of this must be run in the BF mode, prior to analyzing the specimen you are diluting it with.  Observe for contamination.

Procedure

Specimen Preparation:  All specimens must go to microbiology first for proper routing and paper work.  Microbiology will place the volume and description for each tube submitted on the worksheet.  If this has not been done, check with microbiology before beginning processing to insure proper information on worksheet.

All body fluids of this type can be run on the XN analyzer.  Follow the XN operating procedure when running body fluids on the analyzer. When cell counts are below instrument linearity, the results may be reported as less than (<) the linearity posted on the analyzer. When cell counts are above linearity, the results may be reported as (>) than the linearity posted on the analyzer.  Specimen results (>) the linear limit will no longer have to be diluted.

Prior to running the synovial fluids on the XN-3100:  All synovial fluids must be treated with hyaluronidase.  Refer to the “CBC: Whole blood and Body Fluid Analysis on the Sysmex XN-3100” procedure.

- Very viscous synovial fluids are treated with hyaluronidase prior to performing the cell count.  Transfer a well-mixed aliquot of synovial fluid to a disposable sample cup or test tube.

- Use a wooden applicator stick to transfer a minute amount of lyophilized   hyalurondidase to the disposable sample cup or tube.  Mix with the synovial fluid specimen using the wooden applicator stick.

Note: Do not contaminate the hyaluronidase.  Use a fresh applicator stick   each time the bottle is entered.

- Allow the treated specimen to sit for 5 minutes.  Mix again just prior to pipetting.  Note: If the synovial fluid is extremely viscous the fluid may need to sit longer than 5 minutes and/or more hyaluronidase may need to be added.

I.  Manual Cell Counts using the hemacytometer method:

- Mix specimen well.

- Observe the clarity of the specimen to aid you in determining dilution to use.

  - If clear, charge both sides of the hemacytometer direct with fluid using a capillary tube.

  - If cloudy, perform a dilution using Nerl saline.  Also, remember to test a portion of your diluting fluid and check the results to make sure there isn’t any contamination. Document on the Hemocytometer &Cell count control document UPMHEMO 24.01 located next to the cytocentrifuge.

- Charge both sides of the hemacytometer.

- Perform both a WBC and RBC enumeration.

  - Count the number of cells, differentiating RBC and WBC, in the 4 corner squares on both sides (8 sqmm).  Document the cell counts on the Body Fluid Worksheet (under number 5, Cell Count).

  - If count is high, count the central square on both sides (2 sqmm).

  - The counts must agree from one side to the other within 15%.

  - If the counts do not agree, re-plate the hemacytometer and recount.

  - Do not perform an RBC on grossly bloody specimens; footnote comment "many RBCs present".

- Calculate the RBC and WBC using the following formula:

Count = total # cells X depth (10) X dilution

total #sqmm counted

6.  Place cell counts on order sheets

II. Differential

1.	Prepare 1or 2 cytospin smears: * Remember to use the hyaluronidase specimen to make your smears.

- Write name and container ID on slide with pencil.  If there is room write the specimen type as well.

- Place filter on slide with frosted side facing the filter.  Be sure filter is square      with bottom of slide (opposite end of frosted area).

- On back of slide, take felt tip pen and mark around the opening in the filter.

- Place slide and filter in holder with frosted side facing out and at the top.

- Place the cup piece next to filter paper and snap in place (slide-filter-cup).

- Place cup-assembly in cytocentrifuge.  Be sure to balance with another cup-      assembly.

- Place 2-4 drops of well-mixed fluid into cup.

- Place cap on cup.

- Close lid on centrifuge and press "MODE" to spin.

  - The speed should be set at 7 and the time at 10.

  - Centrifuge will buzz at completion of spin cycle.

- Remove the cup-assembly from centrifuge.

- Disassemble, throwing cup and filter in biohazard container.

2.	Dry slide(s) thoroughly.  To enhance drying, you may place slides on warm view      box.

3.	Stain slides with Wright's stain by placing in stainer.

4.	Perform a leukocyte differential cell count using 50X and/or 100X oil.  Include monocytes, macrophages/histiocytes cells under mononuclears.

a.  Solid tumor malignant cells should not be included in the differential, rather footnoted as being present.

b.  If abnormal cells (tumor cells, immature cells) are present, take to a     pathologist for review prior to verification.  This will allow for appropriate correlation with other ancillary tests (ex. Cytology)

5.	Correlate # cells seen on slide to cell count.  The proportion of RBCs and WBCs as well as the # of cells seen are QC checks.  If there is any discrepancy, repeat both the cell count and the preparation of the smear.

6.	Place remaining specimen in refrigerator unless needed for serology and/or      chemistry.

III.  Crystal Examination

- Make a wet preparation for polarizing microscopy by adding 1 drop of synovial fluid to patient labeled slide and add 1 drop of saline.  Place cover slip over drop.

- Place slide in moist container to prevent drying.

- Take to pathologist on clinical for polarized microscopic review for crystals.

- If no pathologist is available (evenings or weekends) to perform microscopic for crystals, place synovial specimen in Hematology's refrigerator after testing is complete.  Leave profile sheet in Hematology for tech to complete the next normal working day.

Calculations

Count = total # cells X depth (10) X dilution

total #sqmm counted

Reporting Results

- Report out in LIS.

- Manual HCT may be ordered on some specimens.  For those hematocrits that are less than 2.0% turn as <2.0

- Results will be reported as TNC (instead of WBC) in the LIS system.

Procedural Notes/Problem-Solving Tips

- Footnote any body fluid that has a clot, clumped cells, or other debris present and indicate that counts may be inaccurate.

- Document cell counts on the body fluid worksheet.

- If cytology is ordered inform the pathologist that a body fluid count and or crystals has been performed on the specimen.

- A reference interval has not been established for this test on the supplied   specimen type.  Please integrate test results into the context of the patient’s clinical condition for interpretation.

5.  If the presence of abnormal/malignant cells are present you must consult pathologist prior to reporting out.  Inform the pathologist of a cytology order if one has been requested.

References

1.	"CLINICAL DIAGNOSIS", Davidsohn & Henry, W.B. Saunders, 18th Edition.

2.	"FUNDAMENTALS OF URINE & BODY FLUID ANALYSIS", Brunzel, W.B.    Saunders, 1994.

3.	"DIAGNOSTIC HEMATOLOGY", Rodak, W.B. Saunders, 1995

| POLICY CREATION: |  |
| --- | --- |
| Author: | September 9, 1992 |
| Medical Director: Horea Baila, MD | August 31, 2010 |

| MEDICAL DIRECTOR |  |  |
| --- | --- | --- |
| DATE | NAME | SIGNATURE |
| February 11, 2017 | Elizabeth A. Bauer-Marsh, M.D. |  |
| SECTION MEDICAL DIRECTOR |  |  |
| August 8, 2012 | Wei Liu, MD, PhD |  |
| July 17, 2014 | Julia Adams, M.D. |  |

| REVISION HISTORY |  |  |  |
| --- | --- | --- | --- |
| Rev | Description of Change | Author | Effective Date |
| 1 | Removed Unopette usage, replaced with Methylene blue diluent steps, Added specimen stability information | T. Mikolajczyk | 8/31/10 |
| 2 | Added document cell counts on body fluid worksheet under cell counts 4a. Added Procedural notes/problem solving tips If cytology is ordered inform the pathologist that a body fluid count and or body fluid crystal have been performed on the specimen. | K. Turpin | 12/11/13 |
| 3 | Deleted New Methylene blue stain from reagent list and QC area and replaced with Nerl saline. Removed report only neutrophils under resulting. Added A reference interval has not been established for this test on the supplied specimen type. Please integrate test results into the context of the patient’s clinical condition for interpretation. | Kim Paige | 9/5/16 |
| 4 | Added Calibration notes for the Sysmex XN-3000 analyzer. Reagent and equipment additions. QC and specimen preparation notes. General language changes. Attached new manual body fluid worksheet. | Kim Paige | 9/10/17 |
| 5 | Changed the way WBC counts are reported and fluids that are above the linear limit will no longer have to be diluted. Minor grammatical changes. Added document references, HCT reporting and specimen preparation. | Kim Paige | 8/7/19 |

Reviewed by

| Lead | Date | Coordinator/ Manager | Date | Medical Director | Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 8/31/10 |  | 8/31/10 |
|  |  |  | 8/6/12 |  | 8/8/12 |
|  |  |  | 1/17/14 |  | 2/4/14 |
| K. Paige | 9/5/16 |  | 9/7/16 |  | 9/8/16 |
| K. Paige | 9/10/17 |  | 9/12/17 9/12/17 |  | 9/19/17 |
| Kim Paige | 8/7/19 |  | 8/19/19 |  | 8/23/19 |