Blood Bank Storage Equipment Alternate Locations

| Equipment | Acceptable Temperature Range | Alternate location | Comments |
| --- | --- | --- | --- |
| BBR1 | 1-6C | BBR4 |  |
| BBR2 | 1-6C | BBR4 or Chemistry Bally |  |
| BBR4 | 1-6C | BBR2 or Chemistry Bally |  |
| Refrigerator Option: Approved shipping box containing wet ice. Document temperature a minimum of every 4 hours |  |  |  |
| BBF1 | <-18C | BBF3 |  |
| BBF3 | <-18C | BBF1 |  |
| Freezer Option: Approved shipping box containing 20 pounds of dry ice. Document temperature a minimum of every 4 hours. |  |  |  |
| BBULF1 | <-40C | Maintain temperature with CO2 tanks until dry ice can be obtained. Approved shipping box containing 30 pounds of dry ice. Document temperature a minimum of every 4 hours. |  |
| BBPL1 | 20-24C | Return to blood supplier, if possible. Approved shipping box containing room temperature gel packs. Document temperature a minimum of every 4 hours. | Pheresis products may be without agitation for a maximum of 24 hours. If necessary, place box on top of countertop rotator. |