Related Department

ALFRED\ALF Quality

| Date: 27/04/2026 | Time: 1400 | Location: Sandringham Lab |
| --- | --- | --- |
| Attendance: | KB. JC, NX |  |

- Agenda Items Outstanding:

- Standing Items:

| Agenda Item # | Discussion | Action Required | Person Responsible | Due Date |
| --- | --- | --- | --- | --- |
| - General Housekeeping | UKG is now live. | Ensure you complete training | all |  |
| - Internal QC | CRP QC has improved with mixing reagent daily. Troponin unexpectedly went out. John remixed the reagent & it fixed the QC without cal. John also remixed the reagent bottles in the fridge. | Nil |  |  |
| - QAP | All good All staff are expected to perform some of these. Haematology keep a register on the shared drive: ‘Staff proficiencies’ |  |  |  |
| - Equipment | Alinity code yellow on 22/04/2026. ICT reservoir not filling due to pinched rubber tubing. XN 10/04/2026 Leaking samples after analysis. Valve assembly replaced. | Engineer fixed it. |  |  |
| - Tests/Methods | Quidell hsTNI in ED. Will begin mid May. Will replace i-stat TNI. | Training in progress | all |  |
| - Quality Management Review | - Internal audits of Alfred departments by Helen & John will be assigned soon. - IQM Incidents: PI-INV-310 Smart Fridge being connect to Challenger system similar to the other blood fridges. - Riskman 145732 Alinity code yellow |  |  |  |
| - Document/ SOP Updates | Please perform regular acknowledgements of SOPs. Even a few a week will help you reduce your IQM workload. |  |  |  |
| - Haematology | AMS haematology is very slow….IT are aware of this. |  |  |  |
| - Blood Bank | Platelet Rocker has been removed. P2 configuration will now be used for platelets. | P2 configuration is valid for 7 hours 25 minutes. Leave platelets in transport box until required for transfusion. |  |  |
| - Biochemistry | Diluted samples coming from G3. Please be aware John has seen several diluted samples recently. These can be identified with unexpected low K, Na, especially Total Protein. | Be alert If suspect diluted, deleted request and recollect. |  |  |
| - Work Health & Safety Issues/ Incidents | Dry Ice: Investigating an outdoor cage for sublimation. |  |  |  |
| - Staff Training | UKG Quidell hsTNI |  |  |  |
| - Staff Competency Assessments | Nil |  |  |  |
| - Continuing Education Program | UKG |  |  |  |
| - Updates from monthly APS Quality Focus Meeting. | Perform acknowledgements for documents. |  |  |  |
| - Updates from monthly APS Management Meeting. | People Matters Survey: Thankyou for participation |  |  |  |
| - Hospital activities/ meetings. | Alison Palmer, Sandy Operations, is retiring mid May. |  |  |  |
| - Other |  |  |  |  |
| - New Business |  |  |  |  |

- Date of Next Meeting: