## Slide 1: Sample Rotor Valve (SRV)
By: Michell McAdams
03/2013

## Slide 2: Cleaning SRV
Operator’s Manual
Section 9
Page 9-4

## Slide 3: Cleaning SRV
WARNING: There is a risk of infection
Make sure you are wearing gloves
Clean hands after with disinfectant
If hands become contaminated by blood, you could be infected by pathogens, seek medical attention.

## Slide 4: Cleaning SRV
Turn off the power of the Main Unit and Pneumatic Unit, wait several minutes until the pressure gauge points at “0”.
Open the Main Unit front cover
Remove tray from the sample rotor valve

## Slide 5: Cleaning SRV
Gently push down the manual rinse cup with both hands.
  Make sure the manual rinse cup is completely removed from the manual aspiration pipette.
  ***NOTE: If manual rinse cup is not completely removed there is a possibility that the manual aspiration pipette may bend when the sample rotor valve is removed.

## Slide 6: Cleaning SRV
Remove the constant-pressure screw
Remove the sample rotor valve assembly
Remove the rotor valve
  DO NOT pull out valve excessively. This is to prevent excess force from applying to the tube that is connected to the rear fixed valve
  When removing the fixed valve, have a towel ready reagent can leak
  Be careful not to bend the manual aspiration pipette.

## Slide 7: Cleaning SRV
Clean the surface of the fixed and rotary valves using distilled water or 1:10 dilution of CELLCLEAN or unscented filtered 5% bleach
  Make sure to clean them with distilled water after cleaning with CELLCLEAN or 5% bleach
  Wipe it off completely to prevent troubles to the unit or other components.
Clean the removed tray in running water

## Slide 8: Cleaning SRV
Make sure the valve contact surfaces are thoroughly free from dirt or dust
  Take care not inflict flaws or scratches on the SRV surfaces. DO NOT use each valve with dust or dirt adhered on its surface (This could cause leakage or incorrect analysis results)

## Slide 9: Cleaning SRV
Assemble the SRV in reverse order of disassembly. Gently push up the manual rinse cup with both hands
  Place rotor valve with notch facing upward and the metal knob positioned between the stoppers
  Confirm the metal knob is placed between the stoppers (Malfunctions can occur if not)
  Make sure that the manual rinse cup is pushed all the way up with manual aspiration pipette inserted (If the power is ON with the rinse cup not in place “Rinse Motor Error” will occur and makes it impossible to continue the operation.)

## Slide 10: Cleaning SRV
Turn on the power to the Main Unit. Background will check automatically. Make sure all values are within acceptable range
Run QC to make sure there are no performance problems.
Reset the SRV count (Chapter 10: 10.5 Status)

## Slide 11: The End 