| Hemoccults |  |
| --- | --- |
| Specimen - Stool specimens will be collected by the provider during a digital rectal exam - The Hemoccult Sensa card will be brought to the lab for testing - The card should be labeled with patient ID information prior to bringing it into the lab for testing - If the card is expired or blue in color, the test must be recollected - Slides can be developed within 3-5 minutes or collection or may be stored at room temperature up to 14 days before developing - Do not refrigerate or freeze cards or developer Test set up/Quality Control - Tests brought from the care unit should be considered “waiting” tests - Any lot# of developer may be used with any lot of cards as long as both are within the exp. date and are the SENSA brand - Run waiting at the clinic if from care unit - The positive and negative controls are part of the card test itself - Add two drops of developer to patient specimens - Read at 60 seconds - Run patients first, then controls to prevent bleeding - Controls require 1 drop of developer, not 2 - Read controls at 10 seconds Reading end results - Any blue color should be reported as positive. - If the slide and developer are functional, the positive control should show a blue color and the negative side will have no blue color. - If either control does not give the expected results, the patient result cannot be recorded and samples should be recollected Logsheet completion - Record lot# and expiration dates of the card and the developer - The provider may provide from 1-3 samples to be tested. Each sample should be resulted separately on the logsheet. - All results are recorded as Pos/Neg. Do not use +/- symbols. Result Entry - Always document the result on the logsheet with tech code before resulting in the computer. - After documenting on the logsheet, result in the computer with the appropriate codes. - If the cards are expired or appear blue upon receiving them, the test should be credited with the appropriate computer code | 1.Direct observation of routine patient test performance, including patient preparation, if applicable, specimen handling, processing and testing. - Collect/prepare/handle specimen - Perform testing per protocol 2.Monitoring the recording and reporting of test results. - Resulting test values 3.Review of intermediate test results or worksheets, quality control records, proficiency testing results, and preventive maintenance records. - Refer to Test result verification, logsheets and worksheets 4.Direct observation of performance of instrument/kit maintenance/internal controls and function checks. - Daily/pre-run maintenance/internal controls completed – refer to test result verification 5.Assessment of test performance through testing previously analyzed specimens, internal blind testing samples or external proficiency testing samples. - PT/Control/Known specimen tested 6.Assessment of problem solving skills. - Critical thinking quiz in MTS – see report Evaluation 3 2 1 Comments:________________________________________________ ________________________________________________________ Reviewer_______________________________Date______________ |