| Subject Stago Satellite Prothrombin | Attachments Yes No |
| --- | --- |
| Key words Coagulation, INR, PT, Satellite | Number GHI-PC-CLINIC LAB-PROCEDURES-Stago Satellite Prothrombin v. 10-2017 |
| Category | Effective Date See Electronic File |
| Manual Coagulation | Last Review Date See Electronic File |
| Issued By Clinic Laboratory Administration | Next Review Date See Electronic File |
| Applicable Clinic Laboratory Staff using the STA Satellite Analyzer | Origination Date October, 2017 |
|  | Retired Date |
| Level of Complexity Moderate | Contact Regional Clinic Laboratories Supervisors |
| Review Responsibility Regional Clinic Laboratories Supervisors | Approval Date See Electronic File |
| APPROVAL(S) Laboratory Medical Director |  |

Stago Satellite Prothrombin Clinic Lab Procedure (Page 3-8) Troubleshooting (Page 8-10) Computer Test (Page 11-12) Purpose To provide instruction for the performance of the prothrombin time (PT) test on the Stago Satellite. Principle STA®-Neoplastin CI PLUS is used for Prothrombin times on the STA ® Satellite. A mixture of thromboplastin is added to citrated plasma and the time of clot formation is determined. The STA® Satellite is a fully automated coagulation instrument, which uses an electromagnetic mechanical clot detection system. The oscillation of a steel ball within the cuvette with the thromboplastin and plasma is monitored by the STA® Satellite. When the oscillation of the steel ball is stopped by clot formation, this registers the time in seconds. The ISI value of a given thromboplastin is determined by performing PT’s on normal plasmas and Coumadin treated patient plasmas with the given thromboplastin and the WHO reference thromboplastin. The slope of this regression curve of the matched pairs is the ISI for the thromboplastin. (The ISI of the WHO reference thromboplastin is 1.0, however this may not be the ISI of your lot of Neoplastin, that information is found in the package insert for that given reagent.) The prothrombin time is a basic coagulation screening test for the assessment of congenital and acquired deficiencies of the extrinsic pathway (factors II, V, VII, X). The prothrombin time can be prolonged in certain clinical states, i.e. warfarin therapy, intestinal reabsorption disorders, liver failure, fibrinolysis and DIC. The prothrombin time is also used to monitor warfarin therapy because of its sensitivity to variations in the concentration of the Vitamin-K dependent factors II, VII , IX and X. Because of the variations in the prothrombin time results with different thromboplastins and instruments, it is recommended that the prothrombin time results be converted to an INR. The INR corresponds to the value of the ratio of the patient’s PT and the geometric mean PT of the normal reference population raised to the ISI (International Sensitivity Index) power. Specimen - Collect 2.7 mL of whole blood in a 3 mL tube of 3.2% Sodium Citrate Plasma (light blue-top tube) filled to fill line on the tube. - If using a butterfly for collection, use a blue top tube to clear the line of air before collecting the sodium citrate tube. - At minimum, tube must be 90% full to be used for testing. (Within 10% of the etched line) Specimen stability - Specimens are stable unspun and capped for 24 hours at room temperature. - DO NOT refrigerate. - If testing is delayed due to an instrument failure, batch and send specimens to Regions hospital for testing. Append –SRH. - If testing will not be completed within 24 hours, freeze an aliquot of plasma at -20° C. Unacceptable Specimens - Samples that do not meet the minimum draw requirement - Samples exceeding maximum draw requirement - Clotted samples - Hemolyzed samples Specimen processing - Centrifuge the blood specimen to obtain platelet poor plasma, as previously assessed per centrifuge, as soon as possible after collection. - Visually check the specimen for a HCT >55% or < 20% by estimating if the proportion of red blood cells to total blood drawn is greater than 55 % or less than 20%. - If HCT visually appears < 55% and > 20% proceed with testing. - If HCT suspected to be >55% or < 20% thoroughly remix sample, perform testing on ABX, Multiply ABX HCT result by 1.1. - If HCT result is <20%, append by free text “Low Hematocrit may affect result” - If HCT result is >55%, collection in a special tube adjusted for volume is required - If patient has a standing order for high hematocrit and a special tube was collected, perform testing on the special tube. - If no previous history and no standing order for special tube: - Cancel test - Notify care team - Contact patient to return for recollection - Contact St. Paul lab for further instruction Reagents Desorb Preparation: - Remove cap and discard rubber stopper - Install Maxi Reducer - Replace cap - Document open date, revised expiration date and tech code on vial Stability: - Unopened: Until manufacturer expiration date when stored at 2-8°C - Opened: 14 days stored on the instrument STA-Neoplastine CI Plus Preparation: - Pour liquid from Reagent Vial 2 into Reagent Vial 1 and tap gently - Allow to stand 30 minutes at room temperature – DO NOT mix before 30 minute incubation - Gently invert to mix - Document open date, revised expiration date and tech code on the vial - Add stir bar to the vial - Tilt vial and add mini reducer careful not to set on top of stir bar Stability: - Unopened: Until manufacturer expiration date when stored at 2-8°C - Reconstituted: 4 days when stored on the instrument STA- Coag Control Normal and Abnormal Preparation: - Reconstitute each vial of control with 1 ml of Reagent Grade water - Allow to stand at room temperature for 30 minutes - DO NOT swirl before 30 minute incubation - Gently swirl vial before use - Document open date and time, revised expiration date and tech code on the vial Stability: - Unopened: Until manufacturer expiration date when stored at 2-8°C - Reconstituted: 8 hours stored on instrument. - Quality control - Normal and abnormal quality control testing is performed every 8 hours each day of patient testing per procedure. - Attach control result printout to worksheet, document tech code - Document control results on quality control logsheet - Verify results are within acceptable range and enter results in OEM. - If controls are out of acceptable range, refer to section V. Corrective action when a control fails to perform as expected NOTE: Controls must be acceptable prior to reporting patient test results. - SATELLITE QUICK GUIDE - Procedure - NOTES: - Instrument to remain on, only shutdown for weekly maintenance. - Press any key to “wake” instrument from save mode (dark screen) - DO NOT open cover without instrument permission - Press ESC to toggle between Main menu and TEST PANEL screen - Blinking red light in front of probe indicates instrument is running – keep clear of probe I. Start up - On main menu screen: - Select STATUS<ENTER> - Use ↑↓to select System <ENTER> - Review and document temperatures “OK” on the daily log sheet - Check number of cuvettes - Check volume of wash solution - Press ESC - “Quit” is highlighted <ENTER> - ESC to return to main menu - Select C (calibration/control) - Use ↑↓to select Calibration <ENTER> - PT+ highlighted <ENTER> - Geo Mean is listed as Reference Time – document on daily log sheet - ISI is located at the bottom right on the screen – document on daily log sheet - ESC – Options - “Quit” is highlighted <ENTER> - ESC – return to TEST PANEL SCREEN II. Reagent loading Prepare reagents as instructed Place reagents in positions 1-11 on carousel to allow viewing of all reagents on one screen. - Desorb – Position 1 – it is required in this position for weekly maintenance – add maxi reducer - Neoplastin – incubate at room temperature 30 minutes - invert to mix add stir bar, tilt and add mini reducer - Controls - incubate at room temperature 30 minutes – swirl to mix NOTE: reagents containing stir bars must not be directly adjacent to one another on the carousel. A. Load by vial: RECOMMENDED On TEST PANEL screen 1. Select F2 Products: vial 2. Open cover when prompted 3. Press any key 4. Use ↑↓ to select position – when highlighted - <ENTER> 5. Selected position will be at bottom of screen 6. Carousel will rotate the selected position to be in front of the barcode reader 7. Place product in carousel with the barcode facing the reader <ENTER> - Use insert for controls - Product information will display across bottom of screen 8. Confirm volume <ENTER> (adjust as needed) 9. Confirm stability <ENTER> (adjust as needed) 10. Repeat for all reagents loaded 11. Close cover 12. ESC - Quit Carousel will initialize and display TEST STATUS screen ESC to return to TEST PANEL screen B. Load by carousel On TEST PANEL screen 1. Select F2 Products: vial 2. Open cover when prompted 3. Press any key 4. Load reagents with barcode facing reader 5. Close cover - Instrument will initialize carousel 6. A reagent will be highlighted a. verify reagent name <ENTER> b. verify reagent volume <ENTER> c. verify reagent stability <ENTER> 8. press ESC III. Quality control testing Controls must be run prior to performing any patient testing and every 8 hours of operation A. Running controls On main menu screen - Select C (calibration/control) - Use ↑↓to select Quality control <Enter> - PT+ highlighted - Press F1 to Select – a letter “S” will appear to the left - Press F10 to confirm – run selected controls - A pop up window will appear: “ Password” will be highlighted in blue, password is QC <ENTER>. - The “S” will change to a yellow rectangle - Press ESC to return to Test Status screen - Instrument will run controls - ESC - ESC – return to TEST PANEL screen - IF A CONTROL FAILS – an alarm will sound and a pop up window will appear- refer to procedure section V. Corrective action when a control fails to perform as expected. The failure must be addressed prior to continuing with testing. To silence alarm – press the space bar B. Reviewing controls 1. ESC to return to main menu screen 2. Select C (calibration/controls). 3. Use ↑↓to select Daily control <Enter> 4. F6 to print results a. attach printout to worksheet b. document tech code 5. Press ESC two times To review Levy Jennings graphs 1. Select C (calibration/controls). 2. Use ↑↓to select Quality control <Enter> 3. PT+ highlighted <ENTER> 4. Use F2/F3 to toggle between levels 5. ESC 6. Pop up window – “Return to result list” highlighted <ENTER> 7. ESC – return to TEST PANEL screen 8. Enter results in computer and record results on quality control log - IV. Patient testing- Loading samples - NOTE: Microtainer tubes must be loaded in MANUAL MODE A. Sample loading by VIAL - MANUAL MODE - cover must remain CLOSED On TEST PANEL screen - Select F1 Samples – sample loading pop up window - Verify MANUAL MODE – if in AUTO MODE, see below to select MANUAL MODE - ESC - Remove cap from sample tube - Place tube with barcode facing reader - gently press tube down to ensure it is seated at bottom of carousel - PT+ highlighted <ENTER> - F10 to validate – sample ID will move to upper box on the screen with carousel position and PT+ indicated - Repeat steps 4--7 for all samples to be tested - ESC – options - “Quit” highlighted <ENTER> - Run will start - End of run – results will print - Attach printout to worksheet and affix patient label - Document tech code - Enter results in LIS Loading a microtainer tube - Place a microtainer adapter on the carousel - Using a disposable pipette, aliquot ~ 1ml of sample into a microtainer tube and place in adapter on carousel - Select F8 to select microtainer tube - Manually enter patient L number <ENTER> - PT+ highlighted <ENTER> - Select F10 validate– sample ID will populate box on the upper screen with position and PT+ - verify presence of a “µ” to the right of the sample carousel position B. Sample loading by VIAL - AUTO MODE On TEST PANEL screen - Select F1 Samples – sample loading pop up window will appear - Verify in AUTO MODE - ESC - Remove cap from sample - Place tube with barcode facing reader – gently press tube down to seat at bottom of carousel - Press any key – sample ID and position will populate lower box - Repeat for all samples - ESC - “Quit” highlighted <ENTER> - Run will begin - End of run – results will print - Attach printout to worksheet and affix patient label - Document tech code - Enter results in LIS C. Sample loading by carousel – use for large batches that do not include microtainer tubes Must be in AUTO MODE – If in MANUAL MODE, see below to select AUTO MODE On TEST PANEL screen - Select F1 Samples – sample loading pop up window will appear - Verify in AUTO MODE - Open cover - Remove cap from sample - Place tube with barcode facing reader- gently press tube down to seat at bottom of carousel - Close cover - Carousel initialization – Reading carousel identity - Run will start - End of run – results will print - Attach printout to worksheet and affix patient label - Document tech code - Enter results in LIS D. Repeat testing On TEST PANEL screen 1. Highlight sample to be repeated <ENTER> 2. Select F4 Delete 3. Select F10 Save – this will erase previous result 4. Select F5 Insert 5. PT+ will be highlighted <ENTER> 6. Select F10 Save 7. Select F10 Run 8. Run will start 9. End of run –results will print 10. Compare repeat result to first result- Verify within acceptable repeatability range 11. Attach printout to worksheet and affix patient label 12. Document tech code 13. Enter repeat result in LIS NOTE: DO NOT use F3 Rerun. This function will report a mean of the two runs and will not report actual repeat result. E. Select MANUAL/AUTO MODE On TEST PANEL screen - Select F1 Samples – sample loading pop up window will appear - ESC - ESC –options - ↑ to highlight MANUAL/AUTO MODE - <ENTER> - ESC - “Quit” highlighted <ENTER> - ESC to return to TEST PANEL screen F. Performing a STAT test Before loading sample on carousel – MANUAL MODE - Select F12 - Load sample with barcode facing reader - The sample ID and position will populate lower box on the screen (normal tube) will change to yellow - PT+ highlighted <ENTER> After sample loaded on carousel – Run not started - ↑ to select sample ID <ENTER> - Sample ID and position will populate lower box - Select F12 STAT/ROUTINE (normal tube) will change to yellow <ENTER> After sample loaded on carousel – Run started - ↑ to highlight sample ID <ENTER> - Select F12 STAT/ROUTINE - F10 to validate - ESC - IV. Testing completed - Unloading samples and products - Unload samples: On TEST PANEL screen 1. Select F1 Sample – sample loading pop up window 2. Open Cover 3. Remove samples 4. Recap and place in storage rack 5. Close cover - Unload product by vial: On TEST PANEL Screen 1. Select F2 Products:vial – product loading pop up window 2. Open cover – press any key 3. Arrow ↑ to highlight product to be removed<ENTER> 4. Carousel will move product to barcode loading position 5. Remove product 6. Repeat steps 4-5 until all desired products removed 7. Close cover 8. ESC 9. ESC-Quit Carousel will initialize – please wait - V. Corrective action when a control fails to perform as expected - An Error pop up window will display and instrument will beep - Quality control: Control out of limits for the test PT+ - ESC to continue – pop up will disappear and alarm will stop - NOTE: On daily control screen, control will be Red and flagged Out of Range - Verify use of correct control and expiration dates - Verify volume and expiration dates of Desorb and Neoplastine - Repeat testing using same vial - Press ESC as needed to return to main menu screen a. Select C (calibration/controls) - b. Use ↑↓to select Quality control <Enter> - Window displays : Control out of range - Arrow ↑ to ACCEPT <ENTER> ( DO NOT SELECT RERUN AS FAILED DATA WILL NOT BE SAVED) - Password: CQ (super user password) <ENTER> - ESC to return to main menu - Repeat control testing per procedure Section III. A. - If results acceptable – document corrective action, record results and proceed with patient testing - If repeat failure – document corrective action and proceed to step 4. - Prepare an new vial of control - Press ESC as needed to return to main menu screen a. Select C (calibration/controls) - b. Use ↑↓to select Quality control <Enter> - Window displays : Control out of range - Arrow ↑ to ACCEPT <ENTER> - Password: CQ (super user password) <ENTER> - ESC to return to main menu - Perform control testing per procedure Section III. A. - If results acceptable – document corrective action, record results and proceed with patient testing - If repeat failure – document corrective action and proceed to step 5. - Prepare an new Desorb and/or Neoplastine - Press ESC as needed to return to main menu screen a. Select C (calibration/controls) - b. Use ↑↓to select Quality control <Enter> - Window displays : Control out of range - Arrow ↑ to ACCEPT <ENTER> - Password: CQ (super user password) <ENTER> - ESC to return to main menu - Perform control testing per procedure Section III. A. - If results acceptable – document corrective action, record results and proceed with patient testing - If repeat failure – document corrective action and proceed to step 6 . - Contact your laboratory supervisor. - CALCULATIONS The INR is automatically calculated by the STA®/STA® Satellite. The ISI is furnished by the manufacturer in the package insert and is stored in the CALIBRATION page for PT (or PT+) along with the geometric mean (reference time). The Geometric Mean is determined from the normal patient study that is completed at the start of each new lot of neoplastin. This mean can be found on the New Lot paperwork for the site. REFERENCE RANGE Reference Range; 12.0-14.5 seconds; No reference range available for children < 6 months.. Critical Range; INR ≥ 5.0; Check for clots using wooden sticks. Repeat Range; INR ≥ 5.0; Check for clots using wooden sticks. Reportable Range. Procedure Notes <mMin; PT is low; Report PT as <10 and INR as <0.7. >mMax; PT is high; Report PT as >100 seconds and INR as >11.5. Check for clots with wooden sticks before reporting either <10 or >100 seconds. Also, check for the presence of hemolysis when reviewing an INR >5.0. If appropriate make an internal comment in the LIS to document hemolysis and clot check.. - LIMITATIONS OF THE PROCEDURE 1. Sample: The slightest coagulation (micro-clots) will induce considerable shortening of the times measured (autocatalytic activation of all the factors) whereas extensive coagulation will prolong the clotting times because of consumption of factors and fibrinogen. Do not keep plasma at 2-8 °C because in this temperature range the factor VII may be activated by the kallikrein system (2). 2. Anticoagulant: Maintain the correct anticoagulant/blood sample volume ratio of 1:9. If there is any considerable variation in hematocrit, modify the quantity of anticoagulant accordingly. 3. Heparins: The STA® - Néoplastine® CI Plus test is insensitive to unfractionated heparin levels up to 1 IU/ml and to low molecular weight heparin levels up to 1.5 anti-Xa IU/ml. 4. Thrombin Inhibitors: Thrombin inhibitors (e.g., hirudin, argatroban...) present in the sample to be tested may lead to a prolonged prothrombin time for this sample. Vitamin K antagonists - Vitamin K antagonists will depress plasma levels of factors II (prothrombin), VII (proconvertin), X (Stuart factor) and IX (antihemophilic factor B). - For the assessment of the vitamin K antagonist therapy, refer to the current recommendations. - REFERENCES - STA Satellite® Operator’s Manual July 2016. - STA Satellite ® User Guide November 2011. For additional information, please refer to the manufacturer’s package inserts. - AUTHOR(S) Karen Kaestner Denise Bergo Marie LaFromboise - APPROVED BY Laboratory Medical Director or Designee

Computer Order and Result Entry for Protime/INR

- Order Code:  APRO

- WORKSHEET:  CO_ _   (COAG  Clinic)

RESULTING (Patient Results):

FUNCTION:  						MEM

TECH: 		 				<< Return>>

SHIFT: 						<< Return>>

WORKSHEET:  					CO - -

DEVICE LAB LOCATION:  (Clinic Lab Location)  	<< Return>>

TEST-1: 			 			<< Return>>

CAP METHODS TO BE USED DURING THIS RESULT ENTRY SESSION

ACCEPT (A) OR MODIFY (M)  A

NUMBER OF SPECIMENS FOR OTHER WORKLOAD, METHOD _ _ COA  	<< Return>>

WORKLOAD DATA FOR  				<< Return>>

ACC. NO.;  						Enter Accession Number

PROX:  						Enter Protime in Seconds

RINR:  						Enter INR

ACCEPT (A),   MODIFY (M),  REJECT (R)

Result Notes:

- Critical  INR is > 5.0.

- Repeat the INR for confirmation

- Notify Centralized INR Hotline (651-451-4195)

- Document with a PHON1 per protocol.

- If the Hematocrit result is < 20, append “Low Hematocrit may affect result”

- Reference range is 12.0-14.5 seconds.  There is no reference range available for children < 6 months

- Reportable Range for PT is 10-100 seconds.  Report PT as <10 or >100 seconds

- Reportable Range for INR is 0.7-11.5.  Report INR as <0.7 or >11.5

RESULTING (Quality Control):

Function:  						MEM

Tech:  							<< Return>>

Shift: 							<< Return>>

Worksheet:  						CO - -

Device Lab Location: (Clinic Lab Location)  		<< Return>>

Test-1:  << Return>>

CAP METHODS TO BE USED DURING THIS RESULT ENTRY SESSION

ACCEPT (A) OR MODIFY (M)  A

NUMBER OF SPECIMENS FOR OTHER WORKLOAD, METHOD _ _ COA  	<< Return>>

WORKLOAD DATA FOR  				<< Return>>

MANUAL RESULT ENTRY

ACC. NO.;   						C-QPTN  (Normal)

C-QPTA  (Abnormal)

PROX:  						Enter Protime in Seconds

ACCEPT (A),   MODIFY (M),  REJECT (R)

Quality Control Notes:

- Controls must be acceptable prior to reporting patient test results.

- All QC Results should be accepted on the instrument and in Sunquest and Documented on the Quality Control Worksheet.

- Attach applicable codes to the results in Sunquest

- Document steps taken to resolve QC issues on the logsheet.

| QC Modifier Codes |  |
| --- | --- |
| AC | After Calibration |
| CAL | Calibration Necessary |
| NB | New Bottle of Control |
| NBC | New Bottle of Calibrators |
| NLR | New Lot of Reagents |
| OKAG | OK according to QC Guidelines |
|  | (Use only if R12S is only QC failure) |
| RNC | Repeated-Same Bottle of Control, New Aliquot |
| RNV | Repeated/New Vial of Control |
| SAC | See Assayed Control |
| TRBL | Troubleshooting - See Log |