| Subject Stago Satellite D-Dimer | Attachments Yes No |
| --- | --- |
| Key words D-Dimer, Satellite | Number [ENTITY CODE - CATEGORY CODE - MANUAL - SEQUENCE NUMBER] |
| Category | Effective Date See Electronic File |
| Manual Clinic Laboratory Procedure Manual | Last Review Date See Electronic File |
| Issued By Clinic Laboratory – Regional Clinic Laboratories Supervisors | Next Review Date See Electronic File |
| Applicable Clinic Laboratory Staff using the STA Satellite Analyzer | Origination Date |
|  | Retired Date |
| Level of Complexity Moderate | Contact Regional Clinic Lab Supervisors |
| Review Responsibility Regional Clinic Lab Supervisors | Approval Date |
| APPROVAL(S) Laboratory Medical Director |  |

| Stago Satellite D-Dimer Specimen (page 2) Quality Control ( pages 3-4) Clinic Lab Procedure (pages 3-8) I. Purpose/principle To provide instruction for performing the d-dimer test on Stago Satellite analyzers. The specific degradation of fibrin (i.e.fibrinolysis) is the reactive mechanism responding to the formation of fibrin. Plasmin is the fibrinolytic enzyme derived from inactive plasminogen. Plasminogen is converted into plasmin by plasminogen activators. The main plasminogen activators are tissue plasminogen activator (tPA) and pro-urokinase which is activated into urokinase (UK) by, among others, the contact system of coagulation. In the bloodstream, plasmin is rapidly and specifically neutralized by alpha 2 anti-plasmin, thereby restricting its fibrinogenolytic activity and localizes the fibrinolysis on the fibrin clot. On the fibrin clot, plasmin degrades fibrin into various products (i.e., D-Dimers). Antibodies specific for these products, which do not recognize fibrinogen, have been developed. The presence of these various fibrin degradation products, among which D-Dimer is the terminal product, is the proof that the fibrinolytic system is in action in response to coagulation activation. Clinical applications for this test are as follows: Disseminated Intravascular Coagulation (DIC), negative predictor for the diagnosis of a thrombotic episode (i.e., DVT, PE), efficacy of treatment for a thrombotic episode and screen for possible re-occurrence (MI), and screen for other activation states of coagulation (i.e., post-operative, cancer, cirrhosis). When a beam of monochromatic light is allowed to transverse a suspension of microlatex particles to which specific antibodies have been attached by covalent bonding and if the wavelength of the light is much greater than the diameter of the latex particles, the light is only slightly absorbed. In the presence of the antigen being tested for, the antibody-coated latex particles agglutinate to form aggregates of a diameter greater than the wavelength of the light, more of the light is absorbed. This increase in light absorption is a function of the antigen level present in the test sample. II. SPECIMEN 3.0 ml sodium citrate (3.2%) blue top tube. Tube contains 2.4 ml sample and 0.3 ml anticoagulant. - At minimum, tube must be 90% full. Do not overfill sodium citrate tube - Stability: 4 hours room temperature – plasma on cells in original tube - 8 hours room temperature – plasma removed from cells - If using a butterfly, use a blue top sodium citrate discard tube to clear the line of air before collecting the sodium citrate tube to ensure proper anticoagulant ratio. - If testing is delayed due to instrument failure, send specimens to Regions for testing Processing: - Centrifuge to obtain platelet poor plasma - Visually check tube for HCT >55% by estimating if the proportion of red blood cells to total blood drawn is greater than 55% - If HCT appears <55% proceed with testing - If HCT suspected to be >55%, thoroughly remix sample, perform testing on the Sysmex. Multiply Sysmex HCT by 1.1 - If HCT results is >55%, collection in a special tube adjusted for volume is required - If patient has a standing order for high hematocrit and a special tube was collected, perform the testing on the special tube. - If no previous history and no standing order for a special tube: - Cancel test - Notify care team - Contact patient to return for a recollection - Contact St. Paul lab for further instruction Unacceptable specimens: Hemolyzed, clotted, lipemic or short draw specimens III. Reagents LIATEST D-Di Buffer; LIATEST D-Di Latex; LIATEST Controls N+P; LIATEST Owren-Koller Buffer. Description; Suspension of microlatex particles coated with two diferent mouse monoclonal anti-human D-Dimer antibodies, then stabilized with bovine albumin; Citrated control plasma, normal and abnormal levels. Reconstitute with; Ready to use; Ready to use; 1.0 ml NERL water; Ready to use. Equilibration time; 15 minutes RT; 15 minutes RT; 30 minutes RT; 30 minutes RT. Stability; Unopened: manufacturer expiration date at 2-8°C Post open: 15 days stored on the Satellite; Unopened: date on label at 2-8°C Post open: 15 days stored on the Satellite; *See note; Per Stago: Aliquot 1 ml into micro cups, store at 2-8 until manufacturer expiration date; use adaptor to place on instrument – 6 days stored on instrument. Notes; Place the STA mini reducer in the vial Replace perforated plastic cap on vial Swirl gently before loading Remove any foam or bubbles; Swirl gently before loading; Used for dilutions. Loading Reagent on Satellite; - From Test panel, press F2. - Open product cover. - Press any key (other than ALT or Symb). - Position cursor on the line corresponding to the vial position. Use the ↑↓ keys and [ENTER]. The analyzer places the selected position in front of the barcode window. - Place the vial in the presented position. Barcode label must face barcode reader window. The analyzer will detect the vial. - Confirm the volume/stability of the reagent by using the [ENTER] key. - Loading a new lot of reagent: - If loading a new lot of reagent, a pop up will display stating “All concerned Calibrations will be definitively invalided. Accept Yes/ No. - Type “YES” <ENTER> - Scan reagent barcode sheet (found inside the box of reagent) - <ESC> to cancel message. Note: you will only have to scan one barcode sheet per lot of reagent - Repeat steps 4-6 to load additional reagent.. *NOTE: Control stability: - Stability studies were conducted to prove that controls can be used for up to 36 hours. - Stability studies were conducted with the controls being refrigerated overnight. The open and expiration dates and initials will be recorded on the bottle. I suggest we validate stored on the instrument - Controls will be used for 36 hours then discarded. IV. CALIBRATION Kit reagents are pre-calibrated: this pre-calibration is valid for all the kits of the same lot. To enter the calibration on the analyzer, scan the barcode printed on the Assay Value insert across the barcode reader of the instrument. The calibration values for the lot of reagents being used will subsequently be validated after the two D-Dimer control levels have been determined. The calibration curve can be examined on the screen of the analyzer in the “calibration” menu. Note: calibration must be ran with each new lot of reagent- see “Running” Calibration” instructions below. A. Running Calibration On main menu screen 1. Select C (calibration/control) 2. Use ↑↓to select Calibration <Enter> 3. Use ↑↓ to select D-Di 4. Press F1 to Select – a letter “S” will appear to the left 5. Press F10 to confirm – run selected calibration 6. A pop up window will appear: “Password” will be highlighted in blue, enter password<ENTER>. o The “S” will change to a yellow rectangle 7. Press ESC to return to Test Status screen 8. Instrument will calibration 9. ESC 10. ESC – return to TEST PANEL screen Note: Quality control will not need to be ran in addition to the calibration. 11. Print Calibration: - From main menu, select C (calibration/ control) - Use ↑↓to select Calibration <Enter> - Use ↑↓ to select D-Di - <Enter> - Press ESC for options - Use ↑↓ to select PRINT - Give printed calibration report to lab supervisor V. QUALITY CONTROL Normal and abnormal quality control testing is performed every 8 hours each day of patient testing. - Results interface into LIS. - Verify results are within acceptable range. NOTE: Controls must be acceptable prior to reporting patient test results. - All QC Results should be accepted on the instrument and in EPIC Beaker. - Document steps taken to resolve QC issues A. Running controls On main menu screen 1. Select C (calibration/control) 2. Use ↑↓to select Quality control <Enter> 3. Move cursor to select D-Dimer 4. Press F1 to Select – a letter “S” will appear to the left 5. Press F10 to confirm – run selected controls 6. A pop up window will appear: “Password” will be highlighted in blue, enter password <ENTER>. o The “S” will change to a yellow rectangle 7. Press ESC to return to Test Status screen 8. Instrument will run controls 9. ESC 10. ESC – return to TEST PANEL screen IF A CONTROL FAILS – an alarm will sound and a pop up window will appear- refer to procedure section V. Corrective action when a control fails to perform as expected. The failure must be addressed prior to continuing with testing. To silence alarm – press the space bar B. Reviewing controls 1. ESC to return to main menu screen 2. Select C (calibration/controls). 3. Use ↑↓to select Daily control <Enter> 4. Press ESC two times To review Levy Jennings graphs 1. Select C (calibration/controls). 2. Use ↑↓to select Quality control <Enter> 3. Move cursor to select D-Dimer <ENTER> 4. Use F2/F3 to toggle between levels 5. ESC 6. Pop up window – “Return to result list” highlighted <ENTER> 7. ESC – return to TEST PANEL screen 8. Results interface into LIS. VI. PROCEDURE NOTE: Microtainer tubes must be loaded in MANUAL MODE Load patient samples – all patients’ plasmas are tested undiluted. If any of the patient results falls outside of the working range of the assay, the instrument automatically retests the sample at an appropriate dilution provided the procedure with dilution has been chosen. A. Sample loading by VIAL - MANUAL MODE - cover must remain CLOSED On TEST PANEL screen 1. Select F1 Samples – sample loading pop up window 2. Verify MANUAL MODE – if in AUTO MODE, see below to select MANUAL MODE 3. ESC 4. Remove cap from sample tube 5. Place tube with barcode facing reader - gently press tube down to ensure it is seated at bottom of carousel 6. Move cursor to select D-Dimer <ENTER> 7. F10 to validate 8. Repeat steps 4--7 for all samples to be tested 9. ESC – options 10. “Quit” highlighted <ENTER> 11. Run will start 12. End of run – results will print – the printout must be retained for 2 years 15. Results interface into LIS. Loading a microtainer tube 1. Place a microtainer adapter on the carousel 2. Using a disposable pipette, aliquot ~ 1ml of sample into a microtainer tube and place in adapter on carousel 3. Select F8 to select microtainer tube 4. Manually enter patient L number <ENTER> 5. Move cursor to select D-Dimer <ENTER> 6. Select F10 validate 7. Verify presence of a “μ” to the right of the sample carousel position B. Sample loading by VIAL - AUTO MODE On TEST PANEL screen 1. Select F1 Samples – sample loading pop up window will appear 2. Verify in AUTO MODE 3. ESC 4. Remove cap from sample 5. Place tube with barcode facing reader – gently press tube down to seat at bottom of carousel 6. Press any key – sample ID and position will populate lower box 7. Repeat for all samples 8. ESC 9. “Quit” highlighted <ENTER> 10. Run will begin 11. End of run – results will print – the printout must be retained for 2 years 12. Results interface into LIS. C. Sample loading by carousel – use for large batches that do not include microtainer tubes Must be in AUTO MODE – If in MANUAL MODE, see below to select AUTO MODE On TEST PANEL screen 1. Select F1 Samples – sample loading pop up window will appear 2. Verify in AUTO MODE 3. Open cover 4. Remove cap from sample 5. Place tube with barcode facing reader- gently press tube down to seat at bottom of carousel 6. Close cover 7. Carousel initialization – Reading carousel identity 8. Run will start 9. End of run – results will print- the printout must be retained for 2 years 10. Results interface into LIS. D. Select MANUAL/AUTO MODE On TEST PANEL screen 1. Select F1 Samples – sample loading pop up window will appear 2. ESC 3. ESC –options 4. ↑ to highlight MANUAL/AUTO MODE 5. <ENTER> 6. ESC 7. “Quit” highlighted <ENTER> 8. ESC to return to TEST PANEL screen Unloading samples and products - Testing completed A. Unload samples: On TEST PANEL screen 1. Select F1 Sample – sample loading pop up window 2. Open Cover 3. Remove samples 4. Recap and place in storage rack 5. Close cover B. Unload product by vial: On TEST PANEL Screen 1. Select F2 Products:vial – product loading pop up window 2. Open cover – press any key 3. Arrow ↑ to highlight product to be removed<ENTER> 4. Carousel will move product to barcode loading position 5. Remove product 6. Repeat steps 4-5 until all desired products removed 7. Close cover 8. ESC 9. ESC-Quit Carousel will initialize – please wait VII. Corrective action when a control fails to perform as expected An Error pop up window will display and instrument will beep Quality control: Control out of limits for the test D-Di ESC to continue – pop up will disappear and alarm will stop NOTE: On daily control screen, control will be Red and flagged Out of Range 1. Verify use of correct control and expiration dates 2. Verify volume and expiration dates of reagents 3. Repeat testing using same vial Press ESC as needed to return to main menu screen a. Select C (calibration/controls) b. Use ↑↓to select Quality control <Enter> Window displays : Control out of range c. Arrow ↑ to ACCEPT <ENTER> ( DO NOT SELECT RERUN AS FAILED DATA WILL NOT BE SAVED) d. Enter password <ENTER> e. ESC to return to main menu f. Repeat control testing per procedure Section III. A. - If results acceptable – document corrective action in LIS, record results and proceed with patient testing - If repeat failure – document corrective action in LIS and proceed to step 4. 4. Prepare an new vial of control Press ESC as needed to return to main menu screen a. Select C (calibration/controls) b. Use ↑↓to select Quality control <Enter> Window displays : Control out of range c. Arrow ↑ to ACCEPT <ENTER> d. Enter password <ENTER> e. ESC to return to main menu f. Perform control testing per procedure Section III. A. - If results acceptable – document corrective action in LIS, record results and proceed with patient testing - If repeat failure – document corrective action in LIS and proceed to step 5. 5. Contact your laboratory supervisor. VIII. CALCULATIONS IX. REFERENCE RANGE Reference Range: ≤ 0.50 ug/ml (FEU) Reportable Range: 0.27-20 ug/ml (FEU) D-Dimer result >mMax report as ≥ 20.0 ug/ml (FEU) X. LIMITATIONS - Cloudy plasmas may lead to an under-estimation of the D-Dimer level. - Concentrations of fibrinogen degradation products greater than 15 ug/ml may lead to an over-estimation of the D-dimer level. - The presence of rheumatoid facto at a level greater than 50 IU.ml may lead to an over –estimation of the D-Dimer level. - The STA® – Liatest® D-DI is insensitive to the following substances: - Hemoglobin ( up to 2 g/l0, conjugated bilirubin 9 up to 290 mg/l), unconjugated bilirubin ( up to 200 mg/l), unfractionated heparin ( up to 2 IU/ml), low molecular weight heparin ( up to 2 anti-Xa IU/ml). Tests performed according to CLSI guideline EP07-A2 - The presence of anti-bovine albumin and/or anti-mouse antibodies in certain subjects may lead to an over-estimation of the D-Dimer level. - Patient with distal DBT may have a normal D-Dimer level. - D-Dimer assay should not be used in patients with high PTP score - The d-dimer level increases during pregnancy. It also rises with age. XI. References 1. STA Satellite® Operator’s Manual July 2016. 2. STA Satellite ® User Guide November 2011. 3. GHI-PC-CLINIC LABPROCEDURES-Stago Satellite Prothrombin v.04-2019 4. Park Nicollet Health Services Procedure Stago Satellite D-Dimer (H) Document# 106226 5. STA® – Liatest® D-DI package insert 6. STA® – Liatest® Control N+P package insert 7. STA® – Liatest® Owen-Koller Buffer package insert XII. Compliance # Failure to comply with this policy or the procedures may result in disciplinary action, up to and including termination. XIII. AUTHOR(S) MLLaFromboise |
| --- |
| XIV. APPROVAL(S) |  |  |  |
| Laboratory Medical Director or Designee |  |  |  |