|  | Pregnancy Test Procedure-Serum | Attachments Yes No |
| --- | --- | --- |
|  | Key words hCG, serum pregnancy, Cardinal | Number GHI-PC-CLINIC LAB-PROCEDURES-Pregnancy Test Serum v. 04-2019 |
|  | Category Provision of Care | Effective Date See Electronic File |
|  | Manual Clinic Laboratory Procedure Manual | Last Review Date See Electronic File |
|  | Issued By Clinic Laboratory Administration – Laboratory Technical Consultants | Next Review Date See Electronic File |
|  | Applicable Clinic Laboratory Staff | Origination Date October 2016 |
|  |  | Retired Date |
|  | Level of Complexity Non-Waived | Contact Regional Clinic Laboratory Supervisors |
|  | Review Responsibility Regional Clinic Laboratory Supervisors | Approved Date |
| APPROVAL(S) Laboratory Medical Director |  |  |

Pregnancy Test – Serum

Purpose/PRINCIPLE

Human Chorionic Gonadotropin (hCG) is a glycoprotein produced by the developing placenta shortly after fertilization.  In normal pregnancy, hCG can be detected in both serum and urine as early as 7-10 days after conception.  hCG levels continue to rise very rapidly, frequently exceeding 100mIU/mL by the first missed menstrual period, and peaking in the 100,000-200,000 mIU/mL range about 10-12 weeks into pregnancy.  The appearance of hCG in both serum and urine soon after conception and its subsequent rapid rise in concentration during early gestational growth, make it an excellent marker for the early detection of pregnancy.

The hCG Combo Rapid Test is a rapid chromatographic immunoassay for the qualitative detection of hCG in serum to aid in the early detection of pregnancy.  The test utilizes a combination of antibodies including mouse monoclonal anti-hCG antibodies and goat polyclonal anti-hCG antibodies to selectively detect elevated levels of hCG.  The assay is conducted by adding serum specimen to the specimen well of the test cassette and observing the formation of colored lines.  The specimen migrates via capillary action along the membrane to react with the colored conjugate.

Positive specimens react with the specific colored antibody conjugates and form a colored line at the test line region of the membrane.  Absence of this colored line suggests a negative result.  To serve as a procedural control, a colored line will always appear at the control line region if the test has been performed properly.

POLICY

Laboratory Staff will follow the approved techniques outlined in this procedure.

REAGENTS/MATERIALS

1.  Individually packaged test cassette each containing one disposable specimen dropper. The test cassette      contains mouse anti-beta hCG antibody conjugated to colloidal gold and goat anti-alpha hCG antibody coated on the membrane

2.  External hCG controls – Sure-Vue® Control Set

3.  Timer

KIT STORAGE AND STABILITY

Store as packaged in the sealed pouch at 2°- 30 °C (36°– 86° F).  The test cassette is stable through the expiration date printed on the sealed pouch.  The test cassette must remain in the sealed pouch until use.  DO NOT FREEZE.  Do not use beyond the expiration date.

QUALITY CONTROL

External quality control:

- Performed first workday of each month, with every new lot/shipment and as a guide in interpretation of a questionable result.

- Negative:  perform testing the same as you would a patient sample

- Positive:  25 mIU/mL -  perform testing the same as you would a patient sample

- Document control test results in LIS.

* Control set includes Negative, Positive 25 mIU/mL and 250 mIU/mL.  Use the 25 mIU/mL for testing.

Storage and stability: Store at 2°- 8°C (36°– 46° F) for the duration of the shelf-life.  The controls must be equilibrated to room temperature 18°- 25°C (64°– 77° F) prior to testing.

External controls must perform as expected prior to reporting patient results.  If a control fails to perform as expected refer to corrective action section below.

Internal quality control: Internal procedural controls are included in the test.

- A red line appearing in the control region (C) is the internal procedural control.  It confirms sufficient specimen volume and correct procedural technique

- A clear background is an internal negative background control.  If the test is working properly, the background in the result area should be white to light pink and not interfere with the ability to read the test result.

Internal controls must perform as expected prior to reporting patient results

SPECIMEN COLLECTION AND STORAGE

Serum:  Blood should be collected aseptically into a clean tube without anticoagulants (serum red/gold tube). Separate the serum from blood as soon as possible to avoid hemolysis.  Use clear non-hemolyzed specimens when possible.

Specimen Stability:  stored at 2°- 8°C (36°– 46° F) for up to 48 hours prior to testing.  For prolonged storage, specimens may be frozen and stored below -20°C (-4° F).  Frozen specimens should be thawed and mixed before testing.

HCGIF specimens should be tested immediately.  If positive, the specimen should be sent to Central Laboratory for quantitative testing and results.

PROCEDURE

Allow the test cassette, serum specimen and/or controls to equilibrate to room temperature 18°- 25°C (64° – 77° F) prior to testing.

- Remove the test cassette from the sealed pouch and use immediately.

- Label the cassette with a patient identification label.

- Place the test cassette on a clean and level surface.

- Using the disposable dropper provided, aspirate specimen and holding the dropper vertically, transfer 3 full drops of serum (approx 100uL)  to the specimen well of the test cassette.  Avoid trapping air bubbles in the specimen well.

- Start the timer for 5-6 minutes.

- Wait for the red line(s) to appear.  Read the result at 5-6 minutes.  DO NOT interpret results after the appropriate read time.  It is important that the background is clear before the result is read.

INTERPRETATION OF RESULTS

Positive:  Two distinct red lines appear.  One line should be in the control region (C) and another line should be in the test region (T).

* Note:  A sample hCG concentration below the cut-off level of this test might result in a weak line appearing in the test region (T) after an extended period of time.  A line in the test region (T) seen after the read time could be indicative of a low hCG level in the sample.

If such results are seen, it is recommended that the test be repeated with a new sample in 48-72 hours or that an alternate confirmation method is used.

Negative:  One red line appears in the control region (C).  No apparent red or pink line appears in the test region (T).

Invalid:  control line fails to appear.  Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure.  Review the procedure and repeat the test with a new test cassette.  If the problem persists, discontinue using the test kit immediately and call Technical services.

REPORTING RESULTS

Test result:  Positive, Negative or Inconclusive

Procedural control:  Valid or Invalid

Background control: Valid or Invalid

If either of the internal controls are invalid, the test cannot be reported and must be repeated.

Refer to EPIC Beaker result entry procedure

CORRECTIVE ACTION WHEN A CONTROL FAILS TO PERFORM AS EXPECTED

- Verify use of  correct control and kit,  correct lot number, expiration dates acceptable

- Remix and repeat

- If acceptable – document corrective action, record results and proceed with patient testing

- If repeat failure – document corrective action and proceed to step 3.

- Open a new  control, mix and test

- If acceptable – document corrective action, record results and proceed with patient testing

- If repeat failure – document corrective action and proceed to step 4.

- Open a new kit

- If acceptable – document corrective action, record results and proceed with patient testing

- If repeat failure – document corrective action and proceed to step 5.

- Contact your laboratory technical consultant

LIMITATIONS

- False negative results may occur when the levels of hCG are below the sensitivity level of the test.  When pregnancy is still suspected, a first morning serum specimen should be collected 48 hours later and tested.

- Very low levels of hCG (less than 50mIU/mL) are present in serum shortly after implantation.  However, because a significant number of first trimester pregnancies terminate for natural reasons, a test result that is weakly positive should be confirmed by retesting with a first morning serum specimen collected 48 hours later.

- This test reliably detects intact hCG up to 500,000 mIU/mL.  It does not reliably detect hCG degradation products, including free-beta hCG and beta core fragments.  Quantitative assays used to detect hCG may detect hCG degradation products and therefore may disagree with the results of this rapid test.

- A number of conditions other than pregnancy, including trophoblastic disease and certain non-trophoblastic neoplasms including testicular tumors, prostate cancer, breast cancer and lung cancer, cause elevated levels of hCG.  Therefore, the presence of hCG in a serum specimen should not be used to diagnose pregnancy unless these conditions have been ruled out.

- As with any assay employing mouse antibodies, the possibility exists for interference by human anti-mouse antibodies (HAMA) in the specimen.  Specimens from patients who have received preparations of monoclonal antibodies for diagnosis or therapy may contain HAMA.  Such specimens may cause false positive or false negative results.

- This test provides a presumptive diagnosis for pregnancy.  A confirmed pregnancy diagnosis should only be made by a physician after all clinical and laboratory findings have been evaluated.

- ReferenceS

- hCG Combo Rapid Test package insert, CardinalHealth DN:  1155904805

- RELATED DOCUMENTS

- APPENDIXES

- NA

AUTHOR/REVIEWER(S)

DJEaton

M. LaFromboise

Definitions

Compliance

Failure to comply with this policy or the procedures may result in disciplinary action, up to and including termination.

ENDORSEMENT

Laboratory Administration