Body Fluid Cell Count (Manual) Procedure

Hematology

| ### Policy Type: Departmental Policy Owner: Director of Laboratory Services Responsible Executive: VP Operations | Policy Contact: Dir. of Lab., ext. 1787 |
| --- | --- |

Policy Statement

Comprehensive laboratory evaluation of different body fluids provides valuable diagnostic information which assists the clinician in differentiating a variety of disease processes.  An integral component of any such fluid analysis is the   determination of the absolute numbers of erythrocytes and leukocytes within the fluid.  These red and white cell counts assist in the evaluation of hemorrhage, infection and other inflammatory disorders.

Purpose or Reason for Policy

Within most body fluids the absolute numbers of different cells are below the level necessary for accurate counting by an automated cell counter.  Therefore such cell counting requires the use of a manual method.  Manual cell counting procedures involve the microscopic visualization and enumeration of cells within an exact volume of fluid (counting chamber).

Procedures

EQUIPMENT:

- Microscope (dry lens)

- WBC/RBC pipet

- Hemacytometer

- Diff counter

- Coverslips

- Pipet suction apparatus

SUPPLIES AND REAGENTS:

- Saline (0.85%)

- Lens paper

- 70% alcohol

- 4 x 4 gauze

- Capillary tubes (plain)

SPECIMEN:

- Cerebrospinal fluid

- Peritoneal fluid

- Pleural fluid

Body Fluid Cell Count (Manual) Procedure

- Synovial fluid

- Pericardial fluid

- Peritoneal lavage fluid

*Test promptly after delivery*  Testing must be performed within 1-2 hours of collection.

If testing is unavoidably delayed, refrigerate specimen and examine ASAP.

PROCEDURE:

1.  All body fluids must be well mixed immediately before analysis begins.

a. fluids submitted in purple top tubes, invert tube 5 times.

b. fluids submitted in CSF containers, hold tube at 45 degree angle with one hand and tap against the

other hand 5 times.

c. fluids submitted in syringe, remove needle and deliver 3-4 ml into a plain red top tube, restopper

tube and invert 5 times.

2.  Clean both the counting chamber and coverglass with 70% alcohol.  Dry chamber and coverglass

with lens paper.

3.  Position clean coverglass onto clean counting chamber and place on flat counter work surface.

4.  Insert a "plain" capillary tube into the well-mixed body fluid allowing the capillary tube to fill with fluid

without air bubbles.

5.  Wipe off exterior of tube with a clean 4 x 4 gauze.

6.  Load both sides of counting chamber with fluid, ensuring no air bubbles enter under the coverglass.

7.  Carefully place counting chamber on microscope stage.

8.  Using low power objective locate ruled area of counting chamber, then switch to high dry objective.

9.  Count both the WBC's and RBC's in entire 9 square ruled chamber area simultaneously, using the diff

counter.  Record results on the Body Fluid Cell Count Worksheet (HF-38).

10.  Count the other side of the chamber in the same manner as above.  Note:  If there is >20% difference

in the counts from side 1 and 2, reload the chambers and repeat the counts to obtain more consistent

results.

11.  Complete formula on the worksheet for total WBC and RBC counts.

12.  Assess the number of both RBC's and WBC's on the counting chamber.  If the Cell density is so

great ( approximately >25 cells per square with high dry objective) that the cells are touching each

other, make a dilution using a WBC pipet and apply the diluted body fluid formula.

A.  If the cell density is such that the specimen requires dilution due to gross number of red blood

cells, but there appears to be only a few white blood cells present, lysis of red blood cells must be

obtained prior to performing the white blood cell count.  Specimens requiring dilution can be diluted

in the same manner as steps 13 - 24, substituting 3% acetic acid for saline.

B.  To prepare a clear (lysed red cells) specimen for counting without dilution, place 4 drops of mixed

specimen in a clean tube.  Rinse a pasteur pipette with glacial acetic acid ( Histology), drain

thoroughly, and draw the 4 drops of fluid into the rinsed pipette.  Allow the pipette to sit for one

minute, mix the solution in the pipette, discard the first drop, and load the hemocytometer.

HEM.35338

13.  When dilutions are required, use fresh 0.85% Blood Bank Saline as the diluent.  Check the saline

visually under the microscope for clarity and document check on worksheet.  Change if indicated.

14.  Use a WBC pipet to dilute both WBS and RBC counts.

15.  Connect the black rubber tip to the open end of the micro pipet suction apparatus. Connect the blunt

end of the pipet to the tapered end of the rubber tip.  Insert the  tapered end of the pipet to a depth of

½ inch below body fluid level.  Using the screw end of the apparatus slowly and carefully aspirate

fluid to the 0.5 mark.

16.  Remove pipet from fluid and wipe excess fluid from exterior being careful not to allow fluid level to

change from the 0.5 mark.

17.  Insert tip of pipet into the diluent (saline) and aspirate diluent to the 11 mark.

18.  Carefully wipe diluent from exterior of pipet. With you finger over the end of the pipet, remove pipet

from the apparatus.

Body Fluid Cell Count (Manual) Procedure

19.  Place pipet filled with body fluid/diluent mixture onto the pipet shaker.  Turn the shaker on.

20.  When shaker stops, remove pipet using the index finger on the top of the pipet as a control valve to

prevent fluid/diluent mixture from escaping.

HEM.35357

21.  Allow 4 drops of the mixture to be expelled onto a gauze, then carefully load both side of clean

counting chamber.

22.  Place chamber on microscope stage.  Using low power, locate ruled area of counting chamber, then

switch to high dry objective.

23.  Count both the RBC and WBC in the 4 corner squares and record results on worksheet for diluted

formula.  Count cells on other side of chamber and record results on worksheet.  Counts from both

sides should agree within 20%, if not reload hemacytometer and recount both sides.

24.  Complete formula on worksheet.

25.  If the cells are still to numerous to count diluting the fluid with a WBC pipet, a further dilution may be

required using a RBC pipet.  Follow the same technique as in steps 13 - 24.

NOTE:  The Worksheet is formulated to use a WBC pipet, when using a RBC pipet you must use a

multiplication factor of "250" instead of 25.

HEM.35319     HEM.35395     HEM35340

QUALITY CONTROL:

1.  Only certified pipets are used when diluting body fluid samples.

2.  Prepare a wright stained cytospin slide as a semi-quantitative means of checking the cell count and

the types of cells present.  Acid rinse is also available, if applicable.

3.  Only those technologists whose names are listed on the Technologist Authorization Record may

perform this procedure.

- All body fluid cytospin slides are reviewed by the pathologist for abnormalities and an interpretive

report, when indicated.

HEM.35376

NOTES AND COMMENTS:

1.  When the body fluid is partially clotted, has cell clumps or debris on the counting chamber, attach the

text "Results of cell count may be inaccurate due to partially clotted specimen" to the cell count report.

REFERENCES:

- Kjeldsberg, C.R. and Krieg, A.F.:  Cerebrospinal Fluid and Other Body Fluids in CLINICAL

DIAGNOSIS AND MANAGEMENT, 18th edition; Henry, B. editor, W.B. Saunders Company,

Philadelphia, 445-469, 1991.

- Kjeldsberg, C.R. and Knight, J. Manual Cell Count in BODY FLUIDS, 3rd ed. American Society of

Clinical Pathologist Press, Chicago, IL, 305-306, 1993.

Forms/Instruction

NA

Definitions

NA

Body Fluid Cell Count (Manual) Procedure

Signatures

| Written By: | Susie Fairfax, MT(AMT,HHS) |  | Original Date: Effective Date: | <1990 November 13, 2010 |  |
| --- | --- | --- | --- | --- | --- |
| Supersedes: | Body Fluid Cell Count (Manual) Procedure (Revised 8/14/2008) |  |  |  |  |
| CRHS | Medical Director |  |  |  | Date |
|  | Mett B. Ausley, Jr., MD |  |  |  | 11/18/10 |
| ACTION |  | Signature |  |  | Date |
| Reviewed |  | S. Fairfax, MT |  |  | 4/23/2012 |
| Reviewed |  | S. Fairfax, MT |  |  | 2/13/2013 |
| Reviewed |  | S. Fairfax< MT |  |  | 1/28/2015 |
| Reviewed |  |  |  |  |  |