## Slide 1: AmnioTest Ruptured Membranes Assessment Module
<?xml version="1.0"?><Settings><answerBulletFormat>Numeric</answerBulletFormat><answerNowAutoInsert>No</answerNowAutoInsert><answerNowStyle>Explosion</answerNowStyle><answerNowText>Answer Now</answerNowText><chartColors>Use PowerPoint Color Scheme</chartColors><chartType>Horizontal</chartType><correctAnswerIndicator>Checkmark</correctAnswerIndicator><countdownAutoInsert>No</countdownAutoInsert><countdownSeconds>10</countdownSeconds><countdownSound>TicToc.wav</countdownSound><countdownStyle>Box</countdownStyle><gridAutoInsert>No</gridAutoInsert><gridFillStyle>Answered</gridFillStyle><gridFillColor>255,255,0</gridFillColor><gridOpacity>50%</gridOpacity><gridTextStyle>Keypad #</gridTextStyle><inputSource>Response Devices</inputSource><multipleResponseDivisor># of Responses</multipleResponseDivisor><participantsLeaderBoard>5</participantsLeaderBoard><percentageDecimalPlaces>0</percentageDecimalPlaces><responseCounterAutoInsert>No</responseCounterAutoInsert><responseCounterStyle>Oval</responseCounterStyle><responseCounterDisplayValue># of Votes Received</responseCounterDisplayValue><insertObjectUsingColor>Red</insertObjectUsingColor><showResults>Yes</showResults><teamColors>Use PowerPoint Color Scheme</teamColors><teamIdentificationType>None</teamIdentificationType><teamScoringType>Voting pads only</teamScoringType><teamScoringDecimalPlaces>1</teamScoringDecimalPlaces><teamIdentificationItem></teamIdentificationItem><teamsLeaderBoard>5</teamsLeaderBoard><teamName1></teamName1><teamName2></teamName2><teamName3></teamName3><teamName4></teamName4><teamName5></teamName5><teamName6></teamName6><teamName7></teamName7><teamName8></teamName8><teamName9></teamName9><teamName10></teamName10><showControlBar>All Slides</showControlBar><defaultCorrectPointValue>0</defaultCorrectPointValue><defaultIncorrectPointValue>0</defaultIncorrectPointValue><chartColor1>187,224,227</chartColor1><chartColor2>51,51,153</chartColor2><chartColor3>0,153,153</chartColor3><chartColor4>153,204,0</chartColor4><chartColor5>128,128,128</chartColor5><chartColor6>0,0,0</chartColor6><chartColor7>0,102,204</chartColor7><chartColor8>204,204,255</chartColor8><chartColor9>255,0,0</chartColor9><chartColor10>255,255,0</chartColor10><teamColor1>187,224,227</teamColor1><teamColor2>51,51,153</teamColor2><teamColor3>0,153,153</teamColor3><teamColor4>153,204,0</teamColor4><teamColor5>128,128,128</teamColor5><teamColor6>0,0,0</teamColor6><teamColor7>0,102,204</teamColor7><teamColor8>204,204,255</teamColor8><teamColor9>255,0,0</teamColor9><teamColor10>255,255,0</teamColor10><displayAnswerImagesDuringVote>Yes</displayAnswerImagesDuringVote><displayAnswerImagesWithResponses>Yes</displayAnswerImagesWithResponses><displayAnswerTextDuringVote>Yes</displayAnswerTextDuringVote><displayAnswerTextWithResponses>Yes</displayAnswerTextWithResponses><questionSlideID></questionSlideID><controlBarState>Expanded</controlBarState><isGridColorKnownColor>True</isGridColorKnownColor><gridColorName>Yellow</gridColorName><AutoRec></AutoRec><AutoRecTimeIntrvl></AutoRecTimeIntrvl></Settings>
<?xml version="1.0"?><AllQuestions />
<?xml version="1.0"?><AllAnswers />
<?xml version="1.0"?><AllResponses />

## Slide 2: AmnioTest™
AmnioTest™ is a qualitative, pH-based, swab screening system intended as an aid in detecting rupture of amniotic membrane in pregnant women.

## Slide 3
Amniotic fluid has a neutral pH while the pH of the upper vagina is normally acidic. Rupture of the amniotic membrane can result in small volumes of amniotic fluid leaking into the upper vagina. The presence of amniotic fluid tends to elevate the pH of the upper vagina. Detection of this pH increase using a pH indicator dye has been shown to assist in determining the presence of amniotic fluid.
A pH of 6.5 or higher is suggestive of membrane rupture
Nitrazine swab qualitatively detects pH levels between 5.0-7.5
How AmnioTest™ Works

## Slide 4
Reagents
  2 buffers, ph 5.0 and 7.5
Frequency
  New shipment/lot, monthly thereafter
  Clinical symptoms do not match results
  Suspicion that product performance is compromised or when kits have not been stored according to manufacturer instructions
Quality Control (QC)

## Slide 5
Storage
  Reagents and swabs are stored at room temperature (15-30oC)
Do not use expired swabs
Do not reuse swabs
Precautions

## Slide 6: PROCEDURE Step 1 - PEEL
Remove sterile swab from individual wrapper by peeling open the AmnioTest™ sleeve.
Note: Do not allow swab tip to come in contact with anything before test is performed.

## Slide 7: PROCEDURE Step 2 - SWAB
Part the labia to expose the cervix. Insert swab so that nitrazine tip gently comes in contact with the upper vaginal tissue for about 15 seconds.
A speculum may be used during this phase of the procedure to limit swab exposure to anything that may affect pH (i.e. urine, bacteria).

## Slide 8: PROCEDURE Step 3 - READ
After carefully removing swab, immediately compare the swab tip color to the interpretation card supplied by Pro-Lab and record the corresponding pH level.
Discard swab.
Exposure to light or humidity may affect performance of the product.
**Be sure to review the technical insert provided by Pro-Lab if you are unsure about any step in the procedure, including interpretation of results.

## Slide 9
Causes of False positive:
  Antibiotic therapy
  Infections of the vagina
  Blood
  Semen
  Alkaline urine
  Other fluids that are alkaline
Causes of False negative:
  Incorrect specimen collection
  Very little amniotic fluid leaking
A negative test result does not preclude the possibility of rupture of the amniotic membrane in pregnant women. Other clinical finds should be considered when interpreting negative results.
**NOTE: Any patient result that is inconsistent with patient’s
clinical condition should be repeated.
Test Limitations

## Slide 10
Congratulations!
You have completed this module. You may proceed to the Post Exam.