P.O. Drawer Z

Hot Springs, VA  24445

(540)839-7000

A Critical Access Hospital

POLICY AND PROCEDURES

Department: Laboratory			        Effective Date:	January 1, 2016

Revised Date:

Reviewed Date:

Policy Name: Rapid Group A Strep Test

PRINCIPLE:

The QuickVue In-Line One-Step Strep A Test allows for the rapid detection of group A

streptococcal antigen directly from patient throat swab specimens. The Test is intended for use

as an aid in the diagnosis of group A streptococcal infection.

Group A Streptococci are organisms that typically cause illnesses such as tonsillitis, pharyngitis

and scarlet fever. These infections can lead to serious complications, including rheumatic fever

and acute glomerulonephritis.1  Rapid diagnosis and appropriate antibiotic therapy of group A

streptococcal infections appear to be the best means of preventing these complications. The

traditional means of detecting group A streptococcal infections involves 24-48 hour culture of

throat swab specimens or other exudates, confirming beta-hemolysis, and showing susceptibility

to bacitracin. Group A, but generally not other groups of beta-hemolytic Streptococci, is

bacitracin susceptible, which provides a presumptive diagnosis of group A streptococcal disease.2

It is also possible to identify group A Streptococci by immunological means. The cell walls of

group A Streptococci share a common, group-specific, antigenic determinant. This determinant is a carbohydrate consisting of polyrhamnose chains terminating in N-acetyl glucosamine residues.3 Antibodies against this carbohydrate antigen have been used in a variety of tests to identify the group A organism.

The QuickVue In-Line One-Step Strep A Test is a lateral-flow immunoassay utilizing an in-the-device antigen extraction. The test, containing a highly specific and sensitive antibody reactive

to the Strep A antigen, is specific to group A with no cross-reactivity from other groups of

Streptococci.

To perform the test, a throat swab specimen is collected and inserted into the Swab Chamber of

the Test Cassette. The Extraction Solutions are mixed, resulting in a green color change, and

added to the swab in the Swab Chamber in order for the antigenic component of the bacteria to

be extracted.

Extraction begins instantaneously, after which the extracted solution flows from the Swab Chamber into the test strip by capillary action. The extracted sample migrates through a label pad consisting of a pink label containing rabbit polyclonal anti-strep A antibody and a blue control label. If the extracted solution contains Strep A antigen, the antigen will bind to the antibody on the pink test label which, in turn, will bind a second rabbit polyclonal anti-Strep A antibody spotted on the membrane, resulting in the formation of a pink-to-purple Test Line. A blue Control Line will also appear next to the letter "C" on the Test Cassette indicating that the reagents were mixed and added properly, proper volume of fluid entered the Test Cassette and capillary flow occurred. A blue Control Line should always appear in a properly functioning Test Cassette. If Strep A is not present, or present at very low levels, only a blue Control Line will be visible.

REAGENTS AND MATERIALS SUPPLIED:

Individually packaged Test Cassettes (25):

Rabbit polyclonal antibody to Strep A (Test Line) and a control rabbit polyclonal antibody capable of binding the blue control label (Control Line).

Extraction Solution Bottles (25):

4M Sodium Nitrite with 0.01% Thimerosal (0.6 mL), and 0.2M Acetic Acid (0.65

mL) inside glass ampule.

Individually packaged sterile rayon-tipped swabs on solid green shafts (25).

Positive Control Swabs (+) (1):

Heat-inactivated group A Streptococcus.

Negative Control Swabs ( - ) (1):

Heat-inactivated group C Streptococcus.

Package Insert (1).

Procedure Card (1).

Extraction Kit (1):

5 Tubes and 5 Disposable Droppers for use with Proficiency Testing Samples only.

Refer to the PROFICIENCY TESTING Section for instructions for use.

WARNINGS AND PRECAUTIONS:

- For In Vitro Diagnostic Use.

- Do not use beyond the expiration date printed on the outside of the box.

- Use appropriate precautions in the collection, handling, storage, and disposal of specimens, Control Swabs and used kit contents.8 Discard used material in a proper biohazard container.

- The Test Cassette must remain sealed in the protective foil pouch until just prior to use.

- The Extraction Solution Bottle contains an acidic solution. If the solution contacts the skin or  eye, flush with large volumes of water.

- The Extraction Solution Bottle contains glass, break cautiously.

- If the Extraction Solution Bottle is missing the glass ampule, or the solution is green prior to the breaking of the ampule, discard and use another Extraction Solution Bottle.

- To obtain accurate results, package insert instructions must be followed.

KIT STORAGE AND STABILITY:

Store kit at room temperature, 59-86°F (15-30°C), out of direct sunlight. Kit contents are stable

until the expiration date printed on the outer box. Do not freeze.

SPECIMEN COLLECTION AND STORAGE:

The sterile rayon swabs supplied with this kit must be the only swabs used for specimen

collection. These special swabs have green print on the paper wrapper and green shafts.

Collect throat swab specimens by standard clinical methods. Depress the tongue with a tongue

blade or spoon. Be careful not to touch the tongue, sides or top of the mouth with the swab.

Rub the swab on the back of the throat, on the tonsils, and in any other area where there is

redness, inflammation or pus. Consult standard reference procedures such as the collection

method described by Facklam.4

Use only the rayon tipped swabs on solid green plastic shafts supplied in the kit to collect throat

specimens. Other swabs, including other rayon swabs, are incompatible with this test due to their small tip size.

It is recommended that swab specimens be processed as soon as possible after collection. Swabs can be held in a clean, dry plastic tube or sleeve up to 4 hours at room temperature (15-30°C) or 24 hours refrigerated (2-8°C) before processing. Performance with transport media has not been assessed, however the use of charcoal or agar medium is not recommended.

If a culture is desired, lightly streak the swab on a 5% sheep blood agar plate before using the

swab in the QuickVue In-Line One-Step Strep A Test. Do not perform the QuickVue In-Line

One-Step Strep A Test before streaking the swab, as the Extraction Solution will destroy the

bacteria on the swab, thereby rendering the organism incapable of successful culturing.  Alternatively, throat swab specimens can be obtained by dual swabs or by two sequential swabs for the culture procedure.

QUALITY CONTROL:

Built-In Control Features

The QuickVue In-Line One-Step Strep A Test contains built-in control features. The

manufacturer's recommendation for daily quality control is to document these controls for the

first sample tested each day. A log sheet is supplied for technician convenience at the conclusion

of this document, but its use is optional.

A control of the extraction procedure is provided by a color change from clear to green as the

extraction solutions are mixed. The color change is an indication of extraction reagent integrity

and is also an indication that the extraction procedure was correctly performed.

The two color result format provides a clear-cut readout for positive and negative results. The

appearance of a blue Control Line next to the letter "C" provides several forms of control. First,

detection components for the specimen and internal control are processed concurrently using

identical procedures; therefore, the appearance of the Control Line ensures that functional

activity of the detection component is maintained. Secondly, the appearance of the Control Line

also ensures that the foil pouch integrity has been maintained and the Test Cassette has been

stored in such a manner as not to compromise its functionality. Third, the appearance of the

Control Line indicates that proper volume of fluid entered the Test Cassette and capillary flow

occurred. This would indicate that the Test Cassette was assembled properly, by acting as a

check for all membrane interfaces and proper positioning of components. If the Control Line

does not develop within 5 minutes, the test result is invalid.

A negative background control is provided by the clearing of background color in the Result

Window and indicates that there were no immunological interfering substances in the specimen.

This area should be white to light pink within 5 minutes and not interfere with the reading of the

test result. If background color remains in the Result Window which interferes with reading the

test result, the result may be invalid. In this case, contact QUIDEL Technical Assistance.

Positive and Negative Quality Control

External controls will be used to assure that each new shipment or lot performs properly, as required by the laboratory's quality assurance plan. Positive and Negative Control

Swabs are supplied in the kit. The Positive Control Swab (+) is stored in the pink-capped tube;

the Negative Control Swab (-) is stored in the blue-capped tube. Additional control swabs may be purchased, if necessary. Quidel QuickVue Strep A Control Swab Set 00345 is available through laboratory suppliers.

To test using a Positive or Negative Control Swab, remove the Control Swab from its container

and insert it into the QuickVue In-Line One-Step Strep A Test Cassette Swab Chamber.

Continue with the assay as instructed in the TEST PROCEDURE section.

A positive and negative external control must be tested when opening a new test kit.

If controls do not perform as expected, do not report the test results. Repeat the test or contact

QUIDEL Technical Assistance.

TEST PROCEDURE:

Important :

- Gloves should be worn when handling human samples.

- Do not use the Extraction Solution if it is green prior to breaking the ampule.

For proficiency testing, follow the alternate test procedure under PROFICIENCY

TESTING SURVEY PROCEDURE.

Before Testing

- Only use the swabs provided in the kit.

- Remove the Test Cassette from foil pouch and place on a clean, dry, level surface. Using the notch at the back of the chamber as a guide, insert the swab completely into the swab chamber.

- Squeeze to crush the glass ampule inside the Extraction Solution Bottle.

Perform The Assay

- Vigorously shake the Bottle five times to mix the solutions. Solution should turn green after the ampule is broken.  Solution must be used immediately.

- Remove the cap. Quickly fill the chamber to the rim (approximately 10 drops).  Begin Timing.

- If liquid has not moved across the Result Window in 1 minute, completely remove the swab and re-insert. If liquid still does not move across, retest with a new specimen, Test Cassette and Extraction Solution Bottle.

- The Test Cassette should not be moved until the assay is complete.

READ RESULTS AT 5 MINUTES.

SOME POSITIVE RESULTS MAY BE SEEN EARLIER.

INTERPRETATION OF RESULTS:

Positive Result :

The appearance of any pink-to-purple line next to the letter “T” in the Result Window, along

with a blue Control Line next to the letter “C” means that the test is positive for group A

Streptococcus.

Negative Result :

The appearance of only the blue Control Line next to the letter “C” in the Result Window means

the test is negative. A negative QuickVue result means that the swab is a presumptive negative

for group A Streptococcus.

Invalid Result :

If the blue Control Line does not appear next to the letter "C" at 5 minutes, the test is considered

INVALID and the test result cannot be used. If this occurs, retest using a fresh swab and a new

QuickVue Test Cassette or contact Technical Assistance.

For a photographic example of test results, please see the procedure card included with the test kit.

PROFICIENCY TESTING SURVEY PROCEDURE:

The testing procedure for Proficiency Survey Swab specimens is outlined below.  This procedure must be followed to ensure accuracy with the QuickVue Test on Proficiency

Survey swab specimens because Proficiency Testing swab tips are smaller in size than the swabs

provided for use with the kit.

- Place a clean tube from the Extraction Kit in a test tube rack.

- Squeeze to crush the glass ampule inside the Extraction Solution Bottle as described in the TEST PROCEDURE Section.

- Dispense 8 DROPS from the Extraction Solution Bottle into the tube. Place the proficiency swab into the tube. Hold the bottom of the tube so the swab head is slightly compressed.

- Rotate the swab three (3) times.

- WAIT ONE (1) MINUTE

- Express all of the liquid from the swab head in the Tube by rolling the swab against the inside of the Tube and pressing slightly as it is withdrawn from the Tube. Discard the swab.

- Fill the Disposable Dropper to the fill line with the solution from the Tube and add the contents into the Test Cassette Swab Chamber.

- Read the result at 5 minutes. See INTERPRETATION OF RESULTS Section.

LIMITATIONS:

The contents of this kit are for use in the qualitative detection of group A streptococcal antigen

from throat swab specimens only. Failure to follow the test procedure and interpretation of test

results may adversely affect performance and/or produce invalid results.

Respiratory infections, including pharyngitis, can be caused by Streptococcus from serogroups

other than group A as well as other pathogens. The QuickVue In-line One-Step Strep A Test will

not differentiate asymptomatic carriers of group A Streptococcus from those exhibiting

streptococcal infection.6  In rare cases, test specimens heavily colonized with Staphylococcus

aureus (>1010) can yield false positive results.

Test results must always be evaluated with other data available to the physician. A negative test

result might occur if the level of extracted antigen in a sample is below the sensitivity of the test.

Additional follow-up testing using the culture method is recommended if the QuickVue test

result is negative.

EXPECTED VALUES:

Group A Streptococci are responsible for about 19% of all upper respiratory tract infections, but

the incidence varies by clinical setting.7  Streptococcal pharyngitis is seasonal in nature with the

highest prevalence found during winter and early spring. The highest incidence of this disease is

found in crowded populations such as military bases and in school-aged children, and is evenly

distributed between males and females.8

PERFORMANCE CHARACTERISTICS:

Clinical Sensitivity and Specificity

The QuickVue In-Line One-Step Strep A Test will yield positive test results with specimens

containing at least 5x105 group A Streptococci organisms per test.

A multi-center evaluation of the QuickVue test was conducted to determine the clinical

performance of the test relative to standard culture techniques. A total of 537 throat swab

specimens were collected from patients presenting with pharyngitis. Prior to performance of the

QuickVue test, each swab specimen was inoculated onto a sheep blood agar plate containing a

bacitracin disk and incubated at 37°C for 48 hours for culture evaluation. All cultures were

confirmed for the presence of group A strep using commercial latex agglutination assays.

Swabs were either tested in the QuickVue test immediately upon collection at the field site (fresh

specimens) or frozen and shipped overnight to QUIDEL. Testing with the QuickVue test was

performed by trained technicians and by users in the field with various levels of work experience

and educational backgrounds.

Of the 537 total specimens, 301 fresh specimens were tested by field users while the other 236

frozen specimens were tested by trained technicians at QUIDEL. Ten (10) additional specimens

tested resulted in uninterpretable results and were eliminated from the analysis.

In the field study, 240 specimens were found to be negative by SBA culture and 225 were also

negative by the QuickVue test. Similarly, 61 specimens were found to be positive by SBA

culture and 53 were also positive by the QuickVue test. Based on this data, specificity was 94%

and sensitivity was 87% for the QuickVue test, 95% confidence intervals were calculated to be

91-97% and 78-95% for specificity and sensitivity, respectively. Overall agreement between

SBA culture and QuickVue was 92% (278/301).

In a separate study conducted at QUIDEL, 136 specimens were found to be negative by SBA

culture and 135 were also negative by the QuickVue test; similarly, 100 specimens were found

to be positive by SBA culture and 92 were also positive by the QuickVue test. Based on this

data, specificity was 99% and sensitivity was 92% for the QuickVue test; 95% confidence

intervals were calculated to be 96-100% and 87-97% for specificity and sensitivity, respectively.

Overall agreement between SBA culture and QuickVue was 96% (227/236).

| SPECIMEN MIX IN POPULATION |  |  |  |  |
| --- | --- | --- | --- | --- |
| Culture Classification | Field Users |  | Trained Technicians |  |
|  | Correct Calls | % Correct | Correct Calls | % Correct |
| Negative | 25/240 2 | 94 | 135/136 | 99 |
| 1+ | 3/9 | 33 | 2/5 | 40 |
| 2+ | 7/9 | 78 | 6/6 | 100 |
| 3+ | 18/18 | 100 | 39/39 | 100 |
| 4+ | 25/25 | 100 | 45/50 | 90 |
| Overall Agreement to Culture | 278/301 | 92% | 227/236 | 96% |

Physician’s Office Laboratory (POL) Studies

An evaluation of the QuickVue In-Line One-Step Strep A Test was conducted at four Physicians'

Offices using a panel of coded specimens. Testing was performed by physician's office

personnel with diverse educational backgrounds and work experience at different locations. The

proficiency panel contained negative, low positive, moderate positive and high positive

specimens. Each specimen level was tested in replicates of five at each site over a period of three

days.

The results obtained at each site ranged from 88 to 100% agreement with the expected results.

No significant differences were observed within run (five replicates), between runs (three

different assay days), or between sites (four POL sites).

Cross-Reactivity

Group C Streptococcus, group G Streptococcus, S. aureus, N. subflava, H. influenza, C. albicans, N. meningitidis, N. gonorrhoea, B. catarrhalis, E. faecalis, S. pneumoniae, and S. mutans were tested in the QuickVue In-Line One-Step Strep A Test at levels exceeding 107/test and did not affect the expected test results.

COMMENTS AND TECHNICAL ASSISTANCE:

If you have any questions regarding the use of this product, please call QUIDEL's Technical

Assistance number, 800-874-1517 (U.S. toll-free) or 858-552-1100, Monday through Friday,

between 7:00 a.m. and 5:00 p.m., Pacific Time. If outside the United States, contact your local

QUIDEL office or distributor.

REFERENCES:

1. Rammelkamp CH, Jr. Principles of Internal Medicine, 8th Edition, 1977, pp. 814-820.

2. Youmans GP, Paterson PY, and Sommers HM. The Biological and Clinical Basis of

Infectious Disease, 1975, pp.172-185.

3. Youmans GP, Paterson PY, and Sommers HM. Ibid., pp. 172-181.

4. Recommendations for the Prevention of HIV Transmission in Health Care Settings,

Morbidity and Mortality Weekly Report, Centers for Disease Control, August 21, 1987.

5. Facklam, RR. and Washington II, JA. Manual of Clinical Microbiology, Balows, A. ed. 5th

Edition, 1991, 29:241.

6. Rammelkamp CH. Jr. Ibid., p. 814.

7. Lauer BA, Reller LB, and Mirrett S. Journal of Clinical Microbiology. 17:338-340, 1983.

- Wannamaker LW. New England Journal of Medicine

Procedure prepared by: Tara Craft 12/16/2015

Approved by: _____________________________ Date: _____________