Laboratory Name

Test Name: a-Amylase EPS ver.2

System information

For cobas c 311/501 analyzers:

AMYL2: ACN 570

SAMY2: ACN 566 (STAT, reaction time: 7)

Intended use

In vitro test for the quantitative determination of a-amylase in human serum, plasma and urine on Roche/Hitachi cobas c systems.

Summary1,2,3,4,5,6,7,8,9

The a-amylases (1,4-a-D-glucanohydrolases, EC 3.2.1.1) catalyze the hydrolytic degradation of polymeric carbohydrates such as amylose, amylopectin and glycogen by cleaving 1,4-a-glucosidic bonds. In polysaccharides and oligosaccharides, several glycosidic bonds are hydrolyzed simultaneously. Maltotriose, the smallest such unit, is converted into maltose and glucose, albeit very slowly. Two types of a-amylases can be distinguished, the pancreatic type (P-type) and the salivary type (S-type). Whereas the P-type can be attributed almost exclusively to the pancreas and is therefore organ-specific, the S-type can originate from a number of sites. As well as appearing in the salivary glands it can also be found in tears, sweat, human milk, amniotic fluid, the lungs, testes and the epithelium of the fallopian tube.

Because of the sparsity of specific clinical symptoms of pancreatic diseases, a-amylase determinations are of considerable importance in pancreatic diagnostics. They are mainly used in the diagnosis and monitoring of acute pancreatitis. Hyperamylasemia does not, however, only occur with acute pancreatitis or in the inflammatory phase of chronic pancreatitis, but also in renal failure (reduced glomerular filtration), tumors of the lungs or ovaries, pulmonary inflammation, diseases of the salivary gland, diabetic ketoacidosis, cerebral trauma, surgical interventions or in the case of macroamylasemia. To confirm pancreatic specificity, it is recommended that an additional pancreas-specific enzyme - lipase or pancreatic-a-amylase - also be determined.

Numerous methods have been described for the determination of a-amylase. These either determine the decrease in the amount of substrate viscometrically, turbidimetrically, nephelometrically and amyloclastically or measure the formation of degradation products saccharogenically or kinetically with the aid of enzyme-catalyzed subsequent reactions. The kinetic method described here is based on the well-proven cleavage of 4,6-ethylidene-(G7)-1,4-nitrophenyl-(G1)- a,D-maltoheptaoside (Ethylidene Protected Substrate = EPS) by a-amylase and subsequent hydrolysis of all the degradation products to p-nitrophenol with the aid of a-glucosidase (100 % chromophore liberation). The results of this method correlate with those obtained by HPLC. This assay follows the recommendation of the IFCC, but was optimized for performance and stability.

Test principle10,11

Enzymatic colorimetric assay acc. to IFCC.

Defined oligosaccharides such as 4,6-ethylidene-(G7) p-nitrophenyl-(G1)-a-D-maltoheptaoside 
(ethylidene-G7PNP) are cleaved under the catalytic action of a-amylases. The G2PNP, G3PNP and G4PNP fragments so formed are completely hydrolyzed to p-nitrophenol and glucose by a-glucosidase.

Simplified reaction scheme:

|  |  | a-amylase |  |
| --- | --- | --- | --- |
| 5 ethylidene-G7PNPa) + 5 H2O |  |  |  |
| 2 ethylidene-G5 + 2 G2PNP + 2 ethylidene-G4 + 2 G3PNP + ethylidene-G3 + G4PNP |  |  |  |
|  |  |  | a-glucosidase |
| 2 G2PNP + 2 G3PNP + G4PNP + 14 H2O |  |  | 5 PNP + 14 Gb) |

a) PNP  p-nitrophenol

b) G  Glucose

The color intensity of the p-nitrophenol formed is directly proportional to the a-amylase activity. It is determined by measuring the increase in absorbance.

Reagents - working solutions

| R1 | HEPES: 52.4 mmol/L; sodium chloride: 87 mmol/L; calcium chloride: 0.08 mmol/L; magnesium chloride: 12.6 mmol/L; a-glucosidase (microbial): ³ 66.8 µkat/L; pH 7.0 (37 °C); preservatives; stabilizers |
| --- | --- |
| R2 | HEPES: 52.4 mmol/L; ethylidene-G7-PNP: 22 mmol/L; pH 7.0 (37 °C); preservatives; stabilizers |

R1 is in position B and R2 is in position C.

Precautions and warnings

For in vitro diagnostic use.

Exercise the normal precautions required for handling all laboratory reagents.

Disposal of all waste material should be in accordance with local guidelines.

Safety data sheet available for professional user on request.

Reagent handling

Ready for use

Storage and stability

| AMYL2 |  |
| --- | --- |
| Shelf life at 2-8 °C: | See expiration date on cobas c pack label. |
| On-board in use and refrigerated on the analyzer: | 12 weeks |
| Diluent NaCl 9 % |  |
| Shelf life at 2-8 °C: | See expiration date on cobas c pack label. |
| On-board in use and refrigerated on the analyzer: | 12 weeks |

Specimen collection and preparation9,12

For specimen collection and preparation only use suitable tubes or collection containers.

Only the specimens listed below were tested and found acceptable.

Serum

Plasma: Li-heparin plasma.

The sample types listed were tested with a selection of sample collection tubes that were commercially available at the time of testing, i.e. not all available tubes of all manufacturers were tested. Sample collection systems from various manufacturers may contain differing materials which could affect the test results in some cases. When processing samples in primary tubes (sample collection systems), follow the instructions of the tube manufacturer.

Centrifuge samples containing precipitates before performing the assay.

Urine: Collect urine without additives. a-Amylase is unstable in acid urine. Assay promptly or adjust pH to alkaline range (just above pH 7) before storage.13

| Stability in serum or plasma:13 | 7 days at 15-25 °C |
| --- | --- |
|  | 1 month at 2-8 °C |
| Stability in urine:14 | 2 days at 15-25 °C |
|  | 10 days at 2-8 °C |

Materials provided

See “Reagents - working solutions” section for reagents.

Materials required (but not provided)

- See “Order information” section

- General laboratory equipment

In addition, other suitable control material can be used.

Assay

For optimum performance of the assay follow the directions given in this document for the analyzer concerned. Refer to the appropriate operator's manual for analyzer-specific assay instructions.

The performance of applications not validated by Roche is not warranted and must be defined by the user.

Application for serum, plasma and urine

| cobas c 501 test definition |  |  |  |
| --- | --- | --- | --- |
| Assay type | Rate A |  |  |
| Reaction time / Assay points | 10 / 30-47 (STAT 7 / 30-47) |  |  |
| Wavelength (sub/main) | 700/415 nm |  |  |
| Reaction direction | Increase |  |  |
| Unit | U/L (µkat/L) |  |  |
| Reagent pipetting |  | Diluent (H2O) |  |
| R1 | 100 µL | – |  |
| R2 | 20 µL | – |  |
| Sample volumes | Sample | Sample dilution |  |
|  |  | Sample | Diluent (NaCl) |
| Normal | 4 µL | – | – |
| Decreased | 8 µL | 15 µL | 135 µL |
| Increased | 4 µL | – | – |

Calibration

| Calibrators | S1: H2O S2: C.f.a.s. |
| --- | --- |
| Calibration mode | Linear |
| Calibration frequency | 2-point calibration - after reagent lot change - as required following quality control procedures |

Traceability: This method has been standardized against Roche system reagent using calibrated pipettes together with a manual photometer providing absolute values and substrate-specific absorptivity, e.

Quality control

At least once daily run solutions at two levels of a quality control material with known concentrations.

Refer to Brown Clinic Quality Control Requirements, Rules and Reviews Policy

Refer to Brown Clinic Quality Control Specialty and Subspecialty Policy

Calculation

Roche/Hitachi cobas c systems automatically calculate the analyte concentration of each sample.

Conversion factor: U/L x 0.0167 = µkat/L

Limitations - interference

A slight change in the yellow coloration of solution 2 does not interfere with the performance of the test.

Do not pipette by mouth, and ensure that the reagent does not come into contact with the skin. Saliva and sweat contain a-amylase!

Criterion: Recovery within ± 10 % of initial value at an amylase activity of 100 U/L (1.67 µkat/L).

Serum/plasma

Icterus:15 No significant interference up to an I index of 60 for conjugated and unconjugated bilirubin (approximate conjugated and unconjugated bilirubin concentration: 1026 µmol/L or 60 mg/dL).

Hemolysis:15 No significant interference up to an H index of 500 (approximate hemoglobin concentration: 310 µmol/L or 500 mg/dL).

Lipemia (Intralipid):15 No significant interference up to an L index of 1500. There is poor correlation between the L index (corresponds to turbidity) and triglycerides concentration.

In rare cases, samples with a combination of elevated turbidity (L-index) and high Amylase activity may cause a >React or >Abs. flags.

Highly turbid and grossly lipemic samples may cause Abs. flags.

Anticoagulants: Interference was found with citrate, fluoride, and EDTA.12

Glucose: No interference from glucose up to 111 mmol/L (2000 mg/dL). Approximately 10 % higher recovery was found at glucose concentrations of 250 mmol/L (4500 mg/dL).

Ascorbic acid: No interference from ascorbic acid up to 5.68 mmol/L (100 mg/dL).

Drugs: No interference was found at therapeutic concentrations using common drug panels.16,17

Exception: Icodextrin-based drugs may lead to decreased amylase results.18

In very rare cases, gammopathy, in particular type IgM (Waldenström’s macroglobulinemia), may cause unreliable results.19

ACTION REQUIRED

Special Wash Programming: The use of special wash steps is mandatory when certain test combinations are run together on Roche/Hitachi cobas c systems. The latest version of the carry-over evasion list can be found with the NaOHD/SMS/Multiclean/SCCS or the NaOHD/SMS/SmpCln1+2/SCCS Method Sheets. For further instructions refer to the operator's manual. cobas c 502 analyzer: All special wash programming necessary for avoiding carry-over is available via the cobas link, manual input is not required.

Where required, special wash/carry-over evasion programming must be implemented prior to reporting results with this test.

Limits and ranges

Measuring range

Serum/plasma

3-1500 U/L (0.05-25.0 µkat/L)

Determine samples having higher activities via the rerun function. Dilution of samples via the rerun function is a 1:5 dilution. Results from samples diluted using the rerun function are automatically multiplied by a factor of 5.

Lower limits of measurement

Lower detection limit of the test

3 U/L (0.05 µkat/L)

The lower detection limit represents the lowest measurable analyte level that can be distinguished from zero. It is calculated as the value lying three standard deviations above that of the lowest standard (standard 
1 + 3 SD, repeatability, n = 21).

Expected values9

| Serum/plasma | Men/Women | 0.47-1.67 µkat/L | 28-100 U/L |
| --- | --- | --- | --- |

Specific performance data

For Known Interfering Substances section refer to package insert.

For Known Non-Interfering Substance refer to package insert.

For Additional Technical Information refer to package insert.

References

- Greiling H, Gressner AM, eds. Lehrbuch der Klinischen Chemie und Pathobiochemie, 3rd ed. Stuttgart/New York: Schattauer Verlag 1995.

- Keller H, ed. Klinisch-chemische Labordiagnostik für die Praxis, 2nd ed. Stuttgart/New York: Georg Thieme Verlag 1991;354-361.

- Salt WB II, Schenker S. Amylase - its clinical significance: a review of the literature [Review]. Medicine 1976;55:269-281.

- Steinberg WM, Goldstein SS, Davies ND, et al. Diagnostic assays in acute pancreatitis [Review]. Ann Intern Med 1985;102:576-580.

- Tietz NW, Huang WY, Rauh DF, et al. Laboratory tests in the differential diagnosis of hyperamylasemia. Clin Chem 1986;32(2):301-307.

- Junge W, Troge B, Klein G, et al. Evaluation of a New Assay for Pancreatic Amylase: Performance Characteristics and Estimation of Reference Intervals. Clin Biochem 1989;22:109-114.

- Rauscher E, von Bülow S, Hägele EO, et al. Ethylidene protected substrate for the assay of human 
a-amylase. Fresenius Z Analyt Chem 1986;324:304-305.

- Kruse-Jarres JD, Hafkenscheid JCM, Hohenwallner W, et al. Evaluation of a New a-Amylase Assay Using 4,6-Ethylidene-(G7)-1-4-nitrophenyl-(G1)-a-D-maltoheptaoside as Substrate. J Clin Chem Clin Biochem 1989;27:103-113.

- Junge W, Wortmann W, Wilke B, et al. Development and evaluation of assays for the determination of total and pancreatic amylase at 37 °C according to the principle recommended by the IFCC. Clin Biochem 2001; 34:607-615. Erratum Clin Biochem 2003;36:161.

- Lorentz K. Approved recommendation on IFCC methods for the measurement of catalytic concentration of enzymes. Part 9. IFCC Method for a-Amylase. (1,4-a-D-Glucan 4-Glucanohydrolase, EC 3.2.1.1). Clin Chem Lab Med 1998;36(3):185-203.

- Kurrle-Weitenhiller A, Hölzel W, Engel D, et al. Method for the determination of total and pancreatic 
a-amylase based on 100 % cleavage of the protected substrate ethylidene-4-nitrophenyl-maltoheptaoside. Clin Chem 1996;42(S6):98.

- Young DS. Effects of Preclinical Variables on Clinical Laboratory Tests. AACC Press 1997, 2nd edition 1997.

- Tietz NW, ed. Clinical Guide to Laboratory Tests, 3rd ed. Philadelphia PA: WB Saunders Company 1995;46-51.

- Hohenwallner W, Hägele EO, Scholer A, et al. Bestimmung von alpha-Amylase mit 
p-Nitrophenylmaltoheptaosid als Substrat. Ber Öster Ges Klin Chem 1983;6:101-112.

- Glick MR, Ryder KW, Jackson SA. Graphical Comparisons of Interferences in Clinical Chemistry Instrumentation. Clin Chem 1986;32:470-475.

- Breuer J. Report on the Symposium “Drug effects in Clinical Chemistry Methods”. Eur J Clin Chem Clin Biochem 1996;34:385-386.

- Sonntag O, Scholer A. Drug interference in clinical chemistry: recommendation of drugs and their concentrations to be used in drug interference studies. Ann Clin Biochem 2001;38:376-385.

- Gokal R, Moberly J, Lindholm B, et al. Metabolic and laboratory effects of icodextrin. Kidney Int 2002;62(81):62-71.

- Bakker AJ, Mücke M. Gammopathy interference in clinical chemistry assays: mechanisms, detection and prevention. Clin Chem Lab Med 2007;45(9):1240-1243.

- Bablok W, Passing H, Bender R, et al. A general regression procedure for method transformation. Application of linear regression procedures for method comparison studies in clinical chemistry, Part III. J Clin Chem Clin Biochem 1988 Nov;26(11);783-790.

Alternative method

Refer to Brown Clinic Back-up Testing Policy

Order information

|  |  |  | Analyzer(s) on which cobas c pack(s) can be used |
| --- | --- | --- | --- |
| 03183742 122 | a-Amylase EPS ver.2 (300 tests) | System-ID 07 6609 7 | Roche/Hitachi cobas c 311, cobas c 501/502 |
| 10759350 190 | Calibrator f.a.s. (12 x 3 mL) | Code 401 |  |
| 10759350 360 | Calibrator f.a.s. (12 x 3 mL, for USA) | Code 401 |  |
| 12149435 122 | Precinorm U plus (10 x 3 mL) | Code 300 |  |
| 12149435 160 | Precinorm U plus (10 x 3 mL, for USA) | Code 300 |  |
| 12149443 122 | Precipath U plus (10 x 3 mL) | Code 301 |  |
| 12149443 160 | Precipath U plus (10 x 3 mL, for USA) | Code 301 |  |
| 10171743 122 | Precinorm U (20 x 5 mL) | Code 300 |  |
| 10171735 122 | Precinorm U (4 x 5 mL) | Code 300 |  |
| 10171778 122 | Precipath U (20 x 5 mL) | Code 301 |  |
| 10171760 122 | Precipath U (4 x 5 mL) | Code 301 |  |
| 04489357 190 | Diluent NaCl 9 % (50 mL) | System-ID 07 6869 3 |  |

Source document

Reagent Name: AMYL2

Package Insert Version: 2014-09, V9.0 English

Effective date

| Effective date for this procedure: |  |
| --- | --- |

Author

Source documentation compiled by Roche Diagnostics

Revised by:	Heather J Hall, MBA, MT(ASCP), CG(ASCP)cm		Date: 4/9/2018

Approved by: 								Date:

REVIEW – REVISION SUMMARY DOCUMENTATION

Date:	By:			Revision Summary: