Laboratory Name

Test Name: Cholesterol Gen.2

- System information

For cobas c 311/501 analyzers:

CHO2I: ACN 798: ID/MS Standardization

CHO2A: ACN 433: Abell/Kendall Standardization

- Intended use

In vitro test for the quantitative determination of cholesterol in human serum and plasma on Roche/Hitachi cobas c systems.

- Summary

Cholesterol is a steroid with a secondary hydroxyl group in the C3 position. It is synthesized in many types of tissue, but particularly in the liver and intestinal wall. Approximately three quarters of cholesterol is newly synthesized and a quarter originates from dietary intake. Cholesterol assays are used for screening for atherosclerotic risk and in the diagnosis and treatment of disorders involving elevated cholesterol levels as well as lipid and lipoprotein metabolic disorders.

Cholesterol analysis was first reported by Liebermann in 1885 followed by Burchard in 1889. In the Liebermann-Burchard reaction, cholesterol forms a blue-green dye from polymeric unsaturated carbohydrates in an acetic acid/acetic anhydride/concentrated sulfuric acid medium. The Abell and Kendall method is specific for cholesterol, but is technically complex and requires the use of corrosive reagents. In 1974, Roeschlau and Allain described the first fully enzymatic method. This method is based on the determination of Δ4-cholestenone after enzymatic cleavage of the cholesterol ester by cholesterol esterase, conversion of cholesterol by cholesterol oxidase, and subsequent measurement by the Trinder reaction of the hydrogen peroxide formed. Optimization of ester cleavage (> 99.5 %) allows standardization using primary and secondary standards and a direct comparison with the CDC and NIST reference methods.1,2,3,4,5,6,7,8,9 Nonfasting sample results may be slightly lower than fasting results.10,11,12

The Roche cholesterol assay meets the 1992 National Institutes of Health (NIH) goal of less than or equal to 3 % for both precision and bias.12

The assay is optionally standardized against Abell/Kendall and isotope dilution/mass spectrometry. The performance claims and data presented here are independent of the standardization.

- Test principle

Enzymatic, colorimetric method.

Cholesterol esters are cleaved by the action of cholesterol esterase to yield free cholesterol and fatty acids. Cholesterol oxidase then catalyzes the oxidation of cholesterol to cholest-4-en-3-one and hydrogen peroxide. In the presence of peroxidase, the hydrogen peroxide formed effects the oxidative coupling of phenol and 4-aminophenazone to form a red quinone-imine dye.

|  | CE |  |
| --- | --- | --- |
| Cholesterol esters + H2O |  | cholesterol + RCOOH |
|  | CHOD |  |
| Cholesterol + O2 |  | cholest-4-en-3-one + H2O2 |
|  | POD |  |
| 2 H2O2 + 4-AAP + phenol |  | quinone-imine dye + 4 H2O |

The color intensity of the dye formed is directly proportional to the cholesterol concentration. It is determined by measuring the increase in absorbance.

- Reagents – working solutions

| R1 | PIPES buffer: 225 mmol/L, pH 6.8; Mg2+: 10 mmol/L; sodium cholate: 0.6 mmol/L; 4-aminophenazone: ≥ 0.45 mmol/L; phenol: ≥ 12.6 mmol/L; fatty alcohol polyglycol ether: 3 %; cholesterol esterase (Pseudomonas spec.): ≥ 25 µkat/L (≥ 1.5 U/mL); cholesterol oxidase (E. coli): ≥ 7.5 µkat/L (≥ 0.45 U/mL); peroxidase (horseradish): ≥ 12.5 µkat/L (≥ 0.75 U/mL); stabilizers; preservative |
| --- | --- |

R1 is in position B.

- Precautions and warnings

For in vitro diagnostic use.
Exercise the normal precautions required for handling all laboratory reagents. Disposal of all waste material should be in accordance with local guidelines.
Safety data sheet available for professional user on request.

For USA: For prescription use only.

This kit contains components classified as follows in accordance with the Regulation (EC) No. 1272/2008:

Warning

| H319 | Causes serious eye irritation. |
| --- | --- |

Prevention:

| P264 | Wash skin thoroughly after handling. |
| --- | --- |

| P280 | Wear eye protection/ face protection. |
| --- | --- |

Response:

| P305 + P351 + P338 | IF IN EYES: Rinse cautiously with water for several minutes. Remove contact lenses, if present and easy to do. Continue rinsing. |
| --- | --- |

| P337 + P313 | If eye irritation persists: Get medical advice/attention. |
| --- | --- |

Product safety labeling primarily follows EU GHS guidance.

Contact phone: all countries: +49-621-7590, USA: 1-800-428-2336

- Reagent handling

Ready for use

- Storage and stability

| CHOL2 |  |
| --- | --- |
| Shelf life at 2-8 °C: | See expiration date on cobas c pack label. |
| On-board in use and refrigerated on the analyzer: | 4 weeks |
| Diluent NaCl 9 % |  |
| Shelf life at 2-8 °C: | See expiration date on cobas c pack label. |
| On-board in use and refrigerated on the analyzer: | 12 weeks |

- Specimen collection and preparation

For specimen collection and preparation only use suitable tubes or collection containers.

Only the specimens listed below were tested and found acceptable.
Serum.
Plasma: Li-heparin and K2-EDTA plasma

Do not use citrate, oxalate or fluoride.13

Fasting and nonfasting samples can be used.11

The sample types listed were tested with a selection of sample collection tubes that were commercially available at the time of testing, i.e. not all available tubes of all manufacturers were tested. Sample collection systems from various manufacturers may contain differing materials which could affect the test results in some cases. When processing samples in primary tubes (sample collection systems), follow the instructions of the tube manufacturer.

Centrifuge samples containing precipitates before performing the assay.

| Stability:14,15 | 7 days at 15-25 °C |
| --- | --- |
|  | 7 days at 2-8 °C |
|  | 3 months at (-15)-(-25) °C |

- Materials provided

See “Reagents – working solutions” section for reagents.

- Materials required (but not provided)

| ▪ | See “Order information” section |
| --- | --- |
| ▪ | General laboratory equipment |

In addition, other suitable control material can be used.

- Assay

For optimum performance of the assay follow the directions given in this document for the analyzer concerned. Refer to the appropriate operator’s manual for analyzer-specific assay instructions.

The performance of applications not validated by Roche is not warranted and must be defined by the user.

- Application for serum and plasma

| cobas c 501 test definition |  |  |  |
| --- | --- | --- | --- |
| Assay type | 1-Point |  |  |
| Reaction time / Assay points | 10 / 70 |  |  |
| Wavelength (sub/main) | 700/505 nm |  |  |
| Reaction direction | Increase |  |  |
| Units | mmol/L (mg/dL, g/L) |  |  |
| Reagent pipetting |  | Diluent (H2O) |  |
| R1 | 47 µL | 93 µL |  |
| Sample volumes | Sample | Sample dilution |  |
|  |  | Sample | Diluent (NaCl) |
| Normal | 2 µL | – | – |
| Decreased | 2 µL | 15 µL | 135 µL |
| Increased | 2 µL | – | – |
| Sample volumes | Sample | Sample dilution |  |
|  |  | Sample | Diluent (NaCl) |
| Normal | 2 µL | – | – |
| Decreased | 2 µL | 15 µL | 135 µL |
| Increased | 4 µL | – | – |

- Calibration

| Calibrators | S1: H2O S2: C.f.a.s. |
| --- | --- |
| Calibration mode | Linear |
| Calibration frequency | 2-point calibration • after reagent lot change • as required following quality control procedures |

Traceability: This method has been standardized according to Abell/Kendall12 and also by isotope dilution/mass spectrometry.16

- Quality control

At least once daily run solutions at two levels of a quality control material with known concentrations. Refer to Brown Clinic Quality Control Requirements, Rules and Reviews Policy Refer to Brown Clinic Quatliy Control Specialty and Subspecialty Policy

- Calculation

Roche/Hitachi cobas c systems automatically calculate the analyte concentration of each sample.

| Conversion factors: | mmol/L x 38.66 = mg/dL |
| --- | --- |
|  | mmol/L x 0.3866 = g/L |
|  | mg/dL x 0.0259 = mmol/L |

- Limitations – interference

Criterion: Recovery within ± 10 % of initial values at a cholesterol concentration of 5.2 mmol/L (200 mg/dL).

Icterus:17 No significant interference up to an I index of 16 for conjugated bilirubin and 14 for unconjugated bilirubin (approximate conjugated bilirubin concentration 274 µmol/L or 16 mg/dL; approximate unconjugated bilirubin concentration 239 µmol/L or 14 mg/dL).

Hemolysis:17 No significant interference up to an H index of 700 (approximate hemoglobin concentration: 435 µmol/L or 700 mg/dL).

Lipemia (Intralipid):17 No significant interference up to an L index of 2000. There is poor correlation between the L index (corresponds to turbidity) and triglycerides concentration.

Drugs:No interference was found at therapeutic concentrations using common drug panels.18,19

Acetaminophen intoxications are frequently treated with N-Acetylcysteine. N-Acetylcysteine at the therapeutic concentration when used as an antidote and the Acetaminophen metabolite N-acetyl-p-benzoquinone imine (NAPQI) independently may cause falsely low results.

Venipuncture should be performed prior to the administration of Metamizole. Venipuncture immediately after or during the administration of Metamizole may lead to falsely low results.

In very rare cases, gammopathy, in particular type IgM (Waldenström’s macroglobulinemia), may cause unreliable results.20

For diagnostic purposes, the results should always be assessed in conjunction with the patient’s medical history, clinical examination and other findings.

ACTION REQUIRED
Special Wash Programming: The use of special wash steps is mandatory when certain test combinations are run together on Roche/Hitachi cobas c systems. The latest version of the carry-over evasion list can be found with the NaOHD-SMS-SmpCln1+2-SCCS

Method Sheets.cobas c 502 analyzer: All special wash programming necessary for avoiding carry-over is available via the cobas link, manual input is not required.

Where required, special wash/carry-over evasion programming must be implemented prior to reporting results with this test.

- Limits and ranges

Measuring range

0.1-20.7 mmol/L (3.86-800 mg/dL)

Determine samples having higher concentrations via the rerun function. Dilution of samples via the rerun function is a 1:10 dilution. Results from samples diluted using the rerun function are automatically multiplied by a factor of 10.

Lower limits of measurement

Lower detection limit of the test

0.1 mmol/L (3.86 mg/dL)

The lower detection limit represents the lowest measurable analyte level that can be distinguished from zero. It is calculated as the value lying 3 standard deviations above that of the lowest standard (standard 1 + 3 SD, repeatability, n = 21).

- Expected values

Clinical interpretation according to the recommendations of the European Atherosclerosis Society:21

|  | mmol/L | mg/dL | Lipid metabolic disorder |
| --- | --- | --- | --- |
| Cholesterol | < 5.2 | (< 200) | No |
| Triglycerides | < 2.3 | (< 200) | No |
| Cholesterol | 5.2-7.8 | (200-300) | Yes, if HDL-cholesterol < 0.9 mmol/L (< 35 mg/dL) |
| Cholesterol | > 7.8 | (> 300) | Yes |
| Triglycerides | > 2.3 | (> 200) | Yes |

Recommendations of the NCEP Adult Treatment Panel for the following risk-cutoff thresholds for the US American population:22

| Desirable cholesterol level | < 5.17 mmol/L | (< 200 mg/dL) |
| --- | --- | --- |
| Borderline high cholesterol | 5.17-6.18 mmol/L | (200-239 mg/dL) |
| High cholesterol | ≥ 6.21 mmol/L | (≥ 240 mg/dL) |

Each laboratory should investigate the transferability of the expected values to its own patient population and if necessary determine its own reference ranges.

- Specific performance data

For Known Interfering Substances section refer to package insert.

For Known Non-Interfering Substance refer to package insert.

For Additional Technical Information refer to package insert.

- References

| 1 | Greiling H, Gressner AM, eds. Lehrbuch der Klinischen Chemie und Pathobiochemie, 3rd ed. Stuttgart/New York: Schattauer Verlag 1995. |
| --- | --- |
| 2 | Liebermann NC. Über das Oxychinoterpen. Ber Dtsch chem Ges 1885;18:1803. |
| 3 | Burchard H. Beiträge zur Kenntnis der Cholesterine. Dissertation, Rostock 1889. |
| 4 | Abell LL, Levy BB, Kendall FE. Cholesterol in serum. In: Seligon D (ed.). Standard Methods of Clinical Chemistry. Vol 2. Academic Press, New York 1958;26-33. |
| 5 | Allain CC, Poon LS, Chan CS, et al. Enzymatic determination of total serum cholesterol. Clin Chem 1974;20(4):470-475. |
| 6 | Roeschlau P, Bernt E, Gruber W. Enzymatic determination of total cholesterol in serum. Z Klin Chem Klin Biochem 1974;12(5):226. |
| 7 | Trinder P. Determination of glucose in blood using glucose oxidase with an alternative oxygen acceptor. Ann Clin Biochem 1969;6:24-27. |
| 8 | Siedel J, Hägele EO, Ziegenhorn J, et al. Reagent for the enzymatic determination of serum total cholesterol with improved lipolytic efficiency. Clin Chem 1983;29:1075-1080. |
| 9 | Wiebe DA, Bernert JT. Influence of incomplete cholesteryl ester hydrolysis on enzymatic measurements of cholesterol. Clin Chem 1984;30:352-356. |
| 10 | Cohn JS, McNamara JR, Schaefer EJ. Lipoprotein Cholesterol Concentrations in the Plasma of Human Subjects as Measured in the Fed and Fasted States. Clin Chem 1988;34:2456-2459. |
| 11 | Pisani T, Gebski CP, Leary ET, et al. Accurate Direct Determination of Low-density Lipoprotein Cholesterol Using an Immunoseparation Reagent and Enzymatic Cholesterol Assay. Arch Pathol Lab Med 1995 Dec;119(12):1127-1135. |
| 12 | Recommendations for Improving Cholesterol Measurement: A Report from the Laboratory Standardization Panel of the National Cholesterol Education Program. NIH Publication No. 90-2964 1990. |
| 13 | Nader R, Dufour DR, Cooper GR. Preanalytical Variation in Lipid, Lipoprotein, and Apolipoprotein Testing. In: Rifai N, Warnick GR, and Dominiczak MH, editors. Handbook of Lipoprotein Testing. 2nd ed. Washington: AACC press p.176. |
| 14 | Tietz NW, ed. Clinical Guide to Laboratory Tests, 3rd ed. Philadelphia PA: WB Saunders Company 1995;130-131. |
| 15 | Use of Anticoagulants in Diagnostic Laboratory Investigations. WHO Publication WHO/DIL/LAB/99.1 Rev. 2. Jan. 2002. |
| 16 | Siekmann L, Hüskes KP, Breuer H. Determination of cholesterol in serum using mass fragmentography - a reference method in clinical chemistry. Z Anal Chem 1976;279:145-146. |
| 17 | Glick MR, Ryder KW, Jackson SA. Graphical Comparisons of Interferences in Clinical Chemistry Instrumentation. Clin Chem 1986;32:470-475. |
| 18 | Breuer J. Report on the Symposium “Drug effects in Clinical Chemistry Methods”. Eur J Clin Chem Clin Biochem 1996;34:385-386. |
| 19 | Sonntag O, Scholer A. Drug interference in clinical chemistry: recommendation of drugs and their concentrations to be used in drug interference studies. Ann Clin Biochem 2001;38:376-385. |
| 20 | Bakker AJ, Mücke M. Gammopathy interference in clinical chemistry assays: mechanisms, detection and prevention. Clin Chem Lab Med 2007;45(9):1240-1243. |
| 21 | Study Group, European Atherosclerosis Society. Strategies for the prevention of coronary heart disease: A policy statement of the European Atherosclerosis Society. European Heart Journal 1987;8:77. |
| 22 | Third Report of the National Cholesterol Education Program (NCEP) Expert Panel on Detection, Evaluation, and Treatment of High Blood Cholesterol in Adults (Adult Treatment Panel III). NIH Publication No 01-3670; May 2001. |
| 23 | Bablok W, Passing H, Bender R, et al. A general regression procedure for method transformation. Application of linear regression procedures for method comparison studies in clinical chemistry, Part III. J Clin Chem Clin Biochem 1988 Nov;26(11):783-790. |

- Alternative method

Refer to Brown Clinic Back-up Testing Policy

- Source document

Reagent Name: CHOL2

Method Sheet Version: V12.0 English

Order information

|  |  |  |  | Analyzer(s) on which cobas c pack(s) can be used |  |
| --- | --- | --- | --- | --- | --- |
| 03039773 190 |  | Cholesterol Gen.2 (400 tests) | System-ID 07 6726 3 | Roche/Hitachi cobas c 311, cobas c 501/502 |  |
| 10759350 190 |  | Calibrator f.a.s. (12 x 3 mL) | Code 401 |  |  |
| 10759350 360 |  | Calibrator f.a.s. (12 x 3 mL, for USA) | Code 401 |  |  |
| 12149435 122 |  | Precinorm U plus (10 x 3 mL) | Code 300 |  |  |
| 12149435 160 |  | Precinorm U plus (10 x 3 mL,or USA) | Code 300 |  |  |
| 12149443 122 |  | Precipath U plus (10 x 3 mL) | Code 301 |  |  |
| 12149443 160 |  | Precipath U plus (10 x 3 mL, for USA) | Code 301 |  |  |
| 10171743 122 |  | Precinorm U (20 x 5 mL) | Code 300 |  |  |
| 10171735 122 |  | Precinorm U (4 x 5 mL) | Code 300 |  |  |
| 10171778 122 |  | Precipath U (20 x 5 mL) | Code 301 |  |  |
| 10171760 122 |  | Precipath U (4 x 5 mL) | Code 301 |  |  |
| 10781827 122 |  | Precinorm L (4 x 3 mL) | Code 304 |  |  |
| 11285874 122 |  | Precipath L (4 x 3 mL) | Code 305 |  |  |
| 04489357 190 |  | Diluent NaCl 9 % (50 mL) | System-ID 07 6869 3 |  |  |

Effective date

Effective date for this procedure:

Author

Source documentation compiled by Roche Diagnostics

Revised by:	Heather J Hall, MBA, MT(ASCP), CG(ASCP)cm		Date: 4/9/2018

Approved by: 								Date:

REVIEW – REVISION SUMMARY DOCUMENTATION

Date:	By:			Revision Summary: