Laboratory Name

Test Name: Creatinine Jaffé Gen.2

- System information

For cobas c 311/501 analyzers:

CREJ2: ACN 690 (Rate blanked, compensated, serum and plasma)

CRJ2U: ACN 691 (Rate blanked, urine)

SCRE2: ACN 773 (STAT, compensated, serum and plasma, reaction time: 4)

SCR2U: ACN 774 (STAT, urine, reaction time: 4)

- Intended use

In vitro test for the quantitative determination of creatinine in human serum, plasma and urine on Roche/Hitachi cobas c systems.

Summary1,2,3,4,5

Chronic kidney disease is a worldwide problem that carries a substantial risk for cardiovascular morbidity and death. Current guidelines define chronic kidney disease as kidney damage or glomerular filtration rate (GFR) less than 60 mL/min per 1.73 m2 for three months or more, regardless of cause.

The assay of creatinine in serum or plasma is the most commonly used test to assess renal function. Creatinine is a break-down product of creatine phosphate in muscle, and is usually produced at a fairly constant rate by the body (depending on muscle mass). It is freely filtered by the glomeruli and, under normal conditions, is not re-absorbed by the tubules to any appreciable extent. A small but significant amount is also actively secreted.

Since a rise in blood creatinine is observed only with marked damage of the nephrons, it is not suited to detect early stage kidney disease. A considerably more sensitive test and better estimation of glomerular filtration rate (GFR) is given by the creatinine clearance test based on creatinine’s concentration in urine and serum or plasma, and urine flow rate. For this test a precisely timed urine collection (usually 24 hours) and a blood sample are needed. However, since this test is prone to error due to the inconvenient collection of timed urine, mathematical attempts to estimate GFR based only on the creatinine concentration in serum or plasma have been made. Among the various approaches suggested, two have found wide recognition: that of Cockroft and Gault and that based on the results of the MDRD trial. While the first equation was derived from data obtained with the conventional Jaffé method, a newer version of the second is usable for IDMS-traceable creatinine methods. Both are applicable for adults. In children, the Bedside Schwartz formula should be used.6,7,8,9

In addition to the diagnosis and treatment of renal disease, the monitoring of renal dialysis, creatinine measurements are used for the calculation of the fractional excretion of other urine analytes (e. g., albumin, α-amylase). Numerous methods were described for determining creatinine. Automated assays established in the routine laboratory include the Jaffé alkaline picrate method in various modifications, as well as enzymatic tests.

Test principle10,11,12

This kinetic colorimetric assay is based on the Jaffé method. In alkaline solution, creatinine forms a yellow-orange complex with picrate. The rate of dye formation is proportional to the creatinine concentration in the specimen. The assay uses “rate-blanking” to minimize interference by bilirubin. To correct for non-specific reaction caused by serum/plasma pseudo-creatinine chromogens, including proteins and ketones, the results for serum or plasma are corrected by -26 µmol/L (-0.3 mg/dL).

|  | Alkaline pH |  |
| --- | --- | --- |
| Creatinine + picric acid |  | yellow-orange complex |

- Reagents - working solutions

| R1 | Potassium hydroxide: 900 mmol/L; phosphate: 135 mmol/L; pH ≥ 13.5; preservative; stabilizer |
| --- | --- |
| R3 (STAT R2) | Picric acid: 38 mmol/L; pH 6.5; non reactive buffer |

R1 is in position B and R3 (STAT R2) is in position C.

- Precautions and warnings

For in vitro diagnostic use.
Exercise the normal precautions required for handling all laboratory reagents.
Disposal of all waste material should be in accordance with local guidelines.
Safety data sheet available for professional user on request.

For USA: For prescription use only.

This kit contains components classified as follows in accordance with the Regulation (EC) No. 1272/2008:

Danger

| H314 | Causes severe skin burns and eye damage. |
| --- | --- |

| H412 | Harmful to aquatic life with long lasting effects. |
| --- | --- |

| EUH 001 | Explosive when dry |
| --- | --- |

Prevention:

| P273 | Avoid release to the environment. |
| --- | --- |

| P280 | Wear protective gloves/ protective clothing/ eye protection/ face protection. |
| --- | --- |

Response:

| P301 + P330 + P331 | IF SWALLOWED: Rinse mouth. Do NOT induce vomiting. |
| --- | --- |

| P303 + P361 + P353 | IF ON SKIN (or hair): Take off immediately all contaminated clothing. Rinse skin with water/shower. |
| --- | --- |

| P304 + P340 + P310 | IF INHALED: Remove person to fresh air and keep comfortable for breathing. Immediately call a POISON CENTER/ doctor. |
| --- | --- |

| P305 + P351 + P338 + P310 | IF IN EYES: Rinse cautiously with water for several minutes. Remove contact lenses, if present and easy to do. Continue rinsing. Immediately call a POISON CENTER/ doctor. |
| --- | --- |

Product safety labeling primarily follows EU GHS guidance.

Contact phone: 1-800-428-2336

- Reagent handling

Ready for use

- Storage and stability

| CREJ2 |  |
| --- | --- |
| Shelf life at 15-25 °C: | See expiration date on cobas c pack label. |
| On-board in use and refrigerated on the analyzer: | 8 weeks |

| Diluent NaCl 9 % |  |
| --- | --- |
| Shelf life at 2-8 °C: | See expiration date on cobas c pack label. |
| On-board in use and refrigerated on the analyzer: | 12 weeks |

Specimen collection and preparation13

For specimen collection and preparation only use suitable tubes or collection containers.

Only the specimens listed below were tested and found acceptable.

Serum.

Plasma: Li-heparin and K2-EDTA plasma.

The sample types listed were tested with a selection of sample collection tubes that were commercially available at the time of testing, i.e. not all available tubes of all manufacturers were tested. Sample collection systems from various manufacturers may contain differing materials which could affect the test results in some cases. When processing samples in primary tubes (sample collection systems), follow the instructions of the tube manufacturer.

Urine.

Collect urine without using additives. If urine must be collected with a preservative for other analytes, only hydrochloric acid (14 to 47 mmol/L urine, e.g. 5 mL 10 % HCl or 5 mL 30 % HCl per liter urine) or boric acid (81 mmol/L, e.g. 5 g per liter urine) may be used.

| Stability in serum/plasma:14 | 7 days at 15-25 °C |
| --- | --- |
|  | 7 days at 2-8 °C |
|  | 3 months at (-15)-(-25) °C |
| Stability in urine (without preservative):14 | 2 days at 15-25 °C |
|  | 6 days at 2-8 °C |
|  | 6 months at (-15)-(-25) °C |

Centrifuge samples containing precipitates before performing the assay.

- Materials provided

See “Reagents – working solutions” section for reagents.

- Materials required (but not provided)

| ▪ | See “Order information” section |
| --- | --- |
| ▪ | General laboratory equipment |

In addition, other suitable control material can be used.

- Assay

For optimum performance of the assay follow the directions given in this document for the analyzer concerned. Refer to the appropriate operator’s manual for analyzer-specific assay instructions.

The performance of applications not validated by Roche is not warranted and must be defined by the user.

- Application for serum and plasma

| cobas c 501/502 test definition |  |  |  |
| --- | --- | --- | --- |
| Assay type | Rate A |  |  |
| Reaction time / Assay points | 10 / 42-52 - 24-34 (STAT 4 / 17-27) |  |  |
| Wavelength (sub/main) | 570/505 nm |  |  |
| Reaction direction | Increase |  |  |
| Units | µmol/L (mg/dL, mmol/L) |  |  |
| Reagent pipetting |  | Diluent (H2O) |  |
| R1 | 13 µL | 77 µL |  |
| R3 | 17 µL | 30 µL |  |
| Sample volumes | Sample | Sample dilution |  |
|  |  | Sample | Diluent (NaCl) |
| Normal | 10 µL | – | – |
| Decreased | 10 µL | 20 µL | 80 µL |
| Increased | 10 µL | – | – |

Enter the correction value for the non-specific protein reaction as the instrument factor y = ax + b for mg/dL or for µmol/L, where a = 1.0 and b = -0.3 (mg/dL) or a = 1.0 and b = -26 (µmol/L).

Application for urine

| cobas c 501 test definition |  |  |  |
| --- | --- | --- | --- |
| Assay type | Rate A |  |  |
| Reaction time / Assay points | 10 / 42-52 - 24-34 (STAT 4 / 17-27) |  |  |
| Wavelength (sub/main) | 570/505 nm |  |  |
| Reaction direction | Increase |  |  |
| Units | µmol/L (mg/dL, mmol/L) |  |  |
| Reagent pipetting |  | Diluent (H2O) |  |
| R1 | 13 µL | 77 µL |  |
| R3 | 17 µL | 30 µL |  |
| Sample volumes | Sample | Sample dilution |  |
|  |  | Sample | Diluent (NaCl) |
| Normal | 10 µL | 6 µL | 144 µL |
| Decreased | 10 µL | 2 µL | 180 µL |
| Increased | 10 µL | 6 µL | 144 µL |

- Calibration

| Calibrators | S1: H2O S2: C.f.a.s. |
| --- | --- |
| Calibration mode | Linear |
| Calibration frequency | 2-point calibration after reagent lot change as required following quality control procedures |

Traceability: This method has been standardized against ID/MS.

For the USA, this method has been standardized against a primary reference material (SRM 914 and SRM 967 (ID/MS)).

- Quality control

At least once daily run solutions at two levels of a quality control material with known concentrations.

Refer to Brown Clinic Quality Control Requirements, Rules and Reviews Policy

Refer to Brown Clinic Quality Control Specialty and Subspecialty Policy

- Calculation

Roche/Hitachi cobas c systems automatically calculate the analyte concentration of each sample.

| Conversion factors: | µmol/L x 0.0113 = mg/dL |
| --- | --- |
|  | µmol/L x 0.001 = mmol/L |

- Limitations – interference

Criterion: Recovery within ± 10 % of initial value at a creatinine concentration of 80 µmol/L (0.90 mg/dL) in serum/plasma and 2500 µmol/L (28.3 mg/dL) in urine.

Serum/plasma

Icterus (CREJ2):16 No significant interference up to an I index of 5 for conjugated bilirubin and 10 for unconjugated bilirubin (approximate conjugated bilirubin concentration: 86 µmol/L or 5 mg/dL; approximate unconjugated bilirubin concentration: 171 µmol/L or 10 mg/dL).

Icterus (SCRE2):16 No significant interference up to an I index of 2 for conjugated bilirubin and 3 for unconjugated bilirubin (approximate conjugated bilirubin concentration: 34 µmol/L or 2 mg/dL; approximate unconjugated bilirubin concentration: 51 µmol/L or 3 mg/dL).

Hemolysis:16 No significant interference up to an H index of 1000 (approximate hemoglobin concentration: 621 µmol/L or 1000 mg/dL).

Lipemia (Intralipid):16 No significant interference up to an L index of 800. There is poor correlation between the L index (corresponds to turbidity) and triglycerides concentration.

Drugs: No interference was found at therapeutic concentrations using common drug panels.17,18

Exception: Antibiotics containing cephalosporin lead to significant false-positive values.19,20

Exception: Cefoxitin causes artificially high creatinine results.

Exception: Cyanokit (Hydroxocobalamin) may cause interference with results.

Values < 15 µmol/L (< 0.17 mg/dL) or negative results are reported in rare cases in children < 3 years and in elderly patients. In such cases use the Creatinine plus test to assay the sample.

Do not use Creatinine Jaffé for the testing of creatinine in hemolyzed samples from neonates, infants or adults with HbF levels ≥ 60 mg/dL for CREJ2 applications (≥ 30 mg/dL for SCRE2 applications).21 In such cases, use the Creatinine plus test (≤ 600 mg/dL HbF) to assay the sample.

Estimation of the Glomerular Filtration Rate (GFR) on the basis of the Schwartz Formula can lead to an overestimation.22

In very rare cases, gammopathy, in particular type IgM (Waldenström’s macroglobulinemia), may cause unreliable results.23

The presence of ketone bodies can cause artificially high results in serum and plasma.

Urine

Icterus: No significant interference up to a conjugated bilirubin concentration of 855 µmol/L or 50 mg/dL.

Hemolysis: No significant interference up to a hemoglobin concentration of 621 µmol/L or 1000 mg/dL.

Glucose < 120 mmol/L (< 2162 mg/dL) and urobilinogen < 676 µmol/L (< 40 mg/dL) do not interfere.

Drugs: No interference was found at therapeutic concentrations using common drug panels.18

Exception: Cyanokit (Hydroxocobalamin) may cause interference with results.

High homogentisic acid concentrations in urine samples lead to false results.

The presence of ketone bodies can cause artificially high results in urine.

For diagnostic purposes, the results should always be assessed in conjunction with the patient’s medical history, clinical examination and other findings.

ACTION REQUIRED

Special Wash Programming: The use of special wash steps is mandatory when certain test combinations are run together on Roche/Hitachi cobas c systems. The latest version of the carry-over evasion list can be found with the NaOHD-SMS-SmpCln1+2-SCCS Method Sheets. For further instructions refer to the operator’s manual. cobas c 502 analyzer: All special wash programming necessary for avoiding carry-over is available via the cobas link, manual input is not required.

Where required, special wash/carry-over evasion programming must be implemented prior to reporting results with this test.

- Limits and ranges

Measuring range

Serum/plasma

15-2200 µmol/L (0.17-24.9 mg/dL)

The technical limit in the instrument setting is defined as 41-2226 µmol/L (0.463-25.2 mg/dL) due to the compensation factor of 26.

Determine samples having higher concentrations via the rerun function. Dilution of samples via the rerun function is a 1:5 dilution. Results from samples diluted using the rerun function are automatically multiplied by a factor of 5.

Urine

375-55000 µmol/L (4.2-622 mg/dL)

Determine samples having higher concentrations via the rerun function. Dilution of samples via the rerun function is a 1:3.6 dilution. Results from samples diluted using the rerun function are automatically multiplied by a factor of 3.6.

Lower limits of measurement

Limit of Blank and Limit of Detection

Serum/plasma (CREJ2)

Limit of Blank = 15 µmol/L (0.17 mg/dL)

Limit of Detection = 15 µmol/L (0.17 mg/dL)

The Limit of Blank and Limit of Detection were determined in accordance with the CLSI (Clinical and Laboratory Standards Institute) EP17-A requirements.

The Limit of Blank is the 95th percentile value from n ≥ 60 measurements of analyte-free samples over several independent series. The Limit of Blank corresponds to the concentration below which analyte-free samples are found with a probability of 95 %.

The Limit of Detection is determined based on the Limit of Blank and the standard deviation of low concentration samples.

The Limit of Detection corresponds to the lowest analyte concentration which can be detected (value above the Limit of Blank with a probability of 95 %).

Lower detection limit of the test

Serum/plasma (SCRE2)

15 µmol/L (0.17 mg/dL)

The lower detection limit represents the lowest measurable analyte level that can be distinguished from zero. It is calculated as the value lying 3 standard deviations above that of the lowest standard (standard 1 + 3 SD, repeatability, n = 21).

Urine (CRJ2U/SCR2U)

375 µmol/L (4.2 mg/dL)

The lower detection limit represents the lowest measurable analyte level that can be distinguished from zero. It is calculated as the value lying 3 standard deviations above that of the lowest standard (standard 1 + 3 SD, repeatability, n = 21).

- Expected values

| Serum/plasma |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Adults24 |  |  |  |  |  |
|  | Females |  | 44-80 µmol/L |  | (0.50-0.90 mg/dL) |
|  | Males |  | 62-106 µmol/L |  | (0.70-1.20 mg/dL) |
| Children25 |  |  |  |  |  |
|  | Neonates (premature) |  | 25-91 µmol/L |  | (0.29-1.04 mg/dL) |
|  | Neonates (full term) |  | 21-75 µmol/L |  | (0.24-0.85 mg/dL) |
|  | 2-12 m |  | 15-37 µmol/L |  | (0.17-0.42 mg/dL) |
|  | 1- < 3 y |  | 21-36 µmol/L |  | (0.24-0.41 mg/dL) |
|  | 3- < 5 y |  | 27-42 µmol/L |  | (0.31-0.47 mg/dL) |
|  | 5- < 7 y |  | 28-52 µmol/L |  | (0.32-0.59 mg/dL) |
|  | 7- < 9 y |  | 35-53 µmol/L |  | (0.40-0.60 mg/dL) |
|  | 9- < 11 y |  | 34-65 µmol/L |  | (0.39-0.73 mg/dL) |
|  | 11- < 13 y |  | 46-70 µmol/L |  | (0.53-0.79 mg/dL) |
|  | 13- < 15 y |  | 50-77 µmol/L |  | (0.57-0.87 mg/dL) |
| Urine |  |  |  |  |  |
| 1st morning urine24 |  |  |  |  |  |
| Females |  | 2470-19200 µmol/L |  | (28-217 mg/dL) |  |
| Males |  | 3450-22900 µmol/L |  | (39-259 mg/dL) |  |
| 24-hour urine26 |  |  |  |  |  |
| Females |  | 7000-14000 µmol/24 h |  | (740-1570 mg/24 h) |  |
| Males |  | 9000-21000 µmol/24 h |  | (1040-2350 mg/24 h) |  |
| Creatinine clearance26,27 |  | 71-151 mL/min |  |  |  |

Refer to reference for a prospective study on creatinine clearance in children.28

Roche has not evaluated reference ranges in a pediatric population.

Each laboratory should investigate the transferability of the expected values to its own patient population and if necessary determine its own reference ranges.

- Specific performance data

For Known Interfering Substances section refer to package insert.

For Known Non-Interfering Substance refer to package insert.

For Additional Technical Information refer to package insert.

- References

| 1 | Thomas C, Thomas L. Labordiagnostik von Erkrankungen der Nieren und ableitenden Harnwege. In: Thomas L, ed. Labor und Diagnose, 6th ed. Frankfurt/Main: TH-Books 2005;520-585. |
| --- | --- |
| 2 | Lamb E, Newman DJ, Price CP. Kidney function tests In: Burtis CA, Ashwood ER, Bruns DE. Tietz textbook of clinical chemistry and molecular diagnostics. 4th ed. St.Louis, MO: Elsevier Saunders 2006;797-835. |
| 3 | http://www.kidney.org/ |
| 4 | http://www.nkdep.nih.gov/ |
| 5 | Lamb EJ, Tomson CRV, Roderick PJ. Estimating kidney function in adults using formulae. Ann Clin Biochem 2005;42:321-345. |
| 6 | Miller WG. Editorial on Estimating glomerular filtration rate. Clin Chem Lab Med 2009;47(9):1017-1019. |
| 7 | Schwartz GJ, Muñoz A, Schneider MF, et al. New Equations to Estimate GFR in Children with CKD. J Am Soc Nephrol 2009;20:629-637. |
| 8 | Schwartz GJ, Work DF. Measurement and Estimation of GFR in Children and Adolescents. Clin J Am Soc Nephrol 2009;4:1832–1843. |
| 9 | Staples A, LeBlond R, Watkins S, et al. Validation of the revised Schwartz estimating equation in a predominantly non-CKD population. Pediatr Nephrol 2010 Jul 22;25:2321-2326. |
| 10 | Jaffé M. Ueber den Niederschlag, welchen Pikrinsäure in normalem Harn erzeugt und über eine neue Reaktion des Kreatinins. Z Physiol Chem 1886;10:391-400. |
| 11 | Fabiny DL, Ertinghausen G. Automated reaction-rate method for determination of serum creatinine with the CentrifiChem Clin Chem. 1971;17:696-700. |
| 12 | Bartels H, Böhmer M. Micro-determination of creatinine. Clin Chim Acta 1971;32:81-85. |
| 13 | Guder WG, Narayanan S, Wisser H, et al. List of Analytes; Preanalytical Variables. Brochure in: Samples: From the Patient to the Laboratory. Darmstadt: GIT-Verlag 1996. |
| 14 | Guder W, Fonseca-Wollheim W, Ehret W, et al. Die Qualität Diagnostischer Proben, 6. Aufl. Heidelberg: BD Diagnostics, 2009. |
| 15 | Data on file at Roche Diagnostics. |
| 16 | Glick MR, Ryder KW, Jackson SA. Graphical Comparisons of Interferences in Clinical Chemistry Instrumentation. Clin Chem 1986;32:470-475. |
| 17 | Breuer J. Report on the Symposium “Drug effects in Clinical Chemistry Methods”. Eur J Clin Chem Clin Biochem 1996;34:385-386. |
| 18 | Sonntag O, Scholer A. Drug interference in clinical chemistry: recommendation of drugs and their concentrations to be used in drug interference studies. Ann Clin Biochem 2001;38:376-385. |
| 19 | Ducharme MP, Smythe M, Strohs G. Drug-induced alterations in serum creatinine concentrations. Annal Pharmacotherapy 1993;27:622-633. |
| 20 | Kroll MH. Some observations on the reaction mechanism of Cefoxitin and Cephalothin with picrate. Michrochem J 1990;42:241-249. |
| 21 | Mazzachi BC, Phillips JW, Peake MJ. Is the Jaffe creatinine assay suitable for neonates? Clin Biochem Revs 1998;19:82. |
| 22 | Filler G, Priem F, Lepage N, et al. β-Trace Protein, Cystatin C, β2-Microglobulin, and Creatinine Compared for Detecting Impaired Glomerular Filtration Rates in Children. Clin Chem 2002;48:729-736. |
| 23 | Bakker AJ, Mücke M. Gammopathy interference in clinical chemistry assays: mechanisms, detection and prevention. Clin Chem Lab Med 2007;45(9):1240-1243. |
| 24 | Mazzachi BC, Peake MJ, Ehrhardt V. Reference Range and Method Comparison Studies for Enzymatic and Jaffé Creatinine Assays in Plasma and Serum and Early Morning Urine. Clin Lab 2000;53-55. |
| 25 | Schlebusch H, Liappis N, Kalina E, et al. High Sensitive CRP and Creatinine: Reference Intervals from Infancy to Childhood. J Lab Med 2002;26:341-346. |
| 26 | Junge W, Wilke B, Halabi A, et al. Determination of reference intervals for serum creatinine, creatinine excretion and creatinine clearance with an enzymatic and a modified Jaffé method. Clin Chim Acta 2004;344:137-148. |
| 27 | Zawta B, Delanghe J, Taes Y, et al. Arithmetic Compensation for Pseudo-Creatinine Interferences of the Creatinine Jaffé Method and its Effect on Creatinine Clearance Results. Clin Chem Part 2, Suppl S June 2001;46(6):487. |
| 28 | Wuyts B, Bernard D, van den Noortgate N, et al Reevaluation of Formulas for Predicting Creatinine Clearance in Adults and Children Using Compensated Creatinine Methods. Clin Chem 2003;49:1011-1014. |
| 29 | Bablok W, Passing H, Bender R, et al. A general regression procedure for method transformation. Application of linear regression procedures for method comparison studies in clinical chemistry, Part III. J Clin Chem Clin Biochem 1988 Nov;26(11):783-790. |

- Alternative method

Refer to Brown Clinic Back-up Testing Policy

- Source document

Reagent Name: CREJ2

Method Sheet Version: V18.0 English

____________________________________________________________________________________

Order information

|  |  |  |  | Analyzer(s) on which cobas c pack(s) can be used |  |
| --- | --- | --- | --- | --- | --- |
| 04810716 190 |  | Creatinine Jaffé Gen.2 700 tests | System ID 07 6928 2 | Roche/Hitachi cobas c 311, cobas c 501/502 |  |
| 10759350 190 |  | Calibrator f.a.s. (12 x 3 mL) | Code 401 |  |  |
| 10759350 360 |  | Calibrator f.a.s. (12 x 3 mL, for USA) | Code 401 |  |  |
| 12149435 122 |  | Precinorm U plus (10 x 3 mL) | Code 300 |  |  |
| 12149435 160 |  | Precinorm U plus (10 x 3 mL, for USA) | Code 300 |  |  |
| 12149443 122 |  | Precipath U plus (10 x 3 mL) | Code 301 |  |  |
| 12149443 160 |  | Precipath U plus (10 x 3 mL, for USA) | Code 301 |  |  |
| 10171743 122 |  | Precinorm U (20 x 5 mL) | Code 300 |  |  |
| 10171778 122 |  | Precipath U (20 x 5 mL) | Code 301 |  |  |
| 03121313 122 |  | Precinorm PUC (4 × 3 mL) | Code 240 |  |  |
| 03121291 122 |  | Precipath PUC (4 × 3 mL) | Code 241 |  |  |
| 04489357 190 |  | Diluent NaCl 9 % (50 mL) | System-ID 07 6869 3 |  |  |

Effective date

Effective date for this procedure:

Author

Source documentation compiled by Roche Diagnostics

Revised by:	Heather J Hall, MBA, MT(ASCP), CG(ASCP)cm		Date: 4/9/2018

Approved by: 								Date:

REVIEW – REVISION SUMMARY DOCUMENTATION

Date:	By:			Revision Summary: